## Calculated Recoveries
Sample Name Date | IDA-13C2 4:2 FTS | IDA-13C2 6:2 FTS | IDA-13C2 8:2 FTS | IDA-13C2_PFDoA | IDA-13C2_PFTeDA | IDA-13C3_HFPO-DA | IDA-13C3_PFBS | IDA-13C3_PFHxS | IDA-13C4_PFBA | IDA-13C4_PFHpA | IDA-13C5_PFHxA | IDA-13C5_PFPeA | IDA-13C6_PFDA | IDA-13C7_PFUdA | IDA-13C8_FOSA | IDA-13C8_PFOA | IDA-13C8_PFOS | IDA-13C9_PFNA | IDA-d-EtFOSA | IDA-d-MeFOSA | IDA-d3-MeFOSAA | IDA-d5-EtFOSAA
PFAS QCS0 0ng/ml_09/17/2024 15:14:43 | 0.69039 | 0.64361 | 1.1258 | 1.2174 | 0.7554 | 0.95471 | 0.91058 | 1.0008 | 0.93295 | 1.0501 | 0.95278 | 0.98227 | 1.3087 | 1.1894 | 1.008 | 1.0584 | 0.93054 | 1.0359 | 0.24696 | 0.22942 | 1.1785 | 1.1979
PFAS QCS0 0ng/ml_09/18/2024 15:10:07 | 0.78197 | 0.74123 | 1.0983 | 0.74859 | 0.69484 | 0.80408 | 0.96403 | 0.99446 | 1.0263 | 0.92002 | 0.78935 | 0.98977 | 0.92757 | 0.89289 | 0.95255 | 0.85577 | 0.95257 | 0.9823 | 0.52359 | 0.55818 | 1.154 | 1.2634
PFBA_72hr_Dup_09/18/2024 13:11:46 | 0.9059 | 1.4089 | 2.2983 | 1.2159 | 1.0795 | 1.0967 | 1.1411 | 1.296 | 1.0821 | 1.2799 | 1.0392 | 1.0471 | 1.3679 | 1.1377 | 1.228 | 1.1967 | 1.1685 | 1.2173 | 1.4898 | 1.05 | 1.8405 | 1.6629
PFBS_72hr_Dup_09/18/2024 13:45:46 | 1.4712 | 1.5172 | 2.4544 | 2.5629 | 2.4157 | 2.2188 | 1.8409 | 2.0116 | 1.9922 | 2.3286 | 2.0325 | 1.8685 | 2.8962 | 2.6583 | 2.1513 | 2.2928 | 1.8864 | 2.327 | 2.668 | 1.8209 | 2.8033 | 2.8241
PFOA_72hr_09/18/2024 13:28:22 | 0.65794 | 0.65787 | 1.3063 | 1.0921 | 1.1119 | 0.93282 | 0.85113 | 0.91818 | 0.94848 | 1.0296 | 0.87916 | 0.89657 | 1.1932 | 1.1885 | 0.96078 | 0.96182 | 0.809 | 0.96531 | 1.1363 | 0.78533 | 1.3667 | 1.2801
PFOS_72hr_Dup_09/18/2024 14:02:27 | 1.6047 | 1.5793 | 2.3423 | 2.7298 | 2.447 | 2.1514 | 1.8878 | 2.1275 | 2.0681 | 2.4414 | 2.1624 | 1.9704 | 2.6709 | 2.5658 | 2.1373 | 2.3134 | 1.9179 | 2.3241 | 2.6741 | 1.9029 | 2.7049 | 2.6648
Standard_check_327ai_343g_09/18/2024 14:53:30 | 0.48822 | 0.49577 | 0.70984 | 0.9548 | 0.82696 | 0.81097 | 0.65412 | 0.71835 | 0.72965 | 0.79431 | 0.74143 | 0.70202 | 0.9226 | 0.81396 | 0.72556 | 0.81803 | 0.68418 | 0.82734 | 0.8799 | 0.65269 | 0.71264 | 0.74824
Standard_check_327aj_343g_09/18/2024 14:19:06 | 0.73995 | 0.74567 | 1.0895 | 1.0212 | 1.0813 | 0.93214 | 0.93156 | 0.97572 | 0.94751 | 1.0435 | 0.89399 | 0.92145 | 1.1812 | 1.2156 | 1.0752 | 0.99407 | 0.91937 | 0.92387 | 1.3959 | 0.99405 | 1.2915 | 1.2655
Standard_check_327aj_343o_09/18/2024 14:36:12 | 1.4128 | 1.4356 | 1.9856 | 2.4578 | 2.6061 | 2.1925 | 1.7628 | 1.9869 | 1.9023 | 2.2772 | 2.1087 | 1.8788 | 2.2361 | 2.4995 | 2.0051 | 2.296 | 1.8178 | 2.2971 | 2.5081 | 1.7952 | 2.4949 | 2.2829

## Reported Recoveries
Sample Name Date | IDA-13C2 4:2 FTS | IDA-13C2 6:2 FTS | IDA-13C2 8:2 FTS | IDA-13C2_PFDoA | IDA-13C2_PFTeDA | IDA-13C3_HFPO-DA | IDA-13C3_PFBS | IDA-13C3_PFHxS | IDA-13C4_PFBA | IDA-13C4_PFHpA | IDA-13C5_PFHxA | IDA-13C5_PFPeA | IDA-13C6_PFDA | IDA-13C7_PFUdA | IDA-13C8_FOSA | IDA-13C8_PFOA | IDA-13C8_PFOS | IDA-13C9_PFNA | IDA-d-EtFOSA | IDA-d-MeFOSA | IDA-d3-MeFOSAA | IDA-d5-EtFOSAA
PFAS QCS0 0ng/ml_09/17/2024 15:14:43 | 0.69039 | 0.64361 | 1.1258 | 1.2777 | 0.79284 | 1.002 | 0.91058 | 1.0008 | 0.93295 | 1.1022 | 0.7989 | 0.98227 | 1.3736 | 1.2484 | 1.008 | 1.1108 | 0.93054 | 1.0872 | 0.24696 | 0.22942 | 1.1785 | 1.1979
PFAS QCS0 0ng/ml_09/18/2024 15:10:07 | 0.78197 | 0.74123 | 1.0983 | 0.94837 | 0.88028 | 1.0187 | 0.96403 | 0.99446 | 1.0263 | 1.1655 | 0.66186 | 0.98977 | 1.1751 | 1.1312 | 0.95255 | 1.0842 | 0.95257 | 1.2444 | 0.52359 | 0.55818 | 1.154 | 1.2634
PFBA_72hr_Dup_09/18/2024 13:11:46 | 0.9059 | 1.4089 | 2.2983 | 1.17 | 1.0388 | 1.0553 | 1.1411 | 1.296 | 1.0821 | 1.2316 | 2.0674 | 1.0471 | 1.3163 | 1.0948 | 1.228 | 1.1516 | 1.1685 | 1.1714 | 1.4898 | 1.05 | 1.8405 | 1.6629
PFBS_72hr_Dup_09/18/2024 13:45:46 | 1.4712 | 1.5172 | 2.4544 | 1.2609 | 1.1885 | 1.0916 | 1.8409 | 2.0116 | 1.9922 | 1.1456 | 2.0542 | 1.8685 | 1.4249 | 1.3079 | 2.1513 | 1.1281 | 1.8864 | 1.1449 | 2.668 | 1.8209 | 2.8033 | 2.8241
PFOA_72hr_09/18/2024 13:28:22 | 0.65794 | 0.65787 | 1.3063 | 1.2423 | 1.2647 | 1.061 | 0.85113 | 0.91818 | 0.94848 | 1.1711 | 2.0085 | 0.89657 | 1.3573 | 1.3519 | 0.96078 | 1.094 | 0.809 | 1.098 | 1.1363 | 0.78533 | 1.3667 | 1.2801
PFOS_72hr_Dup_09/18/2024 14:02:27 | 1.6047 | 1.5793 | 2.3423 | 1.2624 | 1.1316 | 0.99493 | 1.8878 | 2.1275 | 2.0681 | 1.129 | 2.1666 | 1.9704 | 1.2352 | 1.1866 | 2.1373 | 1.0698 | 1.9179 | 1.0748 | 2.6741 | 1.9029 | 2.7049 | 2.6648
Standard_check_327ai_343g_09/18/2024 14:53:30 | 0.48822 | 0.49577 | 0.70984 | 1.2878 | 1.1154 | 1.0938 | 0.65412 | 0.71835 | 0.72965 | 1.0713 | 1.239 | 0.70202 | 1.2444 | 1.0978 | 0.72556 | 1.1033 | 0.68418 | 1.1159 | 0.8799 | 0.65269 | 0.71264 | 0.74824
Standard_check_327aj_343g_09/18/2024 14:19:06 | 0.73995 | 0.74567 | 1.0895 | 1.1423 | 1.2095 | 1.0427 | 0.93156 | 0.97572 | 0.94751 | 1.1672 | 2.1799 | 0.92145 | 1.3213 | 1.3598 | 1.0752 | 1.1119 | 0.91937 | 1.0334 | 1.3959 | 0.99405 | 1.2915 | 1.2655
Standard_check_327aj_343o_09/18/2024 14:36:12 | 1.4128 | 1.4356 | 1.9856 | 1.1655 | 1.2358 | 1.0397 | 1.7628 | 1.9869 | 1.9023 | 1.0799 | 2.1566 | 1.8788 | 1.0604 | 1.1853 | 2.0051 | 1.0888 | 1.8178 | 1.0893 | 2.5081 | 1.7952 | 2.4949 | 2.2829

## MQL_Values
| Sample ID | Sample Type | PFBA | PFPeA | PFHxA | PFHpA | PFOA | PFNA | PFDA | PFUdA | PFDoA | PFTrDA | PFTeDA | 5:3 FTCA | 7:3 FTCA | 9:3 FTCA | PFPrSA | FPeSA | PFEA | PFPrA | PFBS | PFPeS | PFHxS | PFHxS_branched | PFHpS | PFOS | PFOS_branched | PFNS | PFDS | MeFBSA | FBSA | FHxSA | FOSA | N-MeFOSAA | N-MeFOSAA_branched | N-EtFOSAA | N-EtFOSAA_branched | MeFOSA | EtFOSA | HFPO-DA | DONA | 9ClPF3ONS | 11ClPF3OUdS | 3,6-OPFHpA | PF4OPeA | PF5OHxA | PFECHS | PFEESA | 4:2 FTS | 6:2 FTS | 8:2 FTS | 10:2 FTS
MQL | 0 | 0 | 0.029767 | 0.020496 | 0.0030158 | 0.0016775 | 0.005176 | 0.0043025 | 0.003482 | 0.0053122 | 0.0016614 | 0.00051296 | 0.020789 | 0 | 0.0040371 | 0 | 0 | 0 | 0 | 0 | 0.014976 | 0.0058103 | 0.013645 | 0.070159 | 0.010187 | 0.022131 | 0.042776 | 0.020544 | 0.02265 | 0.018093 | 0.027742 | 0.0011706 | 0.0041309 | 0.02068 | 0.33487 | 0.012544 | 0.37524 | 0.0069635 | 0.01626 | 0.0082165 | 0.0013446 | 0.00020026 | 0.0030472 | 0.0042148 | 0.0088009 | 0.013214 | 0.00022287 | 0.015198 | 0.036083 | 0.00061379 | 0.0017626 | 0

## Concentration_filtered_MDL
Sample Name | 4:2 FTS_TOF MS | Br-N-MeFOSAA_TOF MS | Br-PFHxS_TOF MS | Br-PFOS | Br-PFOS_TOF MS | EtFOSA_TOF MS | FHpSA_TOF MS | FOSA | FOSA_TOF MS | L-N-MeFOSAA_TOF MS | L-PFHxS_TOF MS | L-PFOS | L-PFOS_TOF MS | MeFOSA_TOF MS | PFBA | PFBA_TOF MS | PFBS | PFBS_TOF MS | PFDoA_TOF MS | PFHpA | PFHpA_TOF MS | PFHpS_TOF MS | PFHxA | PFHxA_TOF MS | PFOA | PFOA_TOF MS
PFAS QCS0 0ng/ml | 0.0489599086416034 | 0.833257294964726 | 0.0190736060125838 | <MQL | 0.00220304523338136 | 0.0651352227034669 | 0.0136137807512454 | <MQL | <MQL | 1.15203968208113 | 0.0202188283844378 | <MQL | 0.0060187568061587 | <MQL | <MQL | <MQL | <MQL | <MQL | <MQL | 0.0353564961081969 | 0.0220117239549493 | <MQL | <MQL | 0.0151374996161742 | 0.0162671931307739 | 0.00783657244185009
PFBA_72hr_Dup | 0.0563922085361553 | 0.581769864717968 | 0.00175802388493325 | <MQL | <MQL | 0.00915581169209159 | 0.0131971597459451 | <MQL | <MQL | 0.80433975681238 | 0.00214384479019904 | <MQL | <MQL | <MQL | 1.78099064689806 | 1.36788815947709 | <MQL | <MQL | <MQL | 0.00329909834357048 | <MQL | <MQL | 0.012862582429304 | 0.0103553126070929 | <MQL | <MQL
PFBS_72hr_Dup | 0.053646516584077 | 0.58303241837264 | 0.00188911068613336 | <MQL | <MQL | 0.00391612301669401 | 0.00716342761763115 | <MQL | 0.00208720345773131 | 0.806085330382187 | 0.00200253715720763 | <MQL | <MQL | <MQL | <MQL | <MQL | 0.995083222314166 | 0.908592111628908 | 0.00474273815948231 | 0.00350146886802616 | <MQL | <MQL | <MQL | <MQL | <MQL | <MQL
PFOA_72hr | 0.0620721312888851 | 0.485541603859533 | 0.00344476232055755 | <MQL | <MQL | 0.00493477411691998 | 0.0148289313319647 | <MQL | <MQL | 0.671297087139426 | 0.00365159362831443 | <MQL | <MQL | 0.00239456902704659 | <MQL | <MQL | <MQL | <MQL | <MQL | 0.00974767629483499 | 0.00833436020033182 | <MQL | 0.0162205061249421 | 0.0110691684000553 | 4.2034233950993 | 4.08329257635442
PFOS_72hr_Dup | 0.0520494328883192 | 0.659475222118103 | 0.002030942553406 | 0.148578268365655 | 0.073582676293974 | 0.00691724375166028 | 0.00690965068312112 | <MQL | <MQL | 0.911773145966256 | 0.00215288493003768 | 0.983040087137568 | 0.673828768873611 | <MQL | <MQL | <MQL | <MQL | <MQL | 0.00331126654452556 | 0.00545501410327587 | 0.00484133986096783 | 0.0178095398189149 | 0.0208537360611477 | 0.0127002798609048 | <MQL | 0.00309929386493177
Standard_check_327ai_343g | 0.0279218909711343 | 0.383839238695754 | 0.0023068104781044 | <MQL | <MQL | 0.00413720223730037 | 0.00932967617931215 | <MQL | 0.00278076392024542 | 0.530686064423881 | 0.00244531658782531 | <MQL | 0.000442888899288632 | <MQL | <MQL | <MQL | <MQL | <MQL | <MQL | <MQL | 0.00281690457070918 | <MQL | <MQL | <MQL | <MQL | <MQL
Standard_check_327aj_343g | 0.0488814210221799 | 0.599696545849589 | 0.00206798027980715 | <MQL | 0.0021418480474113 | 0.00696500689362919 | 0.01457117491849 | 0.0139661654663983 | 0.00934295988508168 | 0.829124715979784 | 0.00219214648516054 | <MQL | 0.0010101403373563 | <MQL | <MQL | <MQL | <MQL | <MQL | <MQL | <MQL | <MQL | <MQL | 0.00966781769528913 | 0.00668749068038553 | <MQL | <MQL
Standard_check_327aj_343o | 0.0560468957535074 | 0.622327171822095 | 0.00231331620382041 | <MQL | 0.0022169238116042 | 0.00658083585365873 | 0.00937177636320419 | <MQL | 0.00723652513692072 | 0.860413225913281 | 0.00245221293200278 | <MQL | 0.00263855889588486 | <MQL | <MQL | <MQL | <MQL | <MQL | <MQL | <MQL | 0.00347912761268485 | <MQL | <MQL | <MQL | <MQL | <MQL

## Calibration IDA_IPS Ratio
Component Name | Average
IDA-13C2 4:2 FTS | 1.084
IDA-13C2 6:2 FTS | 0.90171
IDA-13C2 8:2 FTS | 1.0659
IDA-13C2_PFDoA | 1.1256
IDA-13C2_PFTeDA | 0.29559
IDA-13C3_HFPO-DA | 1.7784
IDA-13C3_PFBS | 2.6503
IDA-13C3_PFHxS | 3.649
IDA-13C4_PFBA | 1.2435
IDA-13C4_PFHpA | 2.312
IDA-13C5_PFHxA | 2.135
IDA-13C5_PFPeA | 0.91711
IDA-13C6_PFDA | 1.6358
IDA-13C7_PFUdA | 1.3176
IDA-13C8_FOSA | 4.3029
IDA-13C8_PFOA | 2.1745
IDA-13C8_PFOS | 2.2928
IDA-13C9_PFNA | 1.8814
IDA-d-EtFOSA | 1.5715
IDA-d-MeFOSA | 2.1237
IDA-d3-MeFOSAA | 2.7209
IDA-d5-EtFOSAA | 0.85512

## Calibration IDA
| IDA-13C2 4:2 FTS | IDA-13C2 6:2 FTS | IDA-13C2 8:2 FTS | IDA-13C2_PFDoA | IDA-13C2_PFTeDA | IDA-13C3_HFPO-DA | IDA-13C3_PFBS | IDA-13C3_PFHxS | IDA-13C4_PFBA | IDA-13C4_PFHpA | IDA-13C5_PFHxA | IDA-13C5_PFPeA | IDA-13C6_PFDA | IDA-13C7_PFUdA | IDA-13C8_FOSA | IDA-13C8_PFOA | IDA-13C8_PFOS | IDA-13C9_PFNA | IDA-d-EtFOSA | IDA-d-MeFOSA | IDA-d3-MeFOSAA | IDA-d5-EtFOSAA
PFAS CS0 0ng/ml_10/26/2023 19:07:37 | 18596.5495 | 24812.8216 | 30501.4886 | 87765.045 | 19212.3256 | 141445.9959 | 42471.3802 | 61500.3109 | 97153.4926 | 178840.2452 | 158429.6564 | 76418.6589 | 115798.6291 | 100773.8028 | 60041.8992 | 168806.0426 | 38652.0472 | 151860.9441 | 25449.334 | 34430.4513 | 47216.4832 | 18869.7147
PFAS CS0 0ng/ml_10/26/2023 19:24:17 | 18180.9415 | 17262.4149 | 24187.3408 | 89747.2246 | 18716.1191 | 138215.1084 | 42420.1472 | 61149.0723 | 108602.7871 | 182262.6724 | 157116.4641 | 78244.8175 | 129998.6259 | 109332.2039 | 70706.0005 | 174231.0153 | 41418.4752 | 159913.3037 | 28218.7843 | 38139.6779 | 48603.6506 | 17479.3326
PFAS CS0 0ng/ml_10/26/2023 19:40:53 | 17750.6935 | 16130.1415 | 20317.4666 | 94628.8826 | 20918.1168 | 135278.7971 | 42930.7713 | 62650.3786 | 103075.2544 | 183954.9165 | 152804.7565 | 77216.9643 | 143100.1392 | 111616.2338 | 76040.0693 | 167597.5352 | 42141.108 | 150834.9551 | 26331.3845 | 34277.2387 | 47344.6738 | 17498.3439
PFAS CS0 0ng/ml_10/26/2023 19:57:28 | 18857.4612 | 16653.2683 | 19237.4184 | 89224.3172 | 20550.4067 | 130705.7531 | 44579.984 | 64589.1226 | 105545.4962 | 181385.1331 | 155936.9998 | 73872.504 | 134859.2051 | 113852.6749 | 76071.9689 | 173638.9233 | 41004.9219 | 141759.0917 | 23875.3196 | 30149.5032 | 47699.6856 | 17116.6458
PFAS CS1 0.01ng/ml_10/26/2023 20:14:03 | 18479.0831 | 16914.6219 | 19268.3378 | 77881.123 | 14536.4205 | 130633.4279 | 42964.0261 | 62108.3868 | 107217.7308 | 175213.492 | 155213.0294 | 75980.5208 | 128038.8338 | 101408.5655 | 72686.8787 | 165423.7252 | 39946.3381 | 138771.3897 | 28783.5088 | 38587.9088 | 45849.0294 | 15844.5916
PFAS CS1 0.01ng/ml_10/26/2023 20:30:40 | 17866.7775 | 16826.6453 | 19923.9469 | 82921.4798 | 16759.8033 | 128199.647 | 41938.1957 | 63412.1961 | 106321.1967 | 175193.0455 | 154218.1035 | 82574.9492 | 134597.911 | 104930.2366 | 72909.1064 | 170937.2545 | 40851.8543 | 144870.5568 | 29123.7698 | 39313.4288 | 46515.0949 | 15692.1863
PFAS CS1 0.01ng/ml_10/26/2023 20:47:16 | 19739.229 | 16214.7795 | 19782.1836 | 85816.3739 | 18738.7578 | 130855.2263 | 42649.6836 | 64383.7242 | 106593.9063 | 183673.0303 | 161294.6955 | 82633.3049 | 140665.593 | 110757.6381 | 72800.0694 | 177200.8763 | 40097.9594 | 149164.7354 | 26173.1171 | 36313.2281 | 47782.0866 | 16804.0676
PFAS CS1 0.01ng/ml_10/26/2023 21:03:55 | 19110.192 | 15906.0357 | 20771.2166 | 83356.2619 | 17814.8014 | 131031.7655 | 42038.6073 | 63890.2909 | 111108.3154 | 179583.5542 | 161553.4971 | 81476.6115 | 129457.4004 | 96218.2584 | 72019.1197 | 174099.5457 | 41655.5941 | 154678.8832 | 27575.4729 | 35620.8581 | 46269.1669 | 16666.2151
PFAS CS10 400ng/ml_10/27/2023 06:11:54 | 13166.9957 | 7112.7353 | 7850.4904 | 48812.4121 | 19233.2773 | 95135.6842 | 37419.6943 | 43019.6637 | 85973.9445 | 107069.8885 | 108615.1139 | 64440.6531 | 59624.6025 | 45795.5743 | 53134.8678 | 82219.6816 | 24597.8343 | 74771.3352 | 24960.3551 | 29535.0146 | 32519.9633 | 7304.4771
PFAS CS10 400ng/ml_10/27/2023 06:28:34 | 13026.8442 | 8034.4463 | 7904.0525 | 50871.8409 | 20087.9886 | 97879.5092 | 37045.1226 | 42307.0447 | 87254.7848 | 107635.2259 | 115424.7956 | 63583.4064 | 57992.0982 | 51560.5605 | 55442.8643 | 81330.6664 | 23920.2901 | 80115.64 | 23473.641 | 28570.4277 | 33054.2744 | 7694.1609
PFAS CS10 400ng/ml_10/27/2023 06:45:11 | 12119.8269 | 7869.1398 | 8474.8916 | 50238.6121 | 17933.6523 | 98402.621 | 34610.8867 | 40734.6572 | 85433.764 | 108858.4273 | 112722.9189 | 66668.1092 | 56738.2389 | 46182.7705 | 54011.0184 | 75554.6356 | 23636.898 | 70650.253 | 19315.2408 | 26249.1628 | 32019.9584 | 7274.3025
PFAS CS10 400ng/ml_10/27/2023 07:01:47 | 12990.7022 | 7981.3432 | 7911.4781 | 50468.3596 | 18048.2518 | 101045.1207 | 37673.9751 | 42886.951 | 85073.9383 | 107454.4491 | 113562.1062 | 63813.822 | 57851.3485 | 51318.9607 | 54079.4915 | 81710.801 | 23207.9564 | 71980.2942 | 19521.8842 | 26227.4234 | 31428.3636 | 7018.2052
PFAS CS2 0.025ng/ml_10/26/2023 21:20:30 | 16877.5802 | 16506.1195 | 20091.4969 | 83956.7749 | 18632.9803 | 126146.5896 | 41819.2594 | 61576.7301 | 102432.3072 | 168819.848 | 145368.61 | 73719.9563 | 133721.8498 | 98767.2731 | 71568.6856 | 164559.4754 | 38105.3886 | 149211.3745 | 24107.6548 | 36101.6246 | 47369.1963 | 15597.1524
PFAS CS2 0.025ng/ml_10/26/2023 21:37:05 | 17290.931 | 15649.8085 | 18959.1294 | 90519.3394 | 21263.5715 | 120895.5741 | 43885.4107 | 61237.1287 | 108415.8957 | 171696.8046 | 154285.1774 | 79293.1328 | 130865.7387 | 103682.319 | 71004.7902 | 166489.8396 | 38744.5288 | 141987.3193 | 22170.9012 | 33315.0239 | 47401.1243 | 16217.2688
PFAS CS2 0.025ng/ml_10/26/2023 21:53:42 | 17026.2027 | 15844.1194 | 19029.3727 | 88319.8322 | 21543.7286 | 122966.8541 | 43144.377 | 60741.9578 | 105417.4537 | 168857.3388 | 157773.8346 | 75784.2152 | 122791.5024 | 96837.021 | 70054.0719 | 173276.5651 | 40606.0261 | 148605.3884 | 21354.6716 | 30737.1521 | 42812.2805 | 15469.5218
PFAS CS2 0.025ng/ml_10/26/2023 22:10:16 | 19257.3943 | 15465.763 | 19395.697 | 79744.9562 | 20027.8467 | 121342.2117 | 43134.558 | 61301.9867 | 109934.8604 | 165975.6614 | 149721.352 | 72665.6101 | 136487.1799 | 102515.8607 | 73799.0823 | 172311.5865 | 40179.0646 | 137093.1435 | 20117.7652 | 31810.6117 | 44650.5397 | 14370.8439
PFAS CS3 0.1ng/ml_10/26/2023 22:26:52 | 18892.798 | 15987.4022 | 19754.0179 | 90153.0487 | 21170.3152 | 124827.9458 | 45587.3318 | 67139.9143 | 112347.3887 | 178396.4075 | 162951.8732 | 79005.4918 | 135939.7079 | 103444.377 | 74843.2582 | 185428.7604 | 41592.5252 | 153181.9932 | 27717.3684 | 36679.0832 | 47033.2701 | 15673.3107
PFAS CS3 0.1ng/ml_10/26/2023 22:43:27 | 18952.4037 | 16854.7723 | 20574.9998 | 93869.3667 | 24811.2587 | 135051.2952 | 44685.5923 | 65258.5583 | 112286.3297 | 173365.6132 | 165655.8991 | 83004.5799 | 134249.0791 | 104722.1508 | 77082.8566 | 179127.9432 | 42104.1526 | 153694.6053 | 25047.5314 | 37740.6625 | 51035.3943 | 16644.266
PFAS CS3 0.1ng/ml_10/26/2023 23:00:04 | 19674.789 | 17026.1685 | 21344.5502 | 97375.8485 | 24671.5778 | 135912.2254 | 44226.44 | 65054.846 | 109186.7898 | 175054.0958 | 164103.1291 | 78367.7279 | 139427.9998 | 119520.0272 | 76702.1975 | 187909.9715 | 40929.9354 | 157388.2311 | 25528.0278 | 35590.8755 | 46977.9541 | 16295.6705
PFAS CS3 0.1ng/ml_10/26/2023 23:16:42 | 18464.7219 | 17014.3868 | 22881.8392 | 89683.8356 | 22839.1301 | 126571.5132 | 47063.1251 | 66568.4228 | 113697.5993 | 188473.6701 | 163206.4451 | 84290.5181 | 143333.1675 | 111548.2222 | 78834.0919 | 188850.1594 | 42197.2566 | 158341.5313 | 25499.2618 | 37167.5886 | 47205.3489 | 17550.3021
PFAS CS4 0.25ng/ml_10/26/2023 23:33:17 | 18465.1754 | 16399.991 | 21500.3351 | 95420.0438 | 20684.7463 | 122464.8635 | 43272.4717 | 64048.1159 | 105908.7739 | 174740.9698 | 152581.4898 | 79173.0487 | 125802.5949 | 104875.777 | 73359.8217 | 169891.2985 | 40526.0745 | 148202.3447 | 23414.2489 | 34593.6586 | 45452.244 | 14955.0403
PFAS CS4 0.25ng/ml_10/26/2023 23:49:51 | 18707.6833 | 16615.8095 | 20407.0697 | 82434.8518 | 22159.8149 | 121075.6634 | 43189.1701 | 61832.1558 | 107032.9424 | 173052.564 | 149327.9125 | 75656.9024 | 131190.8076 | 119238.0358 | 77622.7051 | 166497.1342 | 39537.8635 | 138842.9236 | 22628.873 | 35560.8103 | 48005.9537 | 16117.444
PFAS CS4 0.25ng/ml_10/27/2023 00:06:26 | 17902.4103 | 15745.9586 | 20528.3818 | 88132.9586 | 23399.7604 | 128152.1013 | 42927.626 | 62319.0598 | 103188.4271 | 169394.9169 | 153302.3179 | 79679.337 | 140508.0097 | 99986.4534 | 68095.2756 | 174116.364 | 40436.7155 | 142236.5531 | 19653.644 | 31333.2564 | 46911.7949 | 15570.71
PFAS CS4 0.25ng/ml_10/27/2023 00:23:00 | 19781.0328 | 16126.2798 | 19418.3038 | 77495.4723 | 18557.8147 | 129833.0647 | 42737.1046 | 62651.2588 | 108643.0531 | 168056.3558 | 154876.6017 | 76941.7127 | 125704.6757 | 106374.609 | 74093.2583 | 177543.7329 | 40682.3813 | 142297.1285 | 20835.1034 | 31353.8768 | 45451.2001 | 15979.5496
PFAS CS5 1ng/ml_10/27/2023 00:39:39 | 18136.81 | 16196.7941 | 19290.7638 | 89085.5844 | 23180.2089 | 124049.213 | 42964.9552 | 60687.3918 | 104425.7097 | 173974.5489 | 151154.1922 | 76359.2956 | 125813.3659 | 106461.9013 | 68993.9029 | 167429.0273 | 39768.555 | 138156.0506 | 24796.2605 | 34719.4377 | 44709.4246 | 14581.8416
PFAS CS5 1ng/ml_10/27/2023 00:56:19 | 17687.9081 | 15334.1234 | 18531.3353 | 85240.3318 | 25176.4138 | 121051.9645 | 40843.6509 | 60507.0471 | 105764.7288 | 165983.8166 | 147146.3666 | 77046.0412 | 123833.1656 | 103970.8211 | 69635.9399 | 165069.6474 | 40782.6799 | 146474.067 | 22559.0169 | 31881.7382 | 44269.4126 | 15007.4201
PFAS CS5 1ng/ml_10/27/2023 01:12:53 | 18179.4988 | 16685.6551 | 17710.3758 | 84148.0839 | 20067.2211 | 120391.596 | 42242.3222 | 61816.9631 | 103514.2252 | 167489.5153 | 150231.1993 | 75627.6054 | 130540.9645 | 91986.0225 | 72623.0597 | 162779.2797 | 38078.6403 | 139864.1004 | 21680.9939 | 29464.2609 | 44564.8902 | 14862.3516
PFAS CS5 1ng/ml_10/27/2023 01:29:29 | 18734.682 | 14934.7715 | 19024.7866 | 80784.9433 | 20686.5725 | 120886.5543 | 42414.9241 | 61842.0875 | 106055.9927 | 164349.8575 | 153927.9712 | 76037.9407 | 126013.4546 | 94174.7869 | 72347.5705 | 166546.9946 | 38134.3896 | 139998.0841 | 20033.779 | 30982.0021 | 44510.3465 | 14607.7761
PFAS CS6 4ng/ml_10/27/2023 01:46:07 | 18050.5527 | 14879.5686 | 16738.9038 | 73632.4221 | 16996.7998 | 122390.3447 | 41284.4982 | 62842.8687 | 105231.9124 | 161171.379 | 148010.5067 | 74908.1656 | 119041.3053 | 97360.1469 | 71158.0786 | 161958.3043 | 38973.7824 | 138605.1891 | 23153.5368 | 33459.5492 | 45083.3356 | 14338.5547
PFAS CS6 4ng/ml_10/27/2023 02:02:41 | 16711.7354 | 15396.3289 | 16212.4131 | 77090.1571 | 18394.2024 | 119714.3816 | 44388.3804 | 61948.2234 | 105709.9514 | 168500.4181 | 150879.7779 | 79168.7964 | 124018.497 | 101870.845 | 74125.3059 | 159316.8195 | 42124.2521 | 147897.5726 | 21079.8956 | 32053.1962 | 44068.7343 | 14438.628
PFAS CS6 4ng/ml_10/27/2023 02:19:16 | 17950.452 | 14958.2535 | 18048.5463 | 79841.8654 | 19469.8065 | 124082.4698 | 43689.6255 | 61653.7787 | 98650.3094 | 165579.7345 | 151302.2724 | 75608.4676 | 108114.514 | 85723.4293 | 69178.6384 | 161931.2028 | 38050.808 | 138176.1994 | 19973.9533 | 30256.4626 | 44494.6909 | 13790.9978
PFAS CS6 4ng/ml_10/27/2023 02:35:53 | 18233.6526 | 15795.0708 | 17163.2099 | 77507.5679 | 19788.4771 | 119925.8399 | 41822.4149 | 61975.6192 | 104899.5987 | 158609.3373 | 148667.2905 | 73891.6482 | 124149.3389 | 89539.5715 | 68661.3366 | 162299.0198 | 39465.2438 | 135797.3077 | 18858.9398 | 30448.3098 | 46499.1531 | 14527.4524
PFAS CS7 10ng/ml_10/27/2023 02:52:28 | 16839.8362 | 12989.1744 | 13010.3994 | 74814.3228 | 21021.1895 | 113749.678 | 40931.4747 | 56316.3847 | 103308.479 | 147500.0559 | 143379.9715 | 71981.4954 | 103270.5643 | 90773.7601 | 64064.8136 | 136602.0582 | 36037.1943 | 122953.2298 | 23847.5536 | 32380.5499 | 40584.5845 | 11856.6009
PFAS CS7 10ng/ml_10/27/2023 03:09:05 | 16755.0878 | 11870.5779 | 14597.0179 | 73063.109 | 20056.1072 | 113878.013 | 41939.6098 | 55833.8533 | 102694.317 | 161627.4657 | 141874.4974 | 71325.7454 | 108383.8286 | 92317.6336 | 65503.1344 | 144206.6392 | 35527.8841 | 121694.5504 | 24591.426 | 33853.7748 | 41509.5152 | 11764.6931
PFAS CS7 10ng/ml_10/27/2023 03:25:43 | 15973.0797 | 13075.6642 | 14416.2919 | 75585.3699 | 22013.5148 | 116264.8512 | 40612.8009 | 58504.515 | 102258.9988 | 153040.6983 | 141529.0144 | 78976.96 | 99364.9933 | 84045.9722 | 66065.4839 | 140705.4504 | 35724.4728 | 126590.1686 | 22344.0768 | 32485.3746 | 42118.1209 | 12788.5111
PFAS CS7 10ng/ml_10/27/2023 03:42:18 | 16620.7224 | 13729.5463 | 13959.2719 | 72798.6808 | 24072.9319 | 116835.0237 | 40419.4586 | 57522.7518 | 98890.4588 | 152975.3688 | 144104.4328 | 73563.5825 | 107461.2291 | 88498.4064 | 65502.9694 | 147077.0943 | 37787.6906 | 130414.6117 | 20560.5088 | 30835.7326 | 41582.068 | 12827.2444
PFAS CS8 40ng/ml_10/27/2023 03:58:55 | 16069.8299 | 10623.5559 | 10213.153 | 67828.7725 | 18368.9268 | 120466.6267 | 42382.3045 | 56889.9734 | 104447.0586 | 144233.6267 | 141170.3218 | 76249.5587 | 85167.1083 | 72255.6146 | 67647.5818 | 116781.2815 | 31793.5476 | 105531.0428 | 30391.7951 | 38724.2588 | 40497.677 | 10422.1934
PFAS CS8 40ng/ml_10/27/2023 04:15:31 | 15534.153 | 10604.5998 | 10881.2456 | 67711.7477 | 19797.3417 | 117848.702 | 41090.8855 | 56768.3318 | 98479.618 | 141670.9016 | 140268.0577 | 77032.685 | 77876.3827 | 65183.4419 | 65085.4989 | 119027.4363 | 31945.4894 | 105676.764 | 29309.8717 | 38333.3254 | 40355.373 | 9453.5744
PFAS CS8 40ng/ml_10/27/2023 04:32:10 | 16546.0501 | 10746.8354 | 11028.8446 | 59576.3479 | 20674.7012 | 123035.7209 | 43772.5864 | 57743.6093 | 106140.7752 | 144745.5202 | 139735.5416 | 76450.7978 | 84701.801 | 69001.6993 | 66074.0897 | 116669.6735 | 32556.0349 | 96260.8396 | 28843.5844 | 36241.3518 | 40173.3288 | 9984.5211
PFAS CS8 40ng/ml_10/27/2023 04:48:50 | 15741.8069 | 10878.1261 | 10735.5922 | 61085.4431 | 20403.4113 | 119810.9004 | 44424.3711 | 56681.7716 | 106732.7587 | 143917.7947 | 143050.7468 | 79823.2614 | 83580.5698 | 76542.2185 | 68270.6224 | 122336.9886 | 32751.252 | 107023.2864 | 26079.203 | 33746.6498 | 41703.6261 | 9843.3646
PFAS CS9 100ng/ml_10/27/2023 05:05:26 | 9661.4262 | 6120.1574 | 6367.8532 | 28200.8749 | 9632.6686 | 67981.5161 | 24933.2453 | 20981.2269 | 55685.5862 | 60519.1053 | 70287.1702 | 43266.2671 | 32256.8936 | 23740.8874 | 32478.8586 | 40428.4221 | 13596.4905 | 38568.3993 | 23028.2578 | 23458.5375 | 21946.7785 | 4347.3703
PFAS CS9 100ng/ml_10/27/2023 05:22:02 | 9503.8182 | 6420.2702 | 7242.0241 | 28639.0836 | 11607.5131 | 71485.5426 | 26159.942 | 22870.54 | 57350.6433 | 66829.4019 | 71082.2641 | 46852.9056 | 34142.7376 | 25061.309 | 35390.6792 | 43045.9773 | 14351.8684 | 42597.4853 | 25981.6791 | 24112.5546 | 21976.2251 | 4498.5589
PFAS CS9 100ng/ml_10/27/2023 05:38:39 | 10436.8482 | 6125.8287 | 9249.8279 | 32210.8352 | 12412.1241 | 70839.6832 | 27138.4293 | 23439.3697 | 58640.5513 | 65035.7132 | 73170.5521 | 47962.8944 | 33432.4184 | 29472.0361 | 37159.4868 | 42708.4881 | 14909.6202 | 43214.9447 | 28665.9384 | 25116.4122 | 22544.5054 | 4550.117
PFAS CS9 100ng/ml_10/27/2023 05:55:18 | 10013.0929 | 6718.9324 | 7840.1314 | 35103.6218 | 13300.7925 | 72186.4849 | 28120.7263 | 23528.445 | 57461.2297 | 67398.2652 | 76547.8814 | 49318.8825 | 36180.435 | 27667.6962 | 38889.6749 | 46920.45 | 14566.4179 | 44422.571 | 28242.6036 | 27150.8636 | 24020.0741 | 5274.9302
Average | 16694.5332 | 13887.2456 | 16415.3976 | 74381.5324 | 19532.5771 | 117521.4469 | 40817.9897 | 56198.1065 | 99413.394 | 152776.383 | 141081.0638 | 73317.2625 | 108092.9206 | 87066.0865 | 66269.5494 | 143695.3134 | 35311.3524 | 124326.5996 | 24202.2668 | 32706.8839 | 41905.0136 | 13169.7733

## Calibration IPS
| IDA-13C2 4:2 FTS | IDA-13C2 6:2 FTS | IDA-13C2 8:2 FTS | IDA-13C2_PFDoA | IDA-13C2_PFTeDA | IDA-13C3_HFPO-DA | IDA-13C3_PFBS | IDA-13C3_PFHxS | IDA-13C4_PFBA | IDA-13C4_PFHpA | IDA-13C5_PFHxA | IDA-13C5_PFPeA | IDA-13C6_PFDA | IDA-13C7_PFUdA | IDA-13C8_FOSA | IDA-13C8_PFOA | IDA-13C8_PFOS | IDA-13C9_PFNA | IDA-d-EtFOSA | IDA-d-MeFOSA | IDA-d3-MeFOSAA | IDA-d5-EtFOSAA
PFAS CS0 0ng/ml_10/26/2023 19:07:37 | 16920.7246 | 16920.7246 | 16920.7246 | 79364.4487 | 79364.4487 | 79364.4487 | 16920.7246 | 16920.7246 | 80046.3212 | 79364.4487 | 79364.4487 | 80046.3212 | 79364.4487 | 79364.4487 | 16920.7246 | 79364.4487 | 16920.7246 | 79364.4487 | 16920.7246 | 16920.7246 | 16920.7246 | 16920.7246
PFAS CS0 0ng/ml_10/26/2023 19:24:17 | 17068.4744 | 17068.4744 | 17068.4744 | 79364.4487 | 79364.4487 | 79364.4487 | 17068.4744 | 17068.4744 | 87164.5549 | 79364.4487 | 79364.4487 | 87164.5549 | 79364.4487 | 79364.4487 | 17068.4744 | 79364.4487 | 17068.4744 | 79364.4487 | 17068.4744 | 17068.4744 | 17068.4744 | 17068.4744
PFAS CS0 0ng/ml_10/26/2023 19:40:53 | 18152.8726 | 18152.8726 | 18152.8726 | 79364.4487 | 79364.4487 | 79364.4487 | 18152.8726 | 18152.8726 | 89313.1499 | 79364.4487 | 79364.4487 | 89313.1499 | 79364.4487 | 79364.4487 | 18152.8726 | 79364.4487 | 18152.8726 | 79364.4487 | 18152.8726 | 18152.8726 | 18152.8726 | 18152.8726
PFAS CS0 0ng/ml_10/26/2023 19:57:28 | 20861.2805 | 20861.2805 | 20861.2805 | 79364.4487 | 79364.4487 | 79364.4487 | 20861.2805 | 20861.2805 | 89691.5494 | 79364.4487 | 79364.4487 | 89691.5494 | 79364.4487 | 79364.4487 | 20861.2805 | 79364.4487 | 20861.2805 | 79364.4487 | 20861.2805 | 20861.2805 | 20861.2805 | 20861.2805
PFAS CS1 0.01ng/ml_10/26/2023 20:14:03 | 18466.4887 | 18466.4887 | 18466.4887 | 77826.5774 | 77826.5774 | 77826.5774 | 18466.4887 | 18466.4887 | 93321.4236 | 77826.5774 | 77826.5774 | 93321.4236 | 77826.5774 | 77826.5774 | 18466.4887 | 77826.5774 | 18466.4887 | 77826.5774 | 18466.4887 | 18466.4887 | 18466.4887 | 18466.4887
PFAS CS1 0.01ng/ml_10/26/2023 20:30:40 | 17323.3791 | 17323.3791 | 17323.3791 | 77826.5774 | 77826.5774 | 77826.5774 | 17323.3791 | 17323.3791 | 87546.9311 | 77826.5774 | 77826.5774 | 87546.9311 | 77826.5774 | 77826.5774 | 17323.3791 | 77826.5774 | 17323.3791 | 77826.5774 | 17323.3791 | 17323.3791 | 17323.3791 | 17323.3791
PFAS CS1 0.01ng/ml_10/26/2023 20:47:16 | 15923.1668 | 15923.1668 | 15923.1668 | 77826.5774 | 77826.5774 | 77826.5774 | 15923.1668 | 15923.1668 | 84574.1553 | 77826.5774 | 77826.5774 | 84574.1553 | 77826.5774 | 77826.5774 | 15923.1668 | 77826.5774 | 15923.1668 | 77826.5774 | 15923.1668 | 15923.1668 | 15923.1668 | 15923.1668
PFAS CS1 0.01ng/ml_10/26/2023 21:03:55 | 18962.9111 | 18962.9111 | 18962.9111 | 77826.5774 | 77826.5774 | 77826.5774 | 18962.9111 | 18962.9111 | 90211.5141 | 77826.5774 | 77826.5774 | 90211.5141 | 77826.5774 | 77826.5774 | 18962.9111 | 77826.5774 | 18962.9111 | 77826.5774 | 18962.9111 | 18962.9111 | 18962.9111 | 18962.9111
PFAS CS10 400ng/ml_10/27/2023 06:11:54 | 11733.9724 | 11733.9724 | 11733.9724 | 37323.2205 | 37323.2205 | 37323.2205 | 11733.9724 | 11733.9724 | 72478.9105 | 37323.2205 | 37323.2205 | 72478.9105 | 37323.2205 | 37323.2205 | 11733.9724 | 37323.2205 | 11733.9724 | 37323.2205 | 11733.9724 | 11733.9724 | 11733.9724 | 11733.9724
PFAS CS10 400ng/ml_10/27/2023 06:28:34 | 11823.176 | 11823.176 | 11823.176 | 37323.2205 | 37323.2205 | 37323.2205 | 11823.176 | 11823.176 | 71130.4744 | 37323.2205 | 37323.2205 | 71130.4744 | 37323.2205 | 37323.2205 | 11823.176 | 37323.2205 | 11823.176 | 37323.2205 | 11823.176 | 11823.176 | 11823.176 | 11823.176
PFAS CS10 400ng/ml_10/27/2023 06:45:11 | 10962.3905 | 10962.3905 | 10962.3905 | 37323.2205 | 37323.2205 | 37323.2205 | 10962.3905 | 10962.3905 | 67921.1602 | 37323.2205 | 37323.2205 | 67921.1602 | 37323.2205 | 37323.2205 | 10962.3905 | 37323.2205 | 10962.3905 | 37323.2205 | 10962.3905 | 10962.3905 | 10962.3905 | 10962.3905
PFAS CS10 400ng/ml_10/27/2023 07:01:47 | 11437.6886 | 11437.6886 | 11437.6886 | 37323.2205 | 37323.2205 | 37323.2205 | 11437.6886 | 11437.6886 | 69603.4009 | 37323.2205 | 37323.2205 | 69603.4009 | 37323.2205 | 37323.2205 | 11437.6886 | 37323.2205 | 11437.6886 | 37323.2205 | 11437.6886 | 11437.6886 | 11437.6886 | 11437.6886
PFAS CS2 0.025ng/ml_10/26/2023 21:20:30 | 14799.5653 | 14799.5653 | 14799.5653 | 81409.8289 | 81409.8289 | 81409.8289 | 14799.5653 | 14799.5653 | 85962.9365 | 81409.8289 | 81409.8289 | 85962.9365 | 81409.8289 | 81409.8289 | 14799.5653 | 81409.8289 | 14799.5653 | 81409.8289 | 14799.5653 | 14799.5653 | 14799.5653 | 14799.5653
PFAS CS2 0.025ng/ml_10/26/2023 21:37:05 | 18895.0581 | 18895.0581 | 18895.0581 | 81409.8289 | 81409.8289 | 81409.8289 | 18895.0581 | 18895.0581 | 85815.9677 | 81409.8289 | 81409.8289 | 85815.9677 | 81409.8289 | 81409.8289 | 18895.0581 | 81409.8289 | 18895.0581 | 81409.8289 | 18895.0581 | 18895.0581 | 18895.0581 | 18895.0581
PFAS CS2 0.025ng/ml_10/26/2023 21:53:42 | 17906.663 | 17906.663 | 17906.663 | 81409.8289 | 81409.8289 | 81409.8289 | 17906.663 | 17906.663 | 89120.2677 | 81409.8289 | 81409.8289 | 89120.2677 | 81409.8289 | 81409.8289 | 17906.663 | 81409.8289 | 17906.663 | 81409.8289 | 17906.663 | 17906.663 | 17906.663 | 17906.663
PFAS CS2 0.025ng/ml_10/26/2023 22:10:16 | 16016.2371 | 16016.2371 | 16016.2371 | 81409.8289 | 81409.8289 | 81409.8289 | 16016.2371 | 16016.2371 | 84903.7754 | 81409.8289 | 81409.8289 | 84903.7754 | 81409.8289 | 81409.8289 | 16016.2371 | 81409.8289 | 16016.2371 | 81409.8289 | 16016.2371 | 16016.2371 | 16016.2371 | 16016.2371
PFAS CS3 0.1ng/ml_10/26/2023 22:26:52 | 19487.4063 | 19487.4063 | 19487.4063 | 81788.2473 | 81788.2473 | 81788.2473 | 19487.4063 | 19487.4063 | 94056.9934 | 81788.2473 | 81788.2473 | 94056.9934 | 81788.2473 | 81788.2473 | 19487.4063 | 81788.2473 | 19487.4063 | 81788.2473 | 19487.4063 | 19487.4063 | 19487.4063 | 19487.4063
PFAS CS3 0.1ng/ml_10/26/2023 22:43:27 | 18657.3031 | 18657.3031 | 18657.3031 | 81788.2473 | 81788.2473 | 81788.2473 | 18657.3031 | 18657.3031 | 88106.4173 | 81788.2473 | 81788.2473 | 88106.4173 | 81788.2473 | 81788.2473 | 18657.3031 | 81788.2473 | 18657.3031 | 81788.2473 | 18657.3031 | 18657.3031 | 18657.3031 | 18657.3031
PFAS CS3 0.1ng/ml_10/26/2023 23:00:04 | 18386.5477 | 18386.5477 | 18386.5477 | 81788.2473 | 81788.2473 | 81788.2473 | 18386.5477 | 18386.5477 | 88835.4907 | 81788.2473 | 81788.2473 | 88835.4907 | 81788.2473 | 81788.2473 | 18386.5477 | 81788.2473 | 18386.5477 | 81788.2473 | 18386.5477 | 18386.5477 | 18386.5477 | 18386.5477
PFAS CS3 0.1ng/ml_10/26/2023 23:16:42 | 18654.128 | 18654.128 | 18654.128 | 81788.2473 | 81788.2473 | 81788.2473 | 18654.128 | 18654.128 | 90523.0402 | 81788.2473 | 81788.2473 | 90523.0402 | 81788.2473 | 81788.2473 | 18654.128 | 81788.2473 | 18654.128 | 81788.2473 | 18654.128 | 18654.128 | 18654.128 | 18654.128
PFAS CS4 0.25ng/ml_10/26/2023 23:33:17 | 17652.038 | 17652.038 | 17652.038 | 82315.3517 | 82315.3517 | 82315.3517 | 17652.038 | 17652.038 | 86267.8949 | 82315.3517 | 82315.3517 | 86267.8949 | 82315.3517 | 82315.3517 | 17652.038 | 82315.3517 | 17652.038 | 82315.3517 | 17652.038 | 17652.038 | 17652.038 | 17652.038
PFAS CS4 0.25ng/ml_10/26/2023 23:49:51 | 18266.266 | 18266.266 | 18266.266 | 82315.3517 | 82315.3517 | 82315.3517 | 18266.266 | 18266.266 | 87554.3546 | 82315.3517 | 82315.3517 | 87554.3546 | 82315.3517 | 82315.3517 | 18266.266 | 82315.3517 | 18266.266 | 82315.3517 | 18266.266 | 18266.266 | 18266.266 | 18266.266
PFAS CS4 0.25ng/ml_10/27/2023 00:06:26 | 17937.6435 | 17937.6435 | 17937.6435 | 82315.3517 | 82315.3517 | 82315.3517 | 17937.6435 | 17937.6435 | 78089.0397 | 82315.3517 | 82315.3517 | 78089.0397 | 82315.3517 | 82315.3517 | 17937.6435 | 82315.3517 | 17937.6435 | 82315.3517 | 17937.6435 | 17937.6435 | 17937.6435 | 17937.6435
PFAS CS4 0.25ng/ml_10/27/2023 00:23:00 | 17861.0371 | 17861.0371 | 17861.0371 | 82315.3517 | 82315.3517 | 82315.3517 | 17861.0371 | 17861.0371 | 83394.0956 | 82315.3517 | 82315.3517 | 83394.0956 | 82315.3517 | 82315.3517 | 17861.0371 | 82315.3517 | 17861.0371 | 82315.3517 | 17861.0371 | 17861.0371 | 17861.0371 | 17861.0371
PFAS CS5 1ng/ml_10/27/2023 00:39:39 | 15609.9546 | 15609.9546 | 15609.9546 | 76300.1869 | 76300.1869 | 76300.1869 | 15609.9546 | 15609.9546 | 81422.1249 | 76300.1869 | 76300.1869 | 81422.1249 | 76300.1869 | 76300.1869 | 15609.9546 | 76300.1869 | 15609.9546 | 76300.1869 | 15609.9546 | 15609.9546 | 15609.9546 | 15609.9546
PFAS CS5 1ng/ml_10/27/2023 00:56:19 | 19746.58 | 19746.58 | 19746.58 | 76300.1869 | 76300.1869 | 76300.1869 | 19746.58 | 19746.58 | 84836.0623 | 76300.1869 | 76300.1869 | 84836.0623 | 76300.1869 | 76300.1869 | 19746.58 | 76300.1869 | 19746.58 | 76300.1869 | 19746.58 | 19746.58 | 19746.58 | 19746.58
PFAS CS5 1ng/ml_10/27/2023 01:12:53 | 17475.6051 | 17475.6051 | 17475.6051 | 76300.1869 | 76300.1869 | 76300.1869 | 17475.6051 | 17475.6051 | 82948.7905 | 76300.1869 | 76300.1869 | 82948.7905 | 76300.1869 | 76300.1869 | 17475.6051 | 76300.1869 | 17475.6051 | 76300.1869 | 17475.6051 | 17475.6051 | 17475.6051 | 17475.6051
PFAS CS5 1ng/ml_10/27/2023 01:29:29 | 16967.6418 | 16967.6418 | 16967.6418 | 76300.1869 | 76300.1869 | 76300.1869 | 16967.6418 | 16967.6418 | 84780.6746 | 76300.1869 | 76300.1869 | 84780.6746 | 76300.1869 | 76300.1869 | 16967.6418 | 76300.1869 | 16967.6418 | 76300.1869 | 16967.6418 | 16967.6418 | 16967.6418 | 16967.6418
PFAS CS6 4ng/ml_10/27/2023 01:46:07 | 14474.5682 | 14474.5682 | 14474.5682 | 70775.7789 | 70775.7789 | 70775.7789 | 14474.5682 | 14474.5682 | 78612.9775 | 70775.7789 | 70775.7789 | 78612.9775 | 70775.7789 | 70775.7789 | 14474.5682 | 70775.7789 | 14474.5682 | 70775.7789 | 14474.5682 | 14474.5682 | 14474.5682 | 14474.5682
PFAS CS6 4ng/ml_10/27/2023 02:02:41 | 15419.8788 | 15419.8788 | 15419.8788 | 70775.7789 | 70775.7789 | 70775.7789 | 15419.8788 | 15419.8788 | 78743.6783 | 70775.7789 | 70775.7789 | 78743.6783 | 70775.7789 | 70775.7789 | 15419.8788 | 70775.7789 | 15419.8788 | 70775.7789 | 15419.8788 | 15419.8788 | 15419.8788 | 15419.8788
PFAS CS6 4ng/ml_10/27/2023 02:19:16 | 14859.0958 | 14859.0958 | 14859.0958 | 70775.7789 | 70775.7789 | 70775.7789 | 14859.0958 | 14859.0958 | 81400.4542 | 70775.7789 | 70775.7789 | 81400.4542 | 70775.7789 | 70775.7789 | 14859.0958 | 70775.7789 | 14859.0958 | 70775.7789 | 14859.0958 | 14859.0958 | 14859.0958 | 14859.0958
PFAS CS6 4ng/ml_10/27/2023 02:35:53 | 16109.9168 | 16109.9168 | 16109.9168 | 70775.7789 | 70775.7789 | 70775.7789 | 16109.9168 | 16109.9168 | 77840.621 | 70775.7789 | 70775.7789 | 77840.621 | 70775.7789 | 70775.7789 | 16109.9168 | 70775.7789 | 16109.9168 | 70775.7789 | 16109.9168 | 16109.9168 | 16109.9168 | 16109.9168
PFAS CS7 10ng/ml_10/27/2023 02:52:28 | 15289.7545 | 15289.7545 | 15289.7545 | 64557.1987 | 64557.1987 | 64557.1987 | 15289.7545 | 15289.7545 | 75223.569 | 64557.1987 | 64557.1987 | 75223.569 | 64557.1987 | 64557.1987 | 15289.7545 | 64557.1987 | 15289.7545 | 64557.1987 | 15289.7545 | 15289.7545 | 15289.7545 | 15289.7545
PFAS CS7 10ng/ml_10/27/2023 03:09:05 | 14442.0219 | 14442.0219 | 14442.0219 | 64557.1987 | 64557.1987 | 64557.1987 | 14442.0219 | 14442.0219 | 80027.5995 | 64557.1987 | 64557.1987 | 80027.5995 | 64557.1987 | 64557.1987 | 14442.0219 | 64557.1987 | 14442.0219 | 64557.1987 | 14442.0219 | 14442.0219 | 14442.0219 | 14442.0219
PFAS CS7 10ng/ml_10/27/2023 03:25:43 | 15030.6068 | 15030.6068 | 15030.6068 | 64557.1987 | 64557.1987 | 64557.1987 | 15030.6068 | 15030.6068 | 76390.1725 | 64557.1987 | 64557.1987 | 76390.1725 | 64557.1987 | 64557.1987 | 15030.6068 | 64557.1987 | 15030.6068 | 64557.1987 | 15030.6068 | 15030.6068 | 15030.6068 | 15030.6068
PFAS CS7 10ng/ml_10/27/2023 03:42:18 | 16021.8401 | 16021.8401 | 16021.8401 | 64557.1987 | 64557.1987 | 64557.1987 | 16021.8401 | 16021.8401 | 78328.3846 | 64557.1987 | 64557.1987 | 78328.3846 | 64557.1987 | 64557.1987 | 16021.8401 | 64557.1987 | 16021.8401 | 64557.1987 | 16021.8401 | 16021.8401 | 16021.8401 | 16021.8401
PFAS CS8 40ng/ml_10/27/2023 03:58:55 | 14899.738 | 14899.738 | 14899.738 | 55812.9447 | 55812.9447 | 55812.9447 | 14899.738 | 14899.738 | 82145.7252 | 55812.9447 | 55812.9447 | 82145.7252 | 55812.9447 | 55812.9447 | 14899.738 | 55812.9447 | 14899.738 | 55812.9447 | 14899.738 | 14899.738 | 14899.738 | 14899.738
PFAS CS8 40ng/ml_10/27/2023 04:15:31 | 12699.2227 | 12699.2227 | 12699.2227 | 55812.9447 | 55812.9447 | 55812.9447 | 12699.2227 | 12699.2227 | 83963.8973 | 55812.9447 | 55812.9447 | 83963.8973 | 55812.9447 | 55812.9447 | 12699.2227 | 55812.9447 | 12699.2227 | 55812.9447 | 12699.2227 | 12699.2227 | 12699.2227 | 12699.2227
PFAS CS8 40ng/ml_10/27/2023 04:32:10 | 14554.7068 | 14554.7068 | 14554.7068 | 55812.9447 | 55812.9447 | 55812.9447 | 14554.7068 | 14554.7068 | 90422.0791 | 55812.9447 | 55812.9447 | 90422.0791 | 55812.9447 | 55812.9447 | 14554.7068 | 55812.9447 | 14554.7068 | 55812.9447 | 14554.7068 | 14554.7068 | 14554.7068 | 14554.7068
PFAS CS8 40ng/ml_10/27/2023 04:48:50 | 13850.6181 | 13850.6181 | 13850.6181 | 55812.9447 | 55812.9447 | 55812.9447 | 13850.6181 | 13850.6181 | 82988.7309 | 55812.9447 | 55812.9447 | 82988.7309 | 55812.9447 | 55812.9447 | 13850.6181 | 55812.9447 | 13850.6181 | 55812.9447 | 13850.6181 | 13850.6181 | 13850.6181 | 13850.6181
PFAS CS9 100ng/ml_10/27/2023 05:05:26 | 6062.4002 | 6062.4002 | 6062.4002 | 19415.8978 | 19415.8978 | 19415.8978 | 6062.4002 | 6062.4002 | 45565.4144 | 19415.8978 | 19415.8978 | 45565.4144 | 19415.8978 | 19415.8978 | 6062.4002 | 19415.8978 | 6062.4002 | 19415.8978 | 6062.4002 | 6062.4002 | 6062.4002 | 6062.4002
PFAS CS9 100ng/ml_10/27/2023 05:22:02 | 5185.5759 | 5185.5759 | 5185.5759 | 19415.8978 | 19415.8978 | 19415.8978 | 5185.5759 | 5185.5759 | 47570.5579 | 19415.8978 | 19415.8978 | 47570.5579 | 19415.8978 | 19415.8978 | 5185.5759 | 19415.8978 | 5185.5759 | 19415.8978 | 5185.5759 | 5185.5759 | 5185.5759 | 5185.5759
PFAS CS9 100ng/ml_10/27/2023 05:38:39 | 7119.4913 | 7119.4913 | 7119.4913 | 19415.8978 | 19415.8978 | 19415.8978 | 7119.4913 | 7119.4913 | 48835.2844 | 19415.8978 | 19415.8978 | 48835.2844 | 19415.8978 | 19415.8978 | 7119.4913 | 19415.8978 | 7119.4913 | 19415.8978 | 7119.4913 | 7119.4913 | 7119.4913 | 7119.4913
PFAS CS9 100ng/ml_10/27/2023 05:55:18 | 7670.2264 | 7670.2264 | 7670.2264 | 19415.8978 | 19415.8978 | 19415.8978 | 7670.2264 | 7670.2264 | 49832.4046 | 19415.8978 | 19415.8978 | 49832.4046 | 19415.8978 | 19415.8978 | 7670.2264 | 19415.8978 | 7670.2264 | 19415.8978 | 7670.2264 | 7670.2264 | 7670.2264 | 7670.2264
Average | 15401.0423 | 15401.0423 | 15401.0423 | 66080.8801 | 66080.8801 | 66080.8801 | 15401.0423 | 15401.0423 | 79943.4778 | 66080.8801 | 66080.8801 | 79943.4778 | 66080.8801 | 66080.8801 | 15401.0423 | 66080.8801 | 15401.0423 | 66080.8801 | 15401.0423 | 15401.0423 | 15401.0423 | 15401.0423

## Sample IDA_IPS Ratio
Sample Name Date | IDA-13C2 4:2 FTS | IDA-13C2 6:2 FTS | IDA-13C2 8:2 FTS | IDA-13C2_PFDoA | IDA-13C2_PFTeDA | IDA-13C3_HFPO-DA | IDA-13C3_PFBS | IDA-13C3_PFHxS | IDA-13C4_PFBA | IDA-13C4_PFHpA | IDA-13C5_PFHxA | IDA-13C5_PFPeA | IDA-13C6_PFDA | IDA-13C7_PFUdA | IDA-13C8_FOSA | IDA-13C8_PFOA | IDA-13C8_PFOS | IDA-13C9_PFNA | IDA-d-EtFOSA | IDA-d-MeFOSA | IDA-d3-MeFOSAA | IDA-d5-EtFOSAA
PFAS QCS0 0ng/ml_09/17/2024 15:14:43 | 0.74837 | 0.58035 | 1.2 | 1.3703 | 0.22329 | 1.6979 | 2.4134 | 3.6519 | 1.1602 | 2.4279 | 2.0342 | 0.90085 | 2.1408 | 1.5672 | 4.3373 | 2.3015 | 2.1335 | 1.949 | 0.38808 | 0.48721 | 3.2065 | 1.0244
PFAS QCS0 0ng/ml_09/18/2024 15:10:07 | 0.84765 | 0.66838 | 1.1706 | 0.84262 | 0.20539 | 1.43 | 2.555 | 3.6288 | 1.2762 | 2.127 | 1.6852 | 0.90773 | 1.5173 | 1.1764 | 4.0987 | 1.8609 | 2.184 | 1.8481 | 0.8228 | 1.1854 | 3.1399 | 1.0804
PFBA_72hr_Dup_09/18/2024 13:11:46 | 0.98198 | 1.2704 | 2.4497 | 1.3686 | 0.31908 | 1.9504 | 3.0243 | 4.7292 | 1.3456 | 2.959 | 2.2187 | 0.9603 | 2.2376 | 1.499 | 5.284 | 2.6023 | 2.6791 | 2.2902 | 2.3412 | 2.2299 | 5.0079 | 1.422
PFBS_72hr_Dup_09/18/2024 13:45:46 | 1.5948 | 1.368 | 2.616 | 2.8849 | 0.71404 | 3.9459 | 4.8791 | 7.3404 | 2.4773 | 5.3835 | 4.3394 | 1.7136 | 4.7375 | 3.5025 | 9.2567 | 4.9859 | 4.325 | 4.3781 | 4.1927 | 3.867 | 7.6275 | 2.415
PFOA_72hr_09/18/2024 13:28:22 | 0.7132 | 0.59321 | 1.3924 | 1.2293 | 0.32865 | 1.659 | 2.2558 | 3.3504 | 1.1795 | 2.3804 | 1.877 | 0.82226 | 1.9519 | 1.5659 | 4.1342 | 2.0915 | 1.8549 | 1.8162 | 1.7857 | 1.6678 | 3.7187 | 1.0946
PFOS_72hr_Dup_09/18/2024 14:02:27 | 1.7395 | 1.4241 | 2.4966 | 3.0727 | 0.7233 | 3.8262 | 5.0033 | 7.7633 | 2.5718 | 5.6445 | 4.6167 | 1.8071 | 4.369 | 3.3806 | 9.1965 | 5.0306 | 4.3973 | 4.3726 | 4.2023 | 4.0412 | 7.3598 | 2.2787
Standard_check_327ai_343g_09/18/2024 14:53:30 | 0.52923 | 0.44704 | 0.75659 | 1.0747 | 0.24444 | 1.4423 | 1.7337 | 2.6212 | 0.90735 | 1.8364 | 1.5829 | 0.64383 | 1.5092 | 1.0724 | 3.122 | 1.7788 | 1.5687 | 1.5566 | 1.3827 | 1.3861 | 1.939 | 0.63984
Standard_check_327aj_343g_09/18/2024 14:19:06 | 0.8021 | 0.67238 | 1.1613 | 1.1495 | 0.31962 | 1.6578 | 2.4689 | 3.5604 | 1.1783 | 2.4125 | 1.9086 | 0.84508 | 1.9322 | 1.6017 | 4.6267 | 2.1616 | 2.1079 | 1.7382 | 2.1937 | 2.111 | 3.514 | 1.0822
Standard_check_327aj_343o_09/18/2024 14:36:12 | 1.5315 | 1.2945 | 2.1164 | 2.7665 | 0.77032 | 3.8992 | 4.6721 | 7.25 | 2.3657 | 5.2648 | 4.5021 | 1.7231 | 3.6577 | 3.2933 | 8.6278 | 4.9928 | 4.1677 | 4.3218 | 3.9414 | 3.8124 | 6.7884 | 1.9522

## Sample IDA
Sample Name Date | IDA-13C2 4:2 FTS | IDA-13C2 6:2 FTS | IDA-13C2 8:2 FTS | IDA-13C2_PFDoA | IDA-13C2_PFTeDA | IDA-13C3_HFPO-DA | IDA-13C3_PFBS | IDA-13C3_PFHxS | IDA-13C4_PFBA | IDA-13C4_PFHpA | IDA-13C5_PFHxA | IDA-13C5_PFPeA | IDA-13C6_PFDA | IDA-13C7_PFUdA | IDA-13C8_FOSA | IDA-13C8_PFOA | IDA-13C8_PFOS | IDA-13C9_PFNA | IDA-d-EtFOSA | IDA-d-MeFOSA | IDA-d3-MeFOSAA | IDA-d5-EtFOSAA
PFAS QCS0 0ng/ml_09/17/2024 15:14:43 | 8991.4696 | 6972.7442 | 14417.6725 | 75926.7521 | 12371.8442 | 94077.2083 | 28995.8935 | 43876.2625 | 65593.8874 | 134525.1585 | 112709.1563 | 50932.2893 | 118617.1563 | 86833.6588 | 52111.9193 | 127520.6872 | 25633.8543 | 107988.0998 | 4662.7147 | 5853.6637 | 38525.2159 | 12307.8294
PFAS QCS0 0ng/ml_09/18/2024 15:10:07 | 9056.2475 | 7140.9448 | 12507.012 | 46688.0317 | 11380.0075 | 79234.4809 | 27297.6296 | 38769.7915 | 67295.7323 | 117855.1481 | 93375.5406 | 47865.9542 | 84069.7071 | 65184.603 | 43790.8638 | 103109.2284 | 23334.3751 | 102401.2248 | 8790.7932 | 12664.6601 | 33547.081 | 11542.7152
PFBA_72hr_Dup_09/18/2024 13:11:46 | 24191.3696 | 31297.4251 | 60348.8967 | 179922.3171 | 41948.4212 | 256411.4452 | 74504.8133 | 116505.4415 | 184192.4946 | 389012.9325 | 291676.7102 | 131449.6908 | 294167.0604 | 197064.8435 | 130171.8486 | 342107.5903 | 66000.9489 | 301085.2875 | 57677.0203 | 54933.9465 | 123371.1496 | 35031.7724
PFBS_72hr_Dup_09/18/2024 13:45:46 | 23986.8673 | 20576.0479 | 39346.4312 | 192663.2304 | 47686.1119 | 263525.5152 | 73383.9531 | 110404.2796 | 190291.2651 | 359533.2758 | 289804.3881 | 131629.0573 | 316387.7007 | 233907.4305 | 139225.8382 | 332974.662 | 65050.8763 | 292383.6446 | 63059.7915 | 58161.2579 | 114722.4972 | 36322.4926
PFOA_72hr_09/18/2024 13:28:22 | 24118.6311 | 20060.7372 | 47085.9556 | 185584.6422 | 49613.6845 | 250443.5546 | 76285.3262 | 113302.2903 | 191398.6091 | 359349.3929 | 283354.8026 | 133431.0283 | 294659.8017 | 236398.5245 | 139807.3418 | 315741.9548 | 62726.8312 | 274173.019 | 60388.8928 | 56400.6112 | 125758.8326 | 37017.0869
PFOS_72hr_Dup_09/18/2024 14:02:27 | 26296.0283 | 21527.4416 | 37740.2252 | 203446.252 | 47889.5837 | 253335.0414 | 75633.6899 | 117356.6527 | 193931.9654 | 373720.7092 | 305670.1793 | 136267.2048 | 289270.1394 | 223832.7351 | 139021.7874 | 333074.495 | 66473.6698 | 289513.2916 | 63525.8646 | 61090.4791 | 111256.7895 | 34446.8206
Standard_check_327ai_343g_09/18/2024 14:53:30 | 14058.2651 | 11875.1635 | 20097.9309 | 118684.1606 | 26993.5225 | 159272.1393 | 46052.364 | 69629.894 | 112485.9828 | 202798.6845 | 174805.5172 | 79816.9997 | 166658.0595 | 118431.4068 | 82933.1302 | 196439.1022 | 41670.1791 | 171896.0395 | 36730.5346 | 36820.2409 | 51507.8838 | 16996.4831
Standard_check_327aj_343g_09/18/2024 14:19:06 | 26066.4652 | 21850.7948 | 37738.4094 | 185222.3506 | 51500.4955 | 267118.1345 | 80235.3109 | 115704.3408 | 200040.6296 | 388734.1287 | 307540.9136 | 143471.9241 | 311341.3659 | 258079.8021 | 150356.0489 | 348306.8339 | 68502.9746 | 280077.1934 | 71290.1481 | 68604.0117 | 114197.9173 | 35167.9045
Standard_check_327aj_343o_09/18/2024 14:36:12 | 25190.9575 | 21293.1337 | 34811.7153 | 186965.2812 | 52059.5634 | 263516.5682 | 76849.1501 | 119252.7419 | 190110.8626 | 355808.8161 | 304262.3628 | 138472.416 | 247192.8422 | 222568.413 | 141915.0531 | 337424.2436 | 68553.1746 | 292074.7418 | 64831.0752 | 62709.058 | 111660.1884 | 32110.7785

## Sample IPS
Sample Name Date | IDA-13C2 4:2 FTS | IDA-13C2 6:2 FTS | IDA-13C2 8:2 FTS | IDA-13C2_PFDoA | IDA-13C2_PFTeDA | IDA-13C3_HFPO-DA | IDA-13C3_PFBS | IDA-13C3_PFHxS | IDA-13C4_PFBA | IDA-13C4_PFHpA | IDA-13C5_PFHxA | IDA-13C5_PFPeA | IDA-13C6_PFDA | IDA-13C7_PFUdA | IDA-13C8_FOSA | IDA-13C8_PFOA | IDA-13C8_PFOS | IDA-13C9_PFNA | IDA-d-EtFOSA | IDA-d-MeFOSA | IDA-d3-MeFOSAA | IDA-d5-EtFOSAA
PFAS QCS0 0ng/ml_09/17/2024 15:14:43 | 12014.7492 | 12014.7492 | 12014.7492 | 55408.0463 | 55408.0463 | 55408.0463 | 12014.7492 | 12014.7492 | 56538.1016 | 55408.0463 | 55408.0463 | 56538.1016 | 55408.0463 | 55408.0463 | 12014.7492 | 55408.0463 | 12014.7492 | 55408.0463 | 12014.7492 | 12014.7492 | 12014.7492 | 12014.7492
PFAS QCS0 0ng/ml_09/18/2024 15:10:07 | 10684.0019 | 10684.0019 | 10684.0019 | 55408.0463 | 55408.0463 | 55408.0463 | 10684.0019 | 10684.0019 | 52731.2687 | 55408.0463 | 55408.0463 | 52731.2687 | 55408.0463 | 55408.0463 | 10684.0019 | 55408.0463 | 10684.0019 | 55408.0463 | 10684.0019 | 10684.0019 | 10684.0019 | 10684.0019
PFBA_72hr_Dup_09/18/2024 13:11:46 | 24635.2135 | 24635.2135 | 24635.2135 | 131465.6548 | 131465.6548 | 131465.6548 | 24635.2135 | 24635.2135 | 136884.2811 | 131465.6548 | 131465.6548 | 136884.2811 | 131465.6548 | 131465.6548 | 24635.2135 | 131465.6548 | 24635.2135 | 131465.6548 | 24635.2135 | 24635.2135 | 24635.2135 | 24635.2135
PFBS_72hr_Dup_09/18/2024 13:45:46 | 15040.5568 | 15040.5568 | 15040.5568 | 66783.8705 | 66783.8705 | 66783.8705 | 15040.5568 | 15040.5568 | 76812.9235 | 66783.8705 | 66783.8705 | 76812.9235 | 66783.8705 | 66783.8705 | 15040.5568 | 66783.8705 | 15040.5568 | 66783.8705 | 15040.5568 | 15040.5568 | 15040.5568 | 15040.5568
PFOA_72hr_09/18/2024 13:28:22 | 33817.5052 | 33817.5052 | 33817.5052 | 150963.2512 | 150963.2512 | 150963.2512 | 33817.5052 | 33817.5052 | 162274.2544 | 150963.2512 | 150963.2512 | 162274.2544 | 150963.2512 | 150963.2512 | 33817.5052 | 150963.2512 | 33817.5052 | 150963.2512 | 33817.5052 | 33817.5052 | 33817.5052 | 33817.5052
PFOS_72hr_Dup_09/18/2024 14:02:27 | 15116.8134 | 15116.8134 | 15116.8134 | 66210.2953 | 66210.2953 | 66210.2953 | 15116.8134 | 15116.8134 | 75406.6811 | 66210.2953 | 66210.2953 | 75406.6811 | 66210.2953 | 66210.2953 | 15116.8134 | 66210.2953 | 15116.8134 | 66210.2953 | 15116.8134 | 15116.8134 | 15116.8134 | 15116.8134
Standard_check_327ai_343g_09/18/2024 14:53:30 | 26563.7717 | 26563.7717 | 26563.7717 | 110431.5446 | 110431.5446 | 110431.5446 | 26563.7717 | 26563.7717 | 123971.4405 | 110431.5446 | 110431.5446 | 123971.4405 | 110431.5446 | 110431.5446 | 26563.7717 | 110431.5446 | 26563.7717 | 110431.5446 | 26563.7717 | 26563.7717 | 26563.7717 | 26563.7717
Standard_check_327aj_343g_09/18/2024 14:19:06 | 32497.7859 | 32497.7859 | 32497.7859 | 161130.8118 | 161130.8118 | 161130.8118 | 32497.7859 | 32497.7859 | 169773.9108 | 161130.8118 | 161130.8118 | 169773.9108 | 161130.8118 | 161130.8118 | 32497.7859 | 161130.8118 | 32497.7859 | 161130.8118 | 32497.7859 | 32497.7859 | 32497.7859 | 32497.7859
Standard_check_327aj_343o_09/18/2024 14:36:12 | 16448.5881 | 16448.5881 | 16448.5881 | 67581.956 | 67581.956 | 67581.956 | 16448.5881 | 16448.5881 | 80362.728 | 67581.956 | 67581.956 | 80362.728 | 67581.956 | 67581.956 | 16448.5881 | 67581.956 | 16448.5881 | 67581.956 | 16448.5881 | 16448.5881 | 16448.5881 | 16448.5881

## IDL_IQL
Sample Name | Sample ID | Sample Type | PFBA | PFPeA | PFHxA | PFHpA | PFOA | PFNA | PFDA | PFUdA | PFDoA | PFTrDA | PFTeDA | 5:3 FTCA | 7:3 FTCA | 9:3 FTCA | PFPrSA | FPeSA | PFEA | PFPrA | PFBS | PFPeS | PFHxS | PFHxS_branched | PFHpS | PFOS | PFOS_branched | PFNS | PFDS | MeFBSA | FBSA | FHxSA | FOSA | N-MeFOSAA | N-MeFOSAA_branched | N-EtFOSAA | N-EtFOSAA_branched | MeFOSA | EtFOSA | HFPO-DA | DONA | 9ClPF3ONS | 11ClPF3OUdS | 3,6-OPFHpA | PF4OPeA | PF5OHxA | PFECHS | PFEESA | 4:2 FTS | 6:2 FTS | 8:2 FTS | 10:2 FTS
IDL |  |  | 0.0099224 | 0.0068319 | 0.0010053 | 0.00055916 | 0.0017253 | 0.0014342 | 0.0011607 | 0.0017707 | 0.00055381 | 0.00017099 | 0.0069295 |  | 0.0013457 | [5 empty] | 0.004992 | 0.0019368 | 0.0045482 | 0.023386 | 0.0033958 | 0.0073769 | 0.014259 | 0.0068479 | 0.00755 | 0.0060309 | 0.0092473 | 0.0003902 | 0.001377 | 0.0068933 | 0.11162 | 0.0041815 | 0.12508 | 0.0023212 | 0.0054201 | 0.0027388 | 0.00044821 | 6.68e-05 | 0.0010157 | 0.0014049 | 0.0029336 | 0.0044046 | 7.43e-05 | 0.005066 | 0.012028 | 0.0002046 | 0.00058753
IQL |  |  | 0.029767 | 0.020496 | 0.0030158 | 0.0016775 | 0.005176 | 0.0043025 | 0.003482 | 0.0053122 | 0.0016614 | 0.00051296 | 0.020789 |  | 0.0040371 | [5 empty] | 0.014976 | 0.0058103 | 0.013645 | 0.070159 | 0.010187 | 0.022131 | 0.042776 | 0.020544 | 0.02265 | 0.018093 | 0.027742 | 0.0011706 | 0.0041309 | 0.02068 | 0.33487 | 0.012544 | 0.37524 | 0.0069635 | 0.01626 | 0.0082165 | 0.0013446 | 0.00020026 | 0.0030472 | 0.0042148 | 0.0088009 | 0.013214 | 0.00022287 | 0.015198 | 0.036083 | 0.00061379 | 0.0017626

## Calibration Data
| Index | Sample Name | Sample Index | Comparison | Original Filename | Sample ID | Barcode | Scanned Barcode | Sample Type | Sample Comment | Rack Number | Plate Number | Vial Number | Dilution Factor | Component Name | Calculated Concentration | Component Type | Acquisition Date & Time | Injection Volume | Operator Name | Acq. Method Name | IS | Component Index | Component Comment | IS Comment | Mass Info | IS Mass Info | IS Name | Component Group Name | Conc. Units | Outlier Reasons | Peak Comment | IS Peak Comment | Actual Concentration | IS Actual Concentration | Concentration Ratio | Reported Recovery | Expected RT | IS Expected RT | Integration Type | IS Integration Type | Area | IS Area | Area Ratio | Height | IS Height | Height Ratio | Area / Height | IS Area / Height | Region Height | IS Region Height | Quality | IS Quality | Retention Time | IS Retention Time | Retention Time Error (%) | Retention Time Delta (min) | Start Time | IS Start Time | End Time | IS End Time | Total Width | IS Total Width | Width at 50% | IS Width at 50% | Signal / Noise | IS Signal / Noise | Baseline Delta / Height | IS Baseline Delta / Height | Modified | Relative RT | Used | Adduct / Charge | Accuracy | Start Time at 5% | End Time at 5% | Width at 5% | Start Time at 10% | End Time at 10% | Width at 10% | Slope of Baseline | Tailing Factor | Asymmetry Factor | Ion Ratio | Expected Ion Ratio | Points Across Baseline | Points Across Half Height | Formula | Precursor Mass | Fragment Mass | Ion Ratio Confidence | Accuracy Acceptance | Concentration Acceptance | Integration Acceptance | XIC Width (Da) | XIC Width (ppm) | MS method | LC method | Std Addition Actual Concentration | Std Addition Calculated Concentration | Std Addition Accuracy | Reportable | Proc. Method Name | Area IDA | Area Ratio of comparison | Expected IDA Area | Mass Error Confidence | Fragment Mass Error Confidence | Area IPS | Expected IPS Area | RT Confidence | IDA Average Response Factor | Isotope Confidence | Calculated IDA Conc. | Library Confidence | Calculated IPS Concentration | Formula Confidence | Found At Mass | Mass Error (ppm) | Mass Error (mDa) | Found At Fragment Mass | Fragment Mass Error (ppm) | Fragment Mass Error (mDa) | Library Hit | Library Score | Combined Score | Formula Finder Results | Formula Finder Score | Isotope Ratio Difference | Non-Targeted Peak | Concentration/IS Concentration | Area/IS Area
1 | 2 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFBA | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 1.49 | 1.48 | 0 | Manual | 0 | 97153.4926 | 0 | 0 | 13855.9087 | 0 | 0 | 7.0117 | 74.3854 | 13898.6643 | 0 | 0.98632 | 0 | 1.5624 | 0 | 0 | 0 | 1.4363 | 0 | 1.8435 | 0 | 0.40725 | 0 | 0.10556 | 0 | 185.7471 | 0 | 0.00065858 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
3 | 4 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFPeA | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 3.18 | 3.21 | 0 | Manual | 0 | 76418.6589 | 0 | 0 | 7392.338 | 0 | 0 | 10.3375 | 58.8951 | 7409.7483 | 0 | 0.98473 | 0 | 3.4753 | 0 | 0 | 0 | 3.2638 | 0 | 3.7259 | 0 | 0.46209 | 0 | 0.15895 | 0 | 581.4486 | 0 | 0.005037 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
5 | 6 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFHxA | 0.0104951410255431 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 6.13 | 6.12 | Manual | Baseline | 386.966 | 158429.6564 | 0.0024425 | 71.8097 | 30080.8569 | 0.0023872 | 5.3888 | 5.2668 | 80.0828 | 30102.9608 | 0.4001 | 0.99482 | 6.306 | 6.2991 | 2.8719 | 0.17605 | 6.2207 | 6.1347 | 6.3934 | 6.4996 | 0.17267 | 0.36498 | 0.10099 | 0.081033 | 12.8334 | 317.2159 | 0.0056041 | 0.0012875 | TRUE | 1.0011 | TRUE | [M-H]- | 0 | 6.2291 | 6.4017 | 0.17261 | 6.2323 | 6.3864 | 0.15409 | -2.3305 | 1.2715 | 1.3823 | 1.5169 | 0 | 20 | 13 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 386.966 | No Control Sample | 203843.5852 | off | off | 386.966 | 203843.5852 | off | 0 | off | 0 | off | 0 | off | 312.9743 | 4.7931 | 1.5001 | 268.9856 | 11.7533 | 3.1614 | 0 | 0 | 0 | 0 | 0 | 7439.5121 | FALSE | 0 | 0.0024425
7 | 8 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFHpA | 0.0147716459096467 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 7.26 | 7.27 | Baseline | Baseline | 661.0323 | 178840.2452 | 0.0036962 | 197.5691 | 43560.136 | 0.0045355 | 3.3458 | 4.1056 | 220.0898 | 43588.7993 | 0.33265 | 0.99521 | 7.376 | 7.3805 | 1.5975 | 0.11598 | 7.2698 | 7.2462 | 7.4332 | 7.5221 | 0.16338 | 0.27592 | 0.04618 | 0.063456 | 22.3361 | 288.345 | 0.063648 | 0.0012973 | FALSE | 0.99939 | TRUE | [M-H]- | 0 | 7.2946 | 7.427 | 0.13244 | 7.3001 | 7.4198 | 0.11962 | 76.9656 | 0.83767 | 0.62727 | 0.75327 | 0 | 20 | 8 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 661.0323 | No Control Sample | 228328.7703 | off | off | 661.0323 | 228328.7703 | off | 0 | off | 0 | off | 0 | off | 362.972 | 6.6936 | 2.4296 | 318.9807 | 5.1703 | 1.6492 | 0 | 0 | 0 | 0 | 0 | 37.814 | FALSE | 0 | 0.0036962
9 | 10 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFOA | 0.0106673075879011 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 7.96 | 7.94 | Baseline | Baseline | 436.7084 | 168806.0426 | 0.002587 | 112.7318 | 37040.6599 | 0.0030435 | 3.8739 | 4.5573 | 120.0913 | 37047.7626 | 0.40725 | 0.99773 | 8.0415 | 8.0511 | 1.0241 | 0.081521 | 7.9781 | 7.901 | 8.1336 | 8.2691 | 0.15557 | 0.36815 | 0.062571 | 0.071654 | 14.8511 | 144.2814 | 0.087933 | 0.0004868 | FALSE | 0.99881 | TRUE | [M-H]- | 0 | 7.9838 | 8.105 | 0.12115 | 7.9919 | 8.0997 | 0.10786 | -63.7212 | 1.1611 | 1.4442 | 1.179 | 0 | 11 | 6 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 436.7084 | No Control Sample | 191475.4056 | off | off | 436.7084 | 191475.4056 | off | 0 | off | 0 | off | 0 | off | 412.9682 | 4.2613 | 1.7598 | 368.9811 | 13.6073 | 5.0208 | 0 | 0 | 0 | 0 | 0 | 28.5686 | FALSE | 0 | 0.002587
11 | 12 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFNA | 0.00555615752474431 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 8.46 | 8.47 | Manual | Baseline | 196.0126 | 151860.9441 | 0.0012907 | 33.0685 | 34114.0584 | 0.00096935 | 5.9275 | 4.4516 | 35.869 | 34130.7826 | 0.55993 | 0.99616 | 8.5517 | 8.5507 | 1.0839 | 0.091701 | 8.4002 | 8.3557 | 8.6331 | 8.7062 | 0.23289 | 0.35048 | 0.067332 | 0.068533 | 12.4128 | 99.4298 | 0.021803 | 0.00082108 | TRUE | 1.0001 | TRUE | [M-H]- | 0 | 8.4007 | 8.6301 | 0.22946 | 8.409 | 8.6236 | 0.21463 | -3.0958 | 0.8027 | 0.59436 | 0 | 0 | 15 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Fail | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 196.0126 | No Control Sample | 159143.3885 | off | off | 196.0126 | 159143.3885 | off | 0 | off | 0 | off | 0 | off | 462.9682 | 10.6452 | 4.9283 | 418.9694 | -9.0966 | -3.8112 | 0 | 0 | 0 | 0 | 0 | 9.8103 | FALSE | 0 | 0.0012907
13 | 14 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFDA | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.87 | 8.86 | 0 | Baseline | 0 | 115798.6291 | 0 | 0 | 24113.0181 | 0 | 0 | 4.8023 | 48.6873 | 24127.6278 | 0 | 0.99573 | 0 | 8.9584 | 0 | 0 | 0 | 8.7664 | 0 | 9.1168 | 0 | 0.35042 | 0 | 0.076208 | 0 | 174.7292 | 0 | 0.0010356 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
15 | 16 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFUdA | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.21 | 9.2 | 0 | Baseline | 0 | 100773.8028 | 0 | 0 | 22240.5517 | 0 | 0 | 4.5311 | 34.1201 | 22248.4676 | 0 | 0.99694 | 0 | 9.2936 | 0 | 0 | 0 | 9.1378 | 0 | 9.4829 | 0 | 0.34517 | 0 | 0.071399 | 0 | 147.1219 | 0 | 0.00076552 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
17 | 18 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFDoA | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.47 | 9.47 | 0 | Baseline | 0 | 87765.045 | 0 | 0 | 21826.6455 | 0 | 0 | 4.021 | 75.043 | 21828.2513 | 0 | 0.99854 | 0 | 9.5555 | 0 | 0 | 0 | 9.418 | 0 | 9.7222 | 0 | 0.3042 | 0 | 0.063084 | 0 | 222.3043 | 0 | 0.00015594 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
19 | 20 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFTrDA | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.68 | 9.47 | 0 | Baseline | 0 | 87765.045 | 0 | 0 | 21826.6455 | 0 | 0 | 4.021 | 68.8024 | 21828.2513 | 0 | 0.99854 | 0 | 9.5555 | 0 | 0 | 0 | 9.418 | 0 | 9.7222 | 0 | 0.3042 | 0 | 0.063084 | 0 | 222.3043 | 0 | 0.00015594 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
21 | 22 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFTeDA | 0.0496232870006124 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 9.87 | 9.86 | Baseline | Baseline | 338.0174 | 19212.3256 | 0.017594 | 100.0858 | 6019.2182 | 0.016628 | 3.3773 | 3.1918 | 102.1622 | 6032.0473 | 0.43565 | 0.98966 | 9.9385 | 9.9442 | 0.69408 | 0.068506 | 9.8849 | 9.8472 | 10.0348 | 10.0223 | 0.14988 | 0.1751 | 0.051595 | 0.050145 | 31.9099 | 764.578 | 0.057225 | 0.0034497 | FALSE | 0.99942 | TRUE | [M-H]- | 0 | 9.8984 | 10.0249 | 0.12657 | 9.9023 | 10.0123 | 0.10992 | 38.2122 | 1.7673 | 2.4515 | 0 | 0 | 19 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 338.0174 | No Control Sample | 0 | off | off | 338.0174 | 0 | off | 0 | off | 0 | off | 0 | off | 713.1848 | 333.1156 | 237.4938 | 668.9633 | 8.0081 | 5.3571 | 0 | 0 | 0 | 0 | 0 | inf | FALSE | 0 | 0.017594
28 | 29 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFBS | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 3.5 | 3.5 | 0 | Manual | 0 | 42471.3802 | 0 | 0 | 3848.7751 | 0 | 0 | 11.035 | 28.776 | 3868.2133 | 0 | 1 | 0 | 4.1837 | 0 | 0 | 0 | 3.9608 | 0 | 4.426 | 0 | 0.46522 | 0 | 0.16111 | 0 | 634.0644 | 0 | 0 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
30 | 31 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFPeS | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 6.38 | 3.5 | 0 | Manual | 0 | 42471.3802 | 0 | 0 | 3848.7751 | 0 | 0 | 11.035 | 13.9894 | 3868.2133 | 0 | 1 | 0 | 4.1837 | 0 | 0 | 0 | 3.9608 | 0 | 4.426 | 0 | 0.46522 | 0 | 0.16111 | 0 | 634.0644 | 0 | 0 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
32 | 33 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | L-PFHxS | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.34 | 7.35 | 0 | Baseline | 0 | 61500.3109 | 0 | 0 | 15584.2995 | 0 | 0 | 3.9463 | 35.412 | 15591.6215 | 0 | 0.99235 | 0 | 7.4484 | 0 | 0 | 0 | 7.3346 | 0 | 7.7068 | 0 | 0.37218 | 0 | 0.058965 | 0 | 210.4971 | 0 | 0.0011916 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
34 | 35 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | Br-PFHxS | 0 | Qualifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.16 | 7.35 | 0 | Baseline | 0 | 61500.3109 | 0 | 0 | 15584.2995 | 0 | 0 | 3.9463 | 35.412 | 15591.6215 | 0 | 0.99235 | 0 | 7.4484 | 0 | 0 | 0 | 7.3346 | 0 | 7.7068 | 0 | 0.37218 | 0 | 0.058965 | 0 | 210.4971 | 0 | 0.0011916 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
36 | 37 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFHpS | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.98 | 7.35 | 0 | Baseline | 0 | 61500.3109 | 0 | 0 | 15584.2995 | 0 | 0 | 3.9463 | 8.0675 | 15591.6215 | 0 | 0.99235 | 0 | 7.4484 | 0 | 0 | 0 | 7.3346 | 0 | 7.7068 | 0 | 0.37218 | 0 | 0.058965 | 0 | 210.4971 | 0 | 0.0011916 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
38 | 39 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | L-PFOS | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.46 | 8.45 | 0 | Baseline | 0 | 38652.0472 | 0 | 0 | 8575.2047 | 0 | 0 | 4.5074 | 25.6912 | 8585.2246 | 0 | 0.99192 | 0 | 8.5479 | 0 | 0 | 0 | 8.3962 | 0 | 8.6909 | 0 | 0.29475 | 0 | 0.069577 | 0 | 84.4406 | 0 | 0.0021327 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
40 | 41 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | Br-PFOS | 0 | Qualifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.3 | 8.45 | 0 | Baseline | 0 | 38652.0472 | 0 | 0 | 8575.2047 | 0 | 0 | 4.5074 | 25.6912 | 8585.2246 | 0 | 0.99192 | 0 | 8.5479 | 0 | 0 | 0 | 8.3962 | 0 | 8.6909 | 0 | 0.29475 | 0 | 0.069577 | 0 | 84.4406 | 0 | 0.0021327 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
42 | 43 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFNS | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.85 | 8.45 | 0 | Baseline | 0 | 38652.0472 | 0 | 0 | 8575.2047 | 0 | 0 | 4.5074 | 17.5717 | 8585.2246 | 0 | 0.99192 | 0 | 8.5479 | 0 | 0 | 0 | 8.3962 | 0 | 8.6909 | 0 | 0.29475 | 0 | 0.069577 | 0 | 84.4406 | 0 | 0.0021327 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
44 | 45 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFDS | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.19 | 8.45 | 0 | Baseline | 0 | 38652.0472 | 0 | 0 | 8575.2047 | 0 | 0 | 4.5074 | 34.7006 | 8585.2246 | 0 | 0.99192 | 0 | 8.5479 | 0 | 0 | 0 | 8.3962 | 0 | 8.6909 | 0 | 0.29475 | 0 | 0.069577 | 0 | 84.4406 | 0 | 0.0021327 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
48 | 49 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FBSA | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 4.6 | 8.86 | 0 | Baseline | 0 | 60041.8992 | 0 | 0 | 12808.9512 | 0 | 0 | 4.6875 | 7.9788 | 12809.5777 | 0 | 0.99771 | 0 | 9.056 | 0 | 0 | 0 | 8.8954 | 0 | 9.2378 | 0 | 0.34238 | 0 | 0.073056 | 0 | 108.8555 | 0 | 9.7593e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
50 | 51 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FPeSA | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7 | 8.86 | 0 | Baseline | 0 | 60041.8992 | 0 | 0 | 12808.9512 | 0 | 0 | 4.6875 | 7.9788 | 12809.5777 | 0 | 0.99771 | 0 | 9.056 | 0 | 0 | 0 | 8.8954 | 0 | 9.2378 | 0 | 0.34238 | 0 | 0.073056 | 0 | 108.8555 | 0 | 9.7593e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
52 | 53 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FHxSA | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.86 | 8.86 | 0 | Baseline | 0 | 60041.8992 | 0 | 0 | 12808.9512 | 0 | 0 | 4.6875 | 17.0375 | 12809.5777 | 0 | 0.99771 | 0 | 9.056 | 0 | 0 | 0 | 8.8954 | 0 | 9.2378 | 0 | 0.34238 | 0 | 0.073056 | 0 | 108.8555 | 0 | 9.7593e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
54 | 55 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FHpSA | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.58 | 8.86 | 0 | Baseline | 0 | 60041.8992 | 0 | 0 | 12808.9512 | 0 | 0 | 4.6875 | 30.5646 | 12809.5777 | 0 | 0.99771 | 0 | 9.056 | 0 | 0 | 0 | 8.8954 | 0 | 9.2378 | 0 | 0.34238 | 0 | 0.073056 | 0 | 108.8555 | 0 | 9.7593e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
56 | 57 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FOSA | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.88 | 8.86 | 0 | Baseline | 0 | 60041.8992 | 0 | 0 | 12808.9512 | 0 | 0 | 4.6875 | 33.704 | 12809.5777 | 0 | 0.99771 | 0 | 9.056 | 0 | 0 | 0 | 8.8954 | 0 | 9.2378 | 0 | 0.34238 | 0 | 0.073056 | 0 | 108.8555 | 0 | 9.7593e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
75 | 76 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 4:2 FTS | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 5.94 | 5.92 | 0 | Baseline | 0 | 18596.5495 | 0 | 0 | 3514.1376 | 0 | 0 | 5.2919 | 7.9815 | 3522.627 | 0 | 0.98627 | 0 | 6.1229 | 0 | 0 | 0 | 5.9947 | 0 | 6.2824 | 0 | 0.28768 | 0 | 0.078569 | 0 | 235.0157 | 0 | 0.0012575 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
77 | 78 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 6:2 FTS | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.92 | 7.92 | 0 | Baseline | 0 | 24812.8216 | 0 | 0 | 5488.9078 | 0 | 0 | 4.5205 | 7.9788 | 5488.9101 | 0 | 0.99819 | 0 | 8.0101 | 0 | 0 | 0 | 7.8624 | 0 | 8.2249 | 0 | 0.36255 | 0 | 0.070124 | 0 | 124.1407 | 0 | 9.7508e-07 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
79 | 80 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 8:2 FTS | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.88 | 8.89 | 0 | Baseline | 0 | 30501.4886 | 0 | 0 | 6230.4892 | 0 | 0 | 4.8955 | 20.7972 | 6233.7213 | 0 | 0.99427 | 0 | 8.9588 | 0 | 0 | 0 | 8.751 | 0 | 9.1439 | 0 | 0.39285 | 0 | 0.077462 | 0 | 134.4544 | 0 | 0.00093627 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
81 | 82 | PFAS CS0 0ng/ml | 1 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 2) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 10:2 FTS | 0 | Quantifiers | 10/26/2023 19:07:37 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.5 | 8.89 | 0 | Baseline | 0 | 30501.4886 | 0 | 0 | 6230.4892 | 0 | 0 | 4.8955 | 46.8438 | 6233.7213 | 0 | 0.99427 | 0 | 8.9588 | 0 | 0 | 0 | 8.751 | 0 | 9.1439 | 0 | 0.39285 | 0 | 0.077462 | 0 | 134.4544 | 0 | 0.00093627 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
142 | 143 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFBA | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 1.49 | 1.48 | 0 | Baseline | 0 | 108602.7871 | 0 | 0 | 16165.1146 | 0 | 0 | 6.7183 | 57.7433 | 16285.7027 | 0 | 0.96849 | 0 | 1.4822 | 0 | 0 | 0 | 1.351 | 0 | 1.6713 | 0 | 0.32028 | 0 | 0.10396 | 0 | 230.5712 | 0 | 0.016644 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
144 | 145 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFPeA | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 3.18 | 3.21 | 0 | Manual | 0 | 78244.8175 | 0 | 0 | 8210.3068 | 0 | 0 | 9.5301 | 63.2503 | 8225.4581 | 0 | 0.99131 | 0 | 3.2485 | 0 | 0 | 0 | 3.0185 | 0 | 3.541 | 0 | 0.52244 | 0 | 0.14158 | 0 | 625.1077 | 0 | 0.00075895 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
146 | 147 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFHxA | 0.0167250088152521 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 6.13 | 6.12 | Manual | Baseline | 611.5558 | 157116.4641 | 0.0038924 | 127.6784 | 28310.6821 | 0.0045099 | 4.7898 | 5.5497 | 131.0903 | 28316.0649 | 0.65978 | 0.99734 | 6.1858 | 6.1722 | 0.91086 | 0.055836 | 6.0503 | 6.0048 | 6.2645 | 6.4045 | 0.21416 | 0.39972 | 0.059768 | 0.086081 | 32.1813 | 283.8592 | 0.015478 | 0.00017924 | TRUE | 1.0022 | TRUE | [M-H]- | 0 | 6.0689 | 6.2559 | 0.18703 | 6.0816 | 6.2522 | 0.17056 | -9.228 | 0.8394 | 0.7287 | 0.77418 | 0 | 24 | 9 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 611.5558 | No Control Sample | 203843.5852 | off | off | 611.5558 | 203843.5852 | off | 0 | off | 0 | off | 0 | off | 312.9733 | 1.6137 | 0.50506 | 268.9809 | -5.4025 | -1.4532 | 0 | 0 | 0 | 0 | 0 | 10374.0624 | FALSE | 0 | 0.0038924
148 | 149 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFHpA | 0.0124573051886192 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 7.26 | 7.27 | Baseline | Baseline | 568.1335 | 182262.6724 | 0.0031171 | 154.8701 | 44581.819 | 0.0034738 | 3.6685 | 4.0883 | 160.331 | 44602.0065 | 0.62122 | 0.99581 | 7.2986 | 7.3086 | 0.53108 | 0.038556 | 7.2426 | 7.1784 | 7.3883 | 7.4593 | 0.14578 | 0.2809 | 0.043589 | 0.06276 | 30.1074 | 243.6554 | 0.040678 | 0.00071217 | FALSE | 0.99863 | TRUE | [M-H]- | 0 | 7.2496 | 7.3775 | 0.12793 | 7.2551 | 7.3736 | 0.11851 | -43.2134 | 1.3986 | 1.9487 | 1.4732 | 0 | 18 | 7 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 568.1335 | No Control Sample | 228328.7703 | off | off | 568.1335 | 228328.7703 | off | 0 | off | 0 | off | 0 | off | 362.9716 | 5.4383 | 1.9739 | 318.9805 | 4.3567 | 1.3897 | 0 | 0 | 0 | 0 | 0 | 25.2282 | FALSE | 0 | 0.0031171
150 | 151 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFOA | 0.0126283269836372 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 7.96 | 7.94 | Baseline | Baseline | 533.6051 | 174231.0153 | 0.0030626 | 164.5303 | 40758.4677 | 0.0040367 | 3.2432 | 4.2747 | 183.5025 | 40766.226 | 0.38242 | 0.99793 | 7.9983 | 7.9911 | 0.48112 | 0.038297 | 7.9517 | 7.8696 | 8.0827 | 8.2337 | 0.131 | 0.36415 | 0.049574 | 0.066598 | 22.2973 | 133.1676 | 0.0011063 | 0.0002016 | FALSE | 1.0009 | TRUE | [M-H]- | 0 | 7.9568 | 8.0727 | 0.11599 | 7.9619 | 8.0631 | 0.1012 | -1.3894 | 1.3611 | 1.7002 | 1.0333 | 0 | 10 | 5 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 533.6051 | No Control Sample | 191475.4056 | off | off | 533.6051 | 191475.4056 | off | 0 | off | 0 | off | 0 | off | 412.9671 | 1.5968 | 0.65943 | 368.9764 | 0.69278 | 0.25562 | 0 | 0 | 0 | 0 | 0 | 23.5614 | FALSE | 0 | 0.0030626
152 | 153 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFNA | 0.00449535572901603 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 8.46 | 8.47 | Manual | Baseline | 166.9983 | 159913.3037 | 0.0010443 | 53.0001 | 33669.0823 | 0.0015741 | 3.1509 | 4.7496 | 56.1071 | 33681.4929 | 0.35951 | 0.99656 | 8.5237 | 8.502 | 0.75338 | 0.063736 | 8.4652 | 8.2994 | 8.5737 | 8.6638 | 0.10843 | 0.36443 | 0.051612 | 0.074391 | 11.1335 | 106.6106 | 0.01064 | 0.00067485 | TRUE | 1.0026 | TRUE | [M-H]- | 0 | 8.4691 | 8.5707 | 0.10157 | 8.474 | 8.5677 | 0.093733 | 5.2008 | 0.96962 | 0.97211 | 0 | 0 | 8 | 5 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Fail | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 166.9983 | No Control Sample | 159143.3885 | off | off | 166.9983 | 159143.3885 | off | 0 | off | 0 | off | 0 | off | 463.2173 | 548.8338 | 254.0899 | 418.9774 | 9.9517 | 4.1695 | 0 | 0 | 0 | 0 | 0 | inf | FALSE | 0 | 0.0010443
154 | 155 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFDA | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.87 | 8.86 | 0 | Baseline | 0 | 129998.6259 | 0 | 0 | 25844.4973 | 0 | 0 | 5.03 | 30.5277 | 25850.968 | 0 | 0.99806 | 0 | 8.9166 | 0 | 0 | 0 | 8.7211 | 0 | 9.1276 | 0 | 0.40657 | 0 | 0.078666 | 0 | 157.7745 | 0 | 4.6467e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
156 | 157 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFUdA | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.21 | 9.2 | 0 | Baseline | 0 | 109332.2039 | 0 | 0 | 25208.2565 | 0 | 0 | 4.3372 | 12.8286 | 25216.9502 | 0 | 0.99684 | 0 | 9.2494 | 0 | 0 | 0 | 9.1084 | 0 | 9.4113 | 0 | 0.30285 | 0 | 0.067755 | 0 | 238.7865 | 0 | 0.000718 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
158 | 159 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFDoA | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.47 | 9.47 | 0 | Baseline | 0 | 89747.2246 | 0 | 0 | 22075.4934 | 0 | 0 | 4.0655 | 25.0031 | 22079.4979 | 0 | 0.99775 | 0 | 9.5161 | 0 | 0 | 0 | 9.3786 | 0 | 9.6798 | 0 | 0.3012 | 0 | 0.063857 | 0 | 261.2318 | 0 | 0.00041455 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
160 | 161 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFTrDA | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.68 | 9.47 | 0 | Baseline | 0 | 89747.2246 | 0 | 0 | 22075.4934 | 0 | 0 | 4.0655 | 9.6882 | 22079.4979 | 0 | 0.99775 | 0 | 9.5161 | 0 | 0 | 0 | 9.3786 | 0 | 9.6798 | 0 | 0.3012 | 0 | 0.063857 | 0 | 261.2318 | 0 | 0.00041455 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
162 | 163 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFTeDA | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.87 | 9.86 | 0 | Baseline | 0 | 18716.1191 | 0 | 0 | 5373.2202 | 0 | 0 | 3.4832 | 69.159 | 5373.2229 | 0 | 0.99722 | 0 | 9.9087 | 0 | 0 | 0 | 9.7903 | 0 | 10.0297 | 0 | 0.23938 | 0 | 0.055114 | 0 | 753.7201 | 0 | 1.0013e-06 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
169 | 170 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFBS | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 3.5 | 3.5 | 0 | Manual | 0 | 42420.1472 | 0 | 0 | 3855.6147 | 0 | 0 | 11.0022 | 17.7463 | 3861.5366 | 0 | 0.99275 | 0 | 3.9033 | 0 | 0 | 0 | 3.6329 | 0 | 4.2097 | 0 | 0.57686 | 0 | 0.16995 | 0 | 859.918 | 0 | 0.0016692 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
171 | 172 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFPeS | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 6.38 | 3.5 | 0 | Manual | 0 | 42420.1472 | 0 | 0 | 3855.6147 | 0 | 0 | 11.0022 | 8.0675 | 3861.5366 | 0 | 0.99275 | 0 | 3.9033 | 0 | 0 | 0 | 3.6329 | 0 | 4.2097 | 0 | 0.57686 | 0 | 0.16995 | 0 | 859.918 | 0 | 0.0016692 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
173 | 174 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | L-PFHxS | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.34 | 7.35 | 0 | Baseline | 0 | 61149.0723 | 0 | 0 | 14844.0878 | 0 | 0 | 4.1194 | 49.8504 | 14868.7307 | 0 | 0.98653 | 0 | 7.3819 | 0 | 0 | 0 | 7.2365 | 0 | 7.5232 | 0 | 0.28663 | 0 | 0.063142 | 0 | 149.0301 | 0 | 0.0024356 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
175 | 176 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | Br-PFHxS | 0 | Qualifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.16 | 7.35 | 0 | Baseline | 0 | 61149.0723 | 0 | 0 | 14844.0878 | 0 | 0 | 4.1194 | 49.8504 | 14868.7307 | 0 | 0.98653 | 0 | 7.3819 | 0 | 0 | 0 | 7.2365 | 0 | 7.5232 | 0 | 0.28663 | 0 | 0.063142 | 0 | 149.0301 | 0 | 0.0024356 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
177 | 178 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFHpS | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.98 | 7.35 | 0 | Baseline | 0 | 61149.0723 | 0 | 0 | 14844.0878 | 0 | 0 | 4.1194 | 7.9788 | 14868.7307 | 0 | 0.98653 | 0 | 7.3819 | 0 | 0 | 0 | 7.2365 | 0 | 7.5232 | 0 | 0.28663 | 0 | 0.063142 | 0 | 149.0301 | 0 | 0.0024356 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
179 | 180 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | L-PFOS | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.46 | 8.45 | 0 | Baseline | 0 | 41418.4752 | 0 | 0 | 9013.4435 | 0 | 0 | 4.5952 | 22.4971 | 9014.2078 | 0 | 0.99913 | 0 | 8.5004 | 0 | 0 | 0 | 8.3362 | 0 | 8.8042 | 0 | 0.46798 | 0 | 0.071593 | 0 | 139.9997 | 0 | 0.00012936 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
181 | 182 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | Br-PFOS | 0 | Qualifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.3 | 8.45 | 0 | Baseline | 0 | 41418.4752 | 0 | 0 | 9013.4435 | 0 | 0 | 4.5952 | 22.4971 | 9014.2078 | 0 | 0.99913 | 0 | 8.5004 | 0 | 0 | 0 | 8.3362 | 0 | 8.8042 | 0 | 0.46798 | 0 | 0.071593 | 0 | 139.9997 | 0 | 0.00012936 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
183 | 184 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFNS | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.85 | 8.45 | 0 | Baseline | 0 | 41418.4752 | 0 | 0 | 9013.4435 | 0 | 0 | 4.5952 | 23.9365 | 9014.2078 | 0 | 0.99913 | 0 | 8.5004 | 0 | 0 | 0 | 8.3362 | 0 | 8.8042 | 0 | 0.46798 | 0 | 0.071593 | 0 | 139.9997 | 0 | 0.00012936 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
185 | 186 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFDS | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.19 | 8.45 | 0 | Baseline | 0 | 41418.4752 | 0 | 0 | 9013.4435 | 0 | 0 | 4.5952 | 13.9894 | 9014.2078 | 0 | 0.99913 | 0 | 8.5004 | 0 | 0 | 0 | 8.3362 | 0 | 8.8042 | 0 | 0.46798 | 0 | 0.071593 | 0 | 139.9997 | 0 | 0.00012936 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
189 | 190 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FBSA | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 4.6 | 8.86 | 0 | Baseline | 0 | 70706.0005 | 0 | 0 | 15153.0039 | 0 | 0 | 4.6661 | 7.9788 | 15161.8131 | 0 | 0.99441 | 0 | 9.0217 | 0 | 0 | 0 | 8.8051 | 0 | 9.2094 | 0 | 0.40433 | 0 | 0.072496 | 0 | 158.3775 | 0 | 0.0011067 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
191 | 192 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FPeSA | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7 | 8.86 | 0 | Baseline | 0 | 70706.0005 | 0 | 0 | 15153.0039 | 0 | 0 | 4.6661 | 7.9788 | 15161.8131 | 0 | 0.99441 | 0 | 9.0217 | 0 | 0 | 0 | 8.8051 | 0 | 9.2094 | 0 | 0.40433 | 0 | 0.072496 | 0 | 158.3775 | 0 | 0.0011067 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
193 | 194 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FHxSA | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.86 | 8.86 | 0 | Baseline | 0 | 70706.0005 | 0 | 0 | 15153.0039 | 0 | 0 | 4.6661 | 20.7998 | 15161.8131 | 0 | 0.99441 | 0 | 9.0217 | 0 | 0 | 0 | 8.8051 | 0 | 9.2094 | 0 | 0.40433 | 0 | 0.072496 | 0 | 158.3775 | 0 | 0.0011067 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
195 | 196 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FHpSA | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.58 | 8.86 | 0 | Baseline | 0 | 70706.0005 | 0 | 0 | 15153.0039 | 0 | 0 | 4.6661 | 12.8183 | 15161.8131 | 0 | 0.99441 | 0 | 9.0217 | 0 | 0 | 0 | 8.8051 | 0 | 9.2094 | 0 | 0.40433 | 0 | 0.072496 | 0 | 158.3775 | 0 | 0.0011067 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
197 | 198 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FOSA | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.88 | 8.86 | 0 | Baseline | 0 | 70706.0005 | 0 | 0 | 15153.0039 | 0 | 0 | 4.6661 | 22.4971 | 15161.8131 | 0 | 0.99441 | 0 | 9.0217 | 0 | 0 | 0 | 8.8051 | 0 | 9.2094 | 0 | 0.40433 | 0 | 0.072496 | 0 | 158.3775 | 0 | 0.0011067 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
216 | 217 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 4:2 FTS | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 5.94 | 5.92 | 0 | Baseline | 0 | 18180.9415 | 0 | 0 | 3450.0958 | 0 | 0 | 5.2697 | 7.9788 | 3464.9505 | 0 | 0.97926 | 0 | 5.9769 | 0 | 0 | 0 | 5.8278 | 0 | 6.1384 | 0 | 0.31055 | 0 | 0.079942 | 0 | 202.5007 | 0 | 0.0043148 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
218 | 219 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 6:2 FTS | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.92 | 7.92 | 0 | Baseline | 0 | 17262.4149 | 0 | 0 | 4262.9973 | 0 | 0 | 4.0494 | 18.2141 | 4263.9397 | 0 | 0.99647 | 0 | 7.954 | 0 | 0 | 0 | 7.831 | 0 | 8.1258 | 0 | 0.29482 | 0 | 0.060275 | 0 | 95.4112 | 0 | 0.00050465 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
220 | 221 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 8:2 FTS | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.88 | 8.89 | 0 | Baseline | 0 | 24187.3408 | 0 | 0 | 4714.7775 | 0 | 0 | 5.1301 | 0 | 4714.8953 | 0 | 0.99804 | 0 | 8.9121 | 0 | 0 | 0 | 8.7449 | 0 | 9.0945 | 0 | 0.34963 | 0 | 0.079622 | 0 | 149.7742 | 0 | 5.0805e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
222 | 223 | PFAS CS0 0ng/ml | 2 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 3) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 10:2 FTS | 0 | Quantifiers | 10/26/2023 19:24:17 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.5 | 8.89 | 0 | Baseline | 0 | 24187.3408 | 0 | 0 | 4714.7775 | 0 | 0 | 5.1301 | 0 | 4714.8953 | 0 | 0.99804 | 0 | 8.9121 | 0 | 0 | 0 | 8.7449 | 0 | 9.0945 | 0 | 0.34963 | 0 | 0.079622 | 0 | 149.7742 | 0 | 5.0805e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
283 | 284 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFBA | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 1.49 | 1.48 | 0 | Baseline | 0 | 103075.2544 | 0 | 0 | 13661.2192 | 0 | 0 | 7.5451 | 66.4158 | 13824.3028 | 0 | 0.95546 | 0 | 1.4955 | 0 | 0 | 0 | 1.3508 | 0 | 1.6808 | 0 | 0.33003 | 0 | 0.11902 | 0 | 168.8468 | 0 | 0.027357 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
285 | 286 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFPeA | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 3.18 | 3.21 | 0 | Baseline | 0 | 77216.9643 | 0 | 0 | 8178.75 | 0 | 0 | 9.4412 | 36.5725 | 8223.0163 | 0 | 0.9796 | 0 | 3.2791 | 0 | 0 | 0 | 3.077 | 0 | 3.5814 | 0 | 0.50443 | 0 | 0.15339 | 0 | 545.9731 | 0 | 0.0034257 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
287 | 288 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFHxA | 0.0144684258637346 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 6.13 | 6.12 | Baseline | Baseline | 514.5247 | 152804.7565 | 0.0033672 | 136.5377 | 26895.958 | 0.0050765 | 3.7684 | 5.6813 | 139.9194 | 26923.837 | 0.59901 | 0.99359 | 6.2063 | 6.2044 | 1.2449 | 0.076313 | 6.1153 | 6.0063 | 6.2986 | 6.3577 | 0.1833 | 0.35142 | 0.053481 | 0.083819 | 30.5211 | 217.221 | 0.039628 | 0.0018661 | FALSE | 1.0003 | TRUE | [M-H]- | 0 | 6.1365 | 6.2787 | 0.14216 | 6.142 | 6.2691 | 0.12703 | 29.5186 | 0.97719 | 0.88878 | 0.79091 | 0 | 20 | 7 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 514.5247 | No Control Sample | 203843.5852 | off | off | 514.5247 | 203843.5852 | off | 0 | off | 0 | off | 0 | off | 312.9746 | 5.6316 | 1.7625 | 268.9839 | 5.7152 | 1.5373 | 0 | 0 | 0 | 0 | 0 | 8738.2937 | FALSE | 0 | 0.0033672
289 | 290 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFHpA | 0.017046717384388 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 7.26 | 7.27 | Baseline | Baseline | 784.6585 | 183954.9165 | 0.0042655 | 171.2868 | 44777.6077 | 0.0038253 | 4.581 | 4.1082 | 175.3242 | 44782.4141 | 0.70838 | 0.99843 | 7.3262 | 7.333 | 0.91145 | 0.066171 | 7.2407 | 7.1957 | 7.4376 | 7.5324 | 0.1969 | 0.3367 | 0.077311 | 0.062969 | 39.7267 | 226.7658 | 0.040329 | 0.00013139 | FALSE | 0.99907 | TRUE | [M-H]- | 0 | 7.2509 | 7.415 | 0.16409 | 7.2607 | 7.4084 | 0.14771 | -35.0826 | 1.0687 | 1.2067 | 1.563 | 0 | 24 | 11 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 784.6585 | No Control Sample | 228328.7703 | off | off | 784.6585 | 228328.7703 | off | 0 | off | 0 | off | 0 | off | 362.9696 | 0.067207 | 0.024394 | 318.9743 | -15.1189 | -4.8226 | 0 | 0 | 0 | 0 | 0 | 28.9936 | FALSE | 0 | 0.0042655
291 | 292 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFOA | 0.0126293601865284 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 7.96 | 7.94 | Baseline | Baseline | 513.3312 | 167597.5352 | 0.0030629 | 109.9027 | 37383.5247 | 0.0029399 | 4.6708 | 4.4832 | 120.4715 | 37385.2505 | 0.4436 | 0.99902 | 8.0234 | 8.0145 | 0.79701 | 0.063442 | 7.949 | 7.8536 | 8.1433 | 8.2144 | 0.19433 | 0.36082 | 0.070783 | 0.069659 | 18.7798 | 105.8424 | 0.053069 | 9.3983e-05 | FALSE | 1.0011 | TRUE | [M-H]- | 0 | 7.9595 | 8.1158 | 0.15624 | 7.9665 | 8.1071 | 0.14055 | -30.0123 | 1.1043 | 1.2037 | 0.95701 | 0 | 14 | 6 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 513.3312 | No Control Sample | 191475.4056 | off | off | 513.3312 | 191475.4056 | off | 0 | off | 0 | off | 0 | off | 412.9669 | 1.2148 | 0.50167 | 368.9774 | 3.457 | 1.2756 | 0 | 0 | 0 | 0 | 0 | 83.2964 | FALSE | 0 | 0.0030629
293 | 294 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFNA | 0.00630870354335214 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 8.46 | 8.47 | Manual | Baseline | 221.0576 | 150834.9551 | 0.0014656 | 52.5678 | 32011.0858 | 0.0016422 | 4.2052 | 4.712 | 54.9145 | 32014.576 | 0.53604 | 0.99883 | 8.5033 | 8.5179 | 0.51156 | 0.043278 | 8.4218 | 8.312 | 8.5706 | 8.8099 | 0.14876 | 0.4979 | 0.065438 | 0.072767 | 16.0148 | 140.6355 | 0.0052495 | 0.00025243 | TRUE | 0.99828 | TRUE | [M-H]- | 0 | 8.4285 | 8.568 | 0.13951 | 8.4365 | 8.5616 | 0.12508 | 1.855 | 1.021 | 1.0741 | 0 | 0 | 10 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Fail | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 221.0576 | No Control Sample | 159143.3885 | off | off | 221.0576 | 159143.3885 | off | 0 | off | 0 | off | 0 | off | 462.9728 | 20.7313 | 9.5978 | 418.9795 | 14.9412 | 6.26 | 0 | 0 | 0 | 0 | 0 | 99.1548 | FALSE | 0 | 0.0014656
295 | 296 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFDA | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.87 | 8.86 | 0 | Baseline | 0 | 143100.1392 | 0 | 0 | 29696.3296 | 0 | 0 | 4.8188 | 47.5134 | 29705.3663 | 0 | 0.99799 | 0 | 8.9302 | 0 | 0 | 0 | 8.7306 | 0 | 9.0803 | 0 | 0.34972 | 0 | 0.075302 | 0 | 146.2253 | 0 | 0.00051041 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
297 | 298 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFUdA | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.21 | 9.2 | 0 | Baseline | 0 | 111616.2338 | 0 | 0 | 24834.9006 | 0 | 0 | 4.4943 | 27.7962 | 24846.4714 | 0 | 0.99696 | 0 | 9.263 | 0 | 0 | 0 | 9.1194 | 0 | 9.4183 | 0 | 0.29892 | 0 | 0.071294 | 0 | 136.8561 | 0 | 0.00099656 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
299 | 300 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFDoA | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.47 | 9.47 | 0 | Baseline | 0 | 94628.8826 | 0 | 0 | 24590.6977 | 0 | 0 | 3.8482 | 41.5942 | 24591.361 | 0 | 0.99875 | 0 | 9.5304 | 0 | 0 | 0 | 9.3691 | 0 | 9.6989 | 0 | 0.32982 | 0 | 0.059423 | 0 | 335.916 | 0 | 5.4645e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
301 | 302 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFTrDA | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.68 | 9.47 | 0 | Baseline | 0 | 94628.8826 | 0 | 0 | 24590.6977 | 0 | 0 | 3.8482 | 44.4055 | 24591.361 | 0 | 0.99875 | 0 | 9.5304 | 0 | 0 | 0 | 9.3691 | 0 | 9.6989 | 0 | 0.32982 | 0 | 0.059423 | 0 | 335.916 | 0 | 5.4645e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
303 | 304 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFTeDA | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.87 | 9.86 | 0 | Baseline | 0 | 20918.1168 | 0 | 0 | 6081.8561 | 0 | 0 | 3.4394 | 108.4498 | 6084.6506 | 0 | 0.99452 | 0 | 9.9199 | 0 | 0 | 0 | 9.8167 | 0 | 10.0274 | 0 | 0.21078 | 0 | 0.055548 | 0 | 739.4245 | 0 | 0.0009483 | TRUE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
310 | 311 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFBS | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 3.5 | 3.5 | 0 | Manual | 0 | 42930.7713 | 0 | 0 | 3706.4166 | 0 | 0 | 11.5828 | 22.9594 | 3718.3044 | 0 | 0.99024 | 0 | 3.9482 | 0 | 0 | 0 | 3.714 | 0 | 4.2657 | 0 | 0.55164 | 0 | 0.17291 | 0 | 750.9889 | 0 | 0.0012413 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
312 | 313 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFPeS | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 6.38 | 3.5 | 0 | Manual | 0 | 42930.7713 | 0 | 0 | 3706.4166 | 0 | 0 | 11.5828 | 17.6577 | 3718.3044 | 0 | 0.99024 | 0 | 3.9482 | 0 | 0 | 0 | 3.714 | 0 | 4.2657 | 0 | 0.55164 | 0 | 0.17291 | 0 | 750.9889 | 0 | 0.0012413 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
314 | 315 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | L-PFHxS | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.34 | 7.35 | 0 | Baseline | 0 | 62650.3786 | 0 | 0 | 15562.2351 | 0 | 0 | 4.0258 | 41.5919 | 15584.0843 | 0 | 0.98825 | 0 | 7.4063 | 0 | 0 | 0 | 7.2709 | 0 | 7.5499 | 0 | 0.27902 | 0 | 0.061538 | 0 | 186.4154 | 0 | 0.0021983 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
316 | 317 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | Br-PFHxS | 0 | Qualifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.16 | 7.35 | 0 | Baseline | 0 | 62650.3786 | 0 | 0 | 15562.2351 | 0 | 0 | 4.0258 | 41.5919 | 15584.0843 | 0 | 0.98825 | 0 | 7.4063 | 0 | 0 | 0 | 7.2709 | 0 | 7.5499 | 0 | 0.27902 | 0 | 0.061538 | 0 | 186.4154 | 0 | 0.0021983 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
318 | 319 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFHpS | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.98 | 7.35 | 0 | Baseline | 0 | 62650.3786 | 0 | 0 | 15562.2351 | 0 | 0 | 4.0258 | 7.9789 | 15584.0843 | 0 | 0.98825 | 0 | 7.4063 | 0 | 0 | 0 | 7.2709 | 0 | 7.5499 | 0 | 0.27902 | 0 | 0.061538 | 0 | 186.4154 | 0 | 0.0021983 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
320 | 321 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | L-PFOS | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.46 | 8.45 | 0 | Baseline | 0 | 42141.108 | 0 | 0 | 9307.1242 | 0 | 0 | 4.5278 | 28.776 | 9309.4546 | 0 | 0.99531 | 0 | 8.5183 | 0 | 0 | 0 | 8.3684 | 0 | 8.6417 | 0 | 0.27328 | 0 | 0.069824 | 0 | 247.7339 | 0 | 0.00043551 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
322 | 323 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | Br-PFOS | 0 | Qualifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.3 | 8.45 | 0 | Baseline | 0 | 42141.108 | 0 | 0 | 9307.1242 | 0 | 0 | 4.5278 | 28.776 | 9309.4546 | 0 | 0.99531 | 0 | 8.5183 | 0 | 0 | 0 | 8.3684 | 0 | 8.6417 | 0 | 0.27328 | 0 | 0.069824 | 0 | 247.7339 | 0 | 0.00043551 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
324 | 325 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFNS | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.85 | 8.45 | 0 | Baseline | 0 | 42141.108 | 0 | 0 | 9307.1242 | 0 | 0 | 4.5278 | 17.0375 | 9309.4546 | 0 | 0.99531 | 0 | 8.5183 | 0 | 0 | 0 | 8.3684 | 0 | 8.6417 | 0 | 0.27328 | 0 | 0.069824 | 0 | 247.7339 | 0 | 0.00043551 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
326 | 327 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFDS | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.19 | 8.45 | 0 | Baseline | 0 | 42141.108 | 0 | 0 | 9307.1242 | 0 | 0 | 4.5278 | 15.9604 | 9309.4546 | 0 | 0.99531 | 0 | 8.5183 | 0 | 0 | 0 | 8.3684 | 0 | 8.6417 | 0 | 0.27328 | 0 | 0.069824 | 0 | 247.7339 | 0 | 0.00043551 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
330 | 331 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FBSA | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 4.6 | 8.86 | 0 | Baseline | 0 | 76040.0693 | 0 | 0 | 16798.2074 | 0 | 0 | 4.5267 | 9.6788 | 16801.7633 | 0 | 0.9969 | 0 | 9.0432 | 0 | 0 | 0 | 8.8356 | 0 | 9.2593 | 0 | 0.42362 | 0 | 0.068583 | 0 | 150.5847 | 0 | 0.00041666 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
332 | 333 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FPeSA | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7 | 8.86 | 0 | Baseline | 0 | 76040.0693 | 0 | 0 | 16798.2074 | 0 | 0 | 4.5267 | 7.9788 | 16801.7633 | 0 | 0.9969 | 0 | 9.0432 | 0 | 0 | 0 | 8.8356 | 0 | 9.2593 | 0 | 0.42362 | 0 | 0.068583 | 0 | 150.5847 | 0 | 0.00041666 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
334 | 335 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FHxSA | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.86 | 8.86 | 0 | Baseline | 0 | 76040.0693 | 0 | 0 | 16798.2074 | 0 | 0 | 4.5267 | 21.8769 | 16801.7633 | 0 | 0.9969 | 0 | 9.0432 | 0 | 0 | 0 | 8.8356 | 0 | 9.2593 | 0 | 0.42362 | 0 | 0.068583 | 0 | 150.5847 | 0 | 0.00041666 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
336 | 337 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FHpSA | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.58 | 8.86 | 0 | Baseline | 0 | 76040.0693 | 0 | 0 | 16798.2074 | 0 | 0 | 4.5267 | 12.8183 | 16801.7633 | 0 | 0.9969 | 0 | 9.0432 | 0 | 0 | 0 | 8.8356 | 0 | 9.2593 | 0 | 0.42362 | 0 | 0.068583 | 0 | 150.5847 | 0 | 0.00041666 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
338 | 339 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FOSA | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.88 | 8.86 | 0 | Baseline | 0 | 76040.0693 | 0 | 0 | 16798.2074 | 0 | 0 | 4.5267 | 68.6702 | 16801.7633 | 0 | 0.9969 | 0 | 9.0432 | 0 | 0 | 0 | 8.8356 | 0 | 9.2593 | 0 | 0.42362 | 0 | 0.068583 | 0 | 150.5847 | 0 | 0.00041666 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
357 | 358 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 4:2 FTS | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 5.94 | 5.92 | 0 | Baseline | 0 | 17750.6935 | 0 | 0 | 3037.5041 | 0 | 0 | 5.8438 | 1.4472 | 3043.1477 | 0 | 0.98865 | 0 | 6.0113 | 0 | 0 | 0 | 5.838 | 0 | 6.1581 | 0 | 0.32008 | 0 | 0.084369 | 0 | 191.5119 | 0 | 0.003422 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
359 | 360 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 6:2 FTS | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.92 | 7.92 | 0 | Baseline | 0 | 16130.1415 | 0 | 0 | 3616.1457 | 0 | 0 | 4.4606 | 7.9788 | 3616.1457 | 0 | 1 | 0 | 7.9764 | 0 | 0 | 0 | 7.8035 | 0 | 8.2183 | 0 | 0.41478 | 0 | 0.068838 | 0 | 92.2843 | 0 | 0 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
361 | 362 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 8:2 FTS | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.88 | 8.89 | 0 | Baseline | 0 | 20317.4666 | 0 | 0 | 4114.577 | 0 | 0 | 4.9379 | 7.9788 | 4120.9638 | 0 | 0.9917 | 0 | 8.9249 | 0 | 0 | 0 | 8.7754 | 0 | 9.0677 | 0 | 0.29227 | 0 | 0.07692 | 0 | 147.6092 | 0 | 0.0030755 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
363 | 364 | PFAS CS0 0ng/ml | 3 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 4) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 10:2 FTS | 0 | Quantifiers | 10/26/2023 19:40:53 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.5 | 8.89 | 0 | Baseline | 0 | 20317.4666 | 0 | 0 | 4114.577 | 0 | 0 | 4.9379 | 23.0168 | 4120.9638 | 0 | 0.9917 | 0 | 8.9249 | 0 | 0 | 0 | 8.7754 | 0 | 9.0677 | 0 | 0.29227 | 0 | 0.07692 | 0 | 147.6092 | 0 | 0.0030755 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
424 | 425 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFBA | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 1.49 | 1.48 | 0 | Manual | 0 | 105545.4962 | 0 | 0 | 14427.4129 | 0 | 0 | 7.3156 | 58.2694 | 14459.5444 | 0 | 0.99007 | 0 | 1.4836 | 0 | 0 | 0 | 1.359 | 0 | 1.7108 | 0 | 0.3518 | 0 | 0.10909 | 0 | 213.977 | 0 | 0 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
426 | 427 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFPeA | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 3.18 | 3.21 | 0 | Baseline | 0 | 73872.504 | 0 | 0 | 8173.6894 | 0 | 0 | 9.0378 | 62.389 | 8335.0583 | 0 | 0.93782 | 0 | 3.2416 | 0 | 0 | 0 | 3.0703 | 0 | 3.4875 | 0 | 0.41728 | 0 | 0.14572 | 0 | 514.4503 | 0 | 0.0061231 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
428 | 429 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFHxA | 0.0142041471606016 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy, Integration Quality | 0 | 0 | 0 | 4 | 0 | 0 | 6.13 | 6.12 | Manual | Baseline | 515.4807 | 155936.9998 | 0.0033057 | 83.476 | 29065.9802 | 0.0028719 | 6.1752 | 5.3649 | 85.1883 | 29099.5688 | 0.701 | 0.99239 | 6.1569 | 6.1615 | 0.4383 | 0.026868 | 6.0177 | 5.9885 | 6.2932 | 6.32 | 0.27557 | 0.33155 | 0.12439 | 0.081955 | 18.9826 | 269.0825 | 0 | 0.0022556 | TRUE | 0.99925 | TRUE | [M-H]- | 0 | 6.0237 | 6.2851 | 0.2614 | 6.0462 | 6.2617 | 0.21554 | 0 | 0.94222 | 0.85505 | 0.78941 | 0 | 31 | 15 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 515.4807 | No Control Sample | 203843.5852 | off | off | 515.4807 | 203843.5852 | off | 0 | off | 0 | off | 0 | off | 312.9723 | -1.5257 | -0.47751 | 268.9831 | 2.5453 | 0.68465 | 0 | 0 | 0 | 0 | 0 | 9751.4677 | FALSE | 0 | 0.0033057
430 | 431 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFHpA | 0.015756535385799 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 7.26 | 7.27 | Baseline | Baseline | 715.1398 | 181385.1331 | 0.0039427 | 193.4081 | 43669.3029 | 0.0044289 | 3.6976 | 4.1536 | 196.9736 | 43720.0342 | 0.71983 | 0.99311 | 7.3056 | 7.2995 | 0.62764 | 0.045566 | 7.2303 | 7.1665 | 7.3776 | 7.4223 | 0.1473 | 0.25585 | 0.057923 | 0.064088 | 43.7788 | 228.4689 | 0.023679 | 0.0016393 | FALSE | 1.0008 | TRUE | [M-H]- | 0 | 7.2345 | 7.3596 | 0.1251 | 7.2398 | 7.3561 | 0.11633 | -31.0916 | 0.94191 | 0.90535 | 1.3017 | 0 | 18 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 715.1398 | No Control Sample | 228328.7703 | off | off | 715.1398 | 228328.7703 | off | 0 | off | 0 | off | 0 | off | 362.9712 | 4.2278 | 1.5345 | 318.9821 | 9.2835 | 2.9612 | 0 | 0 | 0 | 0 | 0 | 55.4193 | FALSE | 0 | 0.0039427
432 | 433 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFOA | 0.0103844741667111 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 7.96 | 7.94 | Manual | Baseline | 437.3008 | 173638.9233 | 0.0025184 | 90.3301 | 39310.8074 | 0.0022978 | 4.8411 | 4.4171 | 98.8869 | 39329.0945 | 0.3794 | 0.99597 | 7.9744 | 7.986 | 0.18119 | 0.014423 | 7.9039 | 7.8543 | 8.0831 | 8.1673 | 0.17918 | 0.31305 | 0.068334 | 0.067915 | 12.0183 | 107.9313 | 0.0055062 | 0.00025029 | TRUE | 0.99855 | TRUE | [M-H]- | 0 | 7.9073 | 8.0756 | 0.16834 | 7.9119 | 8.0704 | 0.15855 | 2.7758 | 1.2742 | 1.5805 | 1.6698 | 0 | 15 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 437.3008 | No Control Sample | 191475.4056 | off | off | 437.3008 | 191475.4056 | off | 0 | off | 0 | off | 0 | off | 412.9677 | 3.1719 | 1.3099 | 368.981 | 13.1554 | 4.854 | 0 | 0 | 0 | 0 | 0 | 26.6191 | FALSE | 0 | 0.0025184
434 | 435 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFNA | 0.00767492690183709 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0 | 4 | 0 | 0 | 8.46 | 8.47 | Manual | Baseline | 252.7484 | 141759.0917 | 0.0017829 | 49.4732 | 29384.3126 | 0.0016837 | 5.1088 | 4.8243 | 53.8102 | 29385.5535 | 0.564 | 0.99932 | 8.491 | 8.4969 | 0.36599 | 0.030963 | 8.3898 | 8.2964 | 8.5954 | 8.6908 | 0.20558 | 0.39443 | 0.061564 | 0.074651 | 12.3905 | 153.4684 | 0.010931 | 7.9535e-05 | TRUE | 0.9993 | TRUE | [M-H]- | 0 | 8.3959 | 8.5951 | 0.19926 | 8.4016 | 8.591 | 0.18935 | 2.6307 | 0.98846 | 0.99262 | 0 | 0 | 13 | 5 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Fail | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 252.7484 | No Control Sample | 159143.3885 | off | off | 252.7484 | 159143.3885 | off | 0 | off | 0 | off | 0 | off | 463.3057 | 739.7742 | 342.4883 | 418.9805 | 17.4808 | 7.324 | 0 | 0 | 0 | 0 | 0 | inf | FALSE | 0 | 0.0017829
436 | 437 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFDA | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.87 | 8.86 | 0 | Baseline | 0 | 134859.2051 | 0 | 0 | 28919.7778 | 0 | 0 | 4.6632 | 36.5411 | 28926.7127 | 0 | 0.99795 | 0 | 8.9135 | 0 | 0 | 0 | 8.7117 | 0 | 9.0802 | 0 | 0.3685 | 0 | 0.072318 | 0 | 126.6932 | 0 | 0.00042587 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
438 | 439 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFUdA | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.21 | 9.2 | 0 | Baseline | 0 | 113852.6749 | 0 | 0 | 24144.8025 | 0 | 0 | 4.7154 | 26.7163 | 24146.2253 | 0 | 0.99922 | 0 | 9.2468 | 0 | 0 | 0 | 9.0811 | 0 | 9.4511 | 0 | 0.37002 | 0 | 0.074545 | 0 | 149.0995 | 0 | 0.00013047 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
440 | 441 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFDoA | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.47 | 9.47 | 0 | Baseline | 0 | 89224.3172 | 0 | 0 | 21903.502 | 0 | 0 | 4.0735 | 45.1012 | 21910.025 | 0 | 0.99722 | 0 | 9.5102 | 0 | 0 | 0 | 9.3691 | 0 | 9.6735 | 0 | 0.30442 | 0 | 0.063327 | 0 | 252.5596 | 0 | 0.00060398 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
442 | 443 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFTrDA | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.68 | 9.47 | 0 | Baseline | 0 | 89224.3172 | 0 | 0 | 21903.502 | 0 | 0 | 4.0735 | 39.5347 | 21910.025 | 0 | 0.99722 | 0 | 9.5102 | 0 | 0 | 0 | 9.3691 | 0 | 9.6735 | 0 | 0.30442 | 0 | 0.063327 | 0 | 252.5596 | 0 | 0.00060398 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
444 | 445 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFTeDA | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.87 | 9.86 | 0 | Baseline | 0 | 20550.4067 | 0 | 0 | 6331.6636 | 0 | 0 | 3.2457 | 78.871 | 6338.1637 | 0 | 0.99164 | 0 | 9.9013 | 0 | 0 | 0 | 9.7877 | 0 | 10.0105 | 0 | 0.22277 | 0 | 0.050467 | 0 | 790.4526 | 0 | 0.0020559 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
451 | 452 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFBS | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 3.5 | 3.5 | 0 | Manual | 0 | 44579.984 | 0 | 0 | 3843.2415 | 0 | 0 | 11.5996 | 16.0463 | 3849.2403 | 0 | 0.9945 | 0 | 3.8947 | 0 | 0 | 0 | 3.6304 | 0 | 4.2369 | 0 | 0.60649 | 0 | 0.1751 | 0 | 878.7391 | 0 | 0.00063228 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
453 | 454 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFPeS | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 6.38 | 3.5 | 0 | Manual | 0 | 44579.984 | 0 | 0 | 3843.2415 | 0 | 0 | 11.5996 | 12.8183 | 3849.2403 | 0 | 0.9945 | 0 | 3.8947 | 0 | 0 | 0 | 3.6304 | 0 | 4.2369 | 0 | 0.60649 | 0 | 0.1751 | 0 | 878.7391 | 0 | 0.00063228 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
455 | 456 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | L-PFHxS | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.34 | 7.35 | 0 | Baseline | 0 | 64589.1226 | 0 | 0 | 16032.0013 | 0 | 0 | 4.0288 | 72.7071 | 16047.8143 | 0 | 0.99123 | 0 | 7.3726 | 0 | 0 | 0 | 7.2433 | 0 | 7.6081 | 0 | 0.36483 | 0 | 0.062346 | 0 | 199.0083 | 0 | 6.1454e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
457 | 458 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | Br-PFHxS | 0 | Qualifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.16 | 7.35 | 0 | Baseline | 0 | 64589.1226 | 0 | 0 | 16032.0013 | 0 | 0 | 4.0288 | 72.7071 | 16047.8143 | 0 | 0.99123 | 0 | 7.3726 | 0 | 0 | 0 | 7.2433 | 0 | 7.6081 | 0 | 0.36483 | 0 | 0.062346 | 0 | 199.0083 | 0 | 6.1454e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
459 | 460 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFHpS | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.98 | 7.35 | 0 | Baseline | 0 | 64589.1226 | 0 | 0 | 16032.0013 | 0 | 0 | 4.0288 | 17.1261 | 16047.8143 | 0 | 0.99123 | 0 | 7.3726 | 0 | 0 | 0 | 7.2433 | 0 | 7.6081 | 0 | 0.36483 | 0 | 0.062346 | 0 | 199.0083 | 0 | 6.1454e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
461 | 462 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | L-PFOS | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.46 | 8.45 | 0 | Baseline | 0 | 41004.9219 | 0 | 0 | 8685.7986 | 0 | 0 | 4.7209 | 15.9577 | 8692.7516 | 0 | 0.99167 | 0 | 8.4963 | 0 | 0 | 0 | 8.3687 | 0 | 8.642 | 0 | 0.2733 | 0 | 0.073903 | 0 | 147.0201 | 0 | 0.0015624 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
463 | 464 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | Br-PFOS | 0 | Qualifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.3 | 8.45 | 0 | Baseline | 0 | 41004.9219 | 0 | 0 | 8685.7986 | 0 | 0 | 4.7209 | 15.9577 | 8692.7516 | 0 | 0.99167 | 0 | 8.4963 | 0 | 0 | 0 | 8.3687 | 0 | 8.642 | 0 | 0.2733 | 0 | 0.073903 | 0 | 147.0201 | 0 | 0.0015624 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
465 | 466 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFNS | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.85 | 8.45 | 0 | Baseline | 0 | 41004.9219 | 0 | 0 | 8685.7986 | 0 | 0 | 4.7209 | 7.9788 | 8692.7516 | 0 | 0.99167 | 0 | 8.4963 | 0 | 0 | 0 | 8.3687 | 0 | 8.642 | 0 | 0.2733 | 0 | 0.073903 | 0 | 147.0201 | 0 | 0.0015624 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
467 | 468 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | PFDS | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.19 | 8.45 | 0 | Baseline | 0 | 41004.9219 | 0 | 0 | 8685.7986 | 0 | 0 | 4.7209 | 7.9788 | 8692.7516 | 0 | 0.99167 | 0 | 8.4963 | 0 | 0 | 0 | 8.3687 | 0 | 8.642 | 0 | 0.2733 | 0 | 0.073903 | 0 | 147.0201 | 0 | 0.0015624 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
471 | 472 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FBSA | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 4.6 | 8.86 | 0 | Baseline | 0 | 76071.9689 | 0 | 0 | 16805.659 | 0 | 0 | 4.5266 | 7.9815 | 16808.6879 | 0 | 0.99743 | 0 | 9.0267 | 0 | 0 | 0 | 8.8358 | 0 | 9.2991 | 0 | 0.46338 | 0 | 0.069523 | 0 | 165.4855 | 0 | 0.00041104 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
473 | 474 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FPeSA | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7 | 8.86 | 0 | Baseline | 0 | 76071.9689 | 0 | 0 | 16805.659 | 0 | 0 | 4.5266 | 7.9788 | 16808.6879 | 0 | 0.99743 | 0 | 9.0267 | 0 | 0 | 0 | 8.8358 | 0 | 9.2991 | 0 | 0.46338 | 0 | 0.069523 | 0 | 165.4855 | 0 | 0.00041104 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
475 | 476 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FHxSA | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.86 | 8.86 | 0 | Baseline | 0 | 76071.9689 | 0 | 0 | 16805.659 | 0 | 0 | 4.5266 | 17.6577 | 16808.6879 | 0 | 0.99743 | 0 | 9.0267 | 0 | 0 | 0 | 8.8358 | 0 | 9.2991 | 0 | 0.46338 | 0 | 0.069523 | 0 | 165.4855 | 0 | 0.00041104 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
477 | 478 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FHpSA | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.58 | 8.86 | 0 | Baseline | 0 | 76071.9689 | 0 | 0 | 16805.659 | 0 | 0 | 4.5266 | 7.9788 | 16808.6879 | 0 | 0.99743 | 0 | 9.0267 | 0 | 0 | 0 | 8.8358 | 0 | 9.2991 | 0 | 0.46338 | 0 | 0.069523 | 0 | 165.4855 | 0 | 0.00041104 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
479 | 480 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | FOSA | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.88 | 8.86 | 0 | Baseline | 0 | 76071.9689 | 0 | 0 | 16805.659 | 0 | 0 | 4.5266 | 24.0366 | 16808.6879 | 0 | 0.99743 | 0 | 9.0267 | 0 | 0 | 0 | 8.8358 | 0 | 9.2991 | 0 | 0.46338 | 0 | 0.069523 | 0 | 165.4855 | 0 | 0.00041104 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
498 | 499 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 4:2 FTS | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 5.94 | 5.92 | 0 | Baseline | 0 | 18857.4612 | 0 | 0 | 3395.5536 | 0 | 0 | 5.5536 | 12.8183 | 3405.4392 | 0 | 0.98432 | 0 | 5.9741 | 0 | 0 | 0 | 5.8212 | 0 | 6.1206 | 0 | 0.29933 | 0 | 0.084893 | 0 | 154.5563 | 0 | 0.002142 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
500 | 501 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 6:2 FTS | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 7.92 | 7.92 | 0 | Baseline | 0 | 16653.2683 | 0 | 0 | 4000.0955 | 0 | 0 | 4.1632 | 7.9788 | 4000.0955 | 0 | 0.99853 | 0 | 7.9481 | 0 | 0 | 0 | 7.7316 | 0 | 8.2029 | 0 | 0.47135 | 0 | 0.062908 | 0 | 95.4295 | 0 | 0 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
502 | 503 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 8:2 FTS | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 8.88 | 8.89 | 0 | Baseline | 0 | 19237.4184 | 0 | 0 | 3942.5033 | 0 | 0 | 4.8795 | 7.1656 | 3942.5033 | 0 | 1 | 0 | 8.9048 | 0 | 0 | 0 | 8.6438 | 0 | 9.2035 | 0 | 0.55975 | 0 | 0.075425 | 0 | 138.9589 | 0 | 0 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
504 | 505 | PFAS CS0 0ng/ml | 4 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 5) | CS0 | 0 | 0 | Standard | 0 | 0 | 0 | 1 | 1 | 10:2 FTS | 0 | Quantifiers | 10/26/2023 19:57:28 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0 | 4 | 0 | 0 | 9.5 | 8.89 | 0 | Baseline | 0 | 19237.4184 | 0 | 0 | 3942.5033 | 0 | 0 | 4.8795 | 19.9471 | 3942.5033 | 0 | 1 | 0 | 8.9048 | 0 | 0 | 0 | 8.6438 | 0 | 9.2035 | 0 | 0.55975 | 0 | 0.075425 | 0 | 138.9589 | 0 | 0 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0 | 0
565 | 566 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFBA | 0 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 1.49 | 1.48 | 0 | Baseline | 0 | 107217.7308 | 0 | 0 | 16340.7542 | 0 | 0 | 6.5614 | 69.7563 | 16400.0835 | 0 | 0.98047 | 0 | 1.4658 | 0 | 0 | 0 | 1.3197 | 0 | 1.6592 | 0 | 0.33947 | 0 | 0.096977 | 0 | 246.1554 | 0 | 0.008842 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
567 | 568 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFPeA | 0 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 3.18 | 3.21 | 0 | Manual | 0 | 75980.5208 | 0 | 0 | 8588.6263 | 0 | 0 | 8.8466 | 69.5571 | 8606.0397 | 0 | 0.98906 | 0 | 3.2265 | 0 | 0 | 0 | 3.004 | 0 | 3.516 | 0 | 0.51198 | 0 | 0.13291 | 0 | 679.2359 | 0 | 0.0007844 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
569 | 570 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFHxA | 0.0189231127132091 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy, Integration Quality | 0 | 0 | 0.01 | 4 | 0.0025 | 0.33533 | 6.13 | 6.12 | Manual | Baseline | 683.5476 | 155213.0294 | 0.0044039 | 120.6719 | 27133.7022 | 0.0044473 | 5.6645 | 5.7203 | 125.2777 | 27143.6225 | 0.68031 | 0.99599 | 6.1368 | 6.147 | 0.11133 | 0.0068248 | 5.9895 | 5.9713 | 6.2413 | 6.3676 | 0.25178 | 0.39637 | 0.069012 | 0.087616 | 22.6948 | 260.1408 | 0.031302 | 0.00083627 | TRUE | 0.99834 | TRUE | [M-H]- | 189.2311 | 5.988 | 6.2393 | 0.25126 | 5.9924 | 6.2349 | 0.24244 | -15.0018 | 0.80497 | 0.5981 | 1.2625 | 0 | 29 | 9 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 683.5476 | No Control Sample | 203843.5852 | off | off | 683.5476 | 203843.5852 | off | 100 | off | 0.01 | off | 3.3533e-05 | off | 312.9754 | 8.2739 | 2.5895 | 268.9841 | 6.2969 | 1.6937 | 0 | 0 | 0 | 0 | 0 | 5983.6748 | FALSE | 0.0025 | 0.0044039
571 | 572 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFHpA | 0.0210455898212623 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.40411 | 7.26 | 7.27 | Baseline | Baseline | 922.6929 | 175213.492 | 0.0052661 | 249.6299 | 42248.3298 | 0.0059086 | 3.6962 | 4.1472 | 264.9834 | 42262.9276 | 0.69516 | 0.99659 | 7.288 | 7.2925 | 0.38584 | 0.028012 | 7.2024 | 7.1385 | 7.3684 | 7.456 | 0.166 | 0.31745 | 0.055824 | 0.063809 | 36.4258 | 224.8697 | 0.10344 | 0.00069427 | FALSE | 0.99938 | TRUE | [M-H]- | 210.4559 | 7.2284 | 7.3665 | 0.13812 | 7.2345 | 7.3475 | 0.11294 | 155.5572 | 1.0697 | 0.93442 | 1.4195 | 0 | 20 | 8 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 922.6929 | No Control Sample | 228328.7703 | off | off | 922.6929 | 228328.7703 | off | 100 | off | 0.01 | off | 4.0411e-05 | off | 362.9735 | 10.7517 | 3.9026 | 318.9794 | 0.94858 | 0.30258 | 0 | 0 | 0 | 0 | 0 | 20.4196 | FALSE | 0.0025 | 0.0052661
573 | 574 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFOA | 0.0199594095983386 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.4182 | 7.96 | 7.94 | Baseline | Baseline | 800.7449 | 165423.7252 | 0.0048406 | 189.7048 | 38819.7875 | 0.0048868 | 4.221 | 4.2613 | 202.8941 | 38841.9901 | 0.50851 | 0.99554 | 7.9798 | 7.9799 | 0.24892 | 0.019814 | 7.9012 | 7.826 | 8.0791 | 8.1182 | 0.17787 | 0.2922 | 0.064637 | 0.066551 | 25.0846 | 123.2805 | 0.12575 | 0.0010865 | FALSE | 0.99998 | TRUE | [M-H]- | 199.5941 | 7.9012 | 8.0515 | 0.15035 | 7.9147 | 8.0456 | 0.13092 | -134.1233 | 0.9083 | 0.8904 | 2.0206 | 0 | 14 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 800.7449 | No Control Sample | 191475.4056 | off | off | 800.7449 | 191475.4056 | off | 100 | off | 0.01 | off | 4.182e-05 | off | 412.9698 | 8.17 | 3.3739 | 368.9787 | 6.9761 | 2.574 | 0 | 0 | 0 | 0 | 0 | 13.2337 | FALSE | 0.0025 | 0.0048406
575 | 576 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFNA | 0.0192424388710375 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.38979 | 8.46 | 8.47 | Baseline | Baseline | 620.3308 | 138771.3897 | 0.0044702 | 145.3495 | 29428.9786 | 0.004939 | 4.2679 | 4.7155 | 155.2371 | 29464.1927 | 0.69617 | 0.99296 | 8.4858 | 8.4919 | 0.30499 | 0.025802 | 8.4116 | 8.3112 | 8.5811 | 8.638 | 0.16947 | 0.32682 | 0.067174 | 0.073824 | 37.1586 | 171.9004 | 0.031782 | 0.0021633 | FALSE | 0.99929 | TRUE | [M-H]- | 192.4244 | 8.4171 | 8.5545 | 0.13738 | 8.4243 | 8.5448 | 0.1205 | -27.2593 | 1.1737 | 1.3481 | 0.59121 | 0 | 10 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Fail | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 620.3308 | No Control Sample | 159143.3885 | off | off | 620.3308 | 159143.3885 | off | 100 | off | 0.01 | off | 3.8979e-05 | off | 462.9661 | 6.2546 | 2.8956 | 418.9764 | 7.7053 | 3.2283 | 0 | 0 | 0 | 0 | 0 | 9.8103 | FALSE | 0.0025 | 0.0044702
577 | 578 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFDA | 0.013283835229288 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.87 | 8.86 | Manual | Baseline | 422.226 | 128038.8338 | 0.0032976 | 88.1021 | 25870.9309 | 0.0034054 | 4.7925 | 4.9491 | 89.4449 | 25873.9254 | 0.91289 | 0.99722 | 8.9064 | 8.9046 | 0.40994 | 0.036362 | 8.8146 | 8.7108 | 8.9949 | 9.0605 | 0.18028 | 0.3497 | 0.077014 | 0.077029 | 61.3176 | 208.0488 | 0.010203 | 0.00021369 | TRUE | 1.0002 | TRUE | [M-H]- | 132.8384 | 8.8278 | 8.9827 | 0.15497 | 8.8355 | 8.9718 | 0.13631 | -4.9865 | 0.99226 | 0.93808 | 2.1563 | 0 | 11 | 5 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Fail | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 422.226 | No Control Sample | 0 | off | off | 422.226 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9635 | 6.6715 | 3.4222 | 468.9773 | 15.4806 | 7.2599 | 0 | 0 | 0 | 0 | 0 | 4616.9634 | FALSE | 0.0025 | 0.0032976
579 | 580 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFUdA | 0.0140625003761255 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.21 | 9.2 | Manual | Baseline | 316.2521 | 101408.5655 | 0.0031186 | 86.8318 | 22365.2168 | 0.0038824 | 3.6421 | 4.5342 | 88.372 | 22367.9198 | 0.7614 | 0.99823 | 9.2491 | 9.2381 | 0.42415 | 0.039064 | 9.1734 | 9.0803 | 9.315 | 9.4675 | 0.14162 | 0.38718 | 0.059049 | 0.071494 | 28.6003 | 213.0162 | 0.0051143 | 0.00029761 | TRUE | 1.0012 | TRUE | [M-H]- | 140.625 | 9.1839 | 9.3065 | 0.12252 | 9.1919 | 9.3019 | 0.10999 | -3.1357 | 0.94705 | 0.93977 | 1.7493 | 0 | 9 | 5 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Fail | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 316.2521 | No Control Sample | 0 | off | off | 316.2521 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9584 | 2.7239 | 1.5334 | 518.9687 | 2.7332 | 1.4185 | 0 | 0 | 0 | 0 | 0 | 11.9735 | FALSE | 0.0025 | 0.0031186
581 | 582 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFDoA | 0.0135426608187828 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.47 | 9.47 | Manual | Baseline | 272.9378 | 77881.123 | 0.0035045 | 63.7861 | 19204.7981 | 0.0033214 | 4.279 | 4.0553 | 64.9296 | 19218.8001 | 0.75017 | 0.99422 | 9.5184 | 9.5062 | 0.51122 | 0.048413 | 9.4318 | 9.3861 | 9.5871 | 9.6158 | 0.15523 | 0.22962 | 0.069211 | 0.063474 | 24.9445 | 305.9297 | 0.0051152 | 0.0013186 | TRUE | 1.0013 | TRUE | [M-H]- | 135.4266 | 9.4431 | 9.5801 | 0.13698 | 9.4524 | 9.5763 | 0.12387 | 2.1019 | 1.0212 | 1.1448 | 0.86354 | 0 | 12 | 7 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 272.9378 | No Control Sample | 0 | off | off | 272.9378 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9563 | 4.3157 | 2.6453 | 568.9557 | -13.4524 | -7.6539 | 0 | 0 | 0 | 0 | 0 | 24738.4141 | FALSE | 0.0025 | 0.0035045
583 | 584 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFTrDA | 0.0135391140982548 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.68 | 9.47 | Manual | Baseline | 198.9002 | 77881.123 | 0.0025539 | 61.1496 | 19204.7981 | 0.0031841 | 3.2527 | 4.0553 | 62.3892 | 19218.8001 | 0.63656 | 0.99422 | 9.7269 | 9.5062 | 0.48458 | 0.046907 | 9.6678 | 9.3861 | 9.786 | 9.6158 | 0.11824 | 0.22962 | 0.050902 | 0.063474 | 32.0928 | 305.9297 | 0.010254 | 0.0013186 | TRUE | 1.0232 | TRUE | [M-H]- | 135.3911 | 9.6724 | 9.7801 | 0.10769 | 9.6769 | 9.776 | 0.099171 | -5.303 | 0.94141 | 0.88023 | 0 | 0 | 13 | 6 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 198.9002 | No Control Sample | 0 | off | off | 198.9002 | 0 | off | 0 | off | 0 | off | 0 | off | 662.944 | -9.7074 | -6.4355 | 618.9602 | -2.3689 | -1.4662 | 0 | 0 | 0 | 0 | 0 | 14.1366 | FALSE | 0.0025 | 0.0025539
585 | 586 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFTeDA | 0.0255255135862019 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.87 | 9.86 | Manual | Baseline | 131.5545 | 14536.4205 | 0.00905 | 45.9897 | 4599.1726 | 0.0099996 | 2.8605 | 3.1607 | 47.1039 | 4608.407 | 0.35603 | 0.98713 | 9.896 | 9.8957 | 0.26329 | 0.025987 | 9.8163 | 9.7863 | 9.9542 | 9.9801 | 0.13794 | 0.1938 | 0.034177 | 0.049428 | 26.8864 | 727.5148 | 0.015441 | 0.0034819 | TRUE | 1 | TRUE | [M-H]- | 255.2551 | 9.8264 | 9.9501 | 0.12376 | 9.8344 | 9.9434 | 0.10892 | 5.1481 | 0.90474 | 0.80477 | 0 | 0 | 16 | 5 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 131.5545 | No Control Sample | 0 | off | off | 131.5545 | 0 | off | 0 | off | 0 | off | 0 | off | 713.1883 | 338.1037 | 241.0501 | 668.9575 | -0.56602 | -0.37865 | 0 | 0 | 0 | 0 | 0 | inf | FALSE | 0.0025 | 0.00905
592 | 593 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFBS | 0 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 3.5 | 3.5 | 0 | Manual | 0 | 42964.0261 | 0 | 0 | 4018.0276 | 0 | 0 | 10.6928 | 35.3207 | 4029.9339 | 0 | 0.9855 | 0 | 3.872 | 0 | 0 | 0 | 3.6327 | 0 | 4.1296 | 0 | 0.49691 | 0 | 0.15757 | 0 | 785.3172 | 0 | 0 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
594 | 595 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFPeS | 0 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 6.38 | 3.5 | 0 | Manual | 0 | 42964.0261 | 0 | 0 | 4018.0276 | 0 | 0 | 10.6928 | 28.8646 | 4029.9339 | 0 | 0.9855 | 0 | 3.872 | 0 | 0 | 0 | 3.6327 | 0 | 4.1296 | 0 | 0.49691 | 0 | 0.15757 | 0 | 785.3172 | 0 | 0 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
596 | 597 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | L-PFHxS | 0.0194646807889381 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.00811 | 4 | 0.0020275 | 0 | 7.34 | 7.35 | Manual | Baseline | 225.8193 | 62108.3868 | 0.0036359 | 84.5479 | 15176.9615 | 0.0055708 | 2.6709 | 4.0923 | 85.6281 | 15180.8482 | 0.71952 | 0.99425 | 7.3657 | 7.367 | 0.35033 | 0.025714 | 7.2957 | 7.2248 | 7.4306 | 7.5629 | 0.13488 | 0.3381 | 0.034622 | 0.064349 | 51.0281 | 196.0244 | 0.0050893 | 0.00024807 | TRUE | 0.99982 | TRUE | [M-H]- | 240.0084 | 7.2947 | 7.4217 | 0.12708 | 7.3148 | 7.4162 | 0.10143 | -3.1901 | 0.9467 | 1.1578 | 9.7952 | 0 | 18 | 6 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 225.8193 | No Control Sample | 0 | off | off | 225.8193 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9381 | 3.7825 | 1.509 | 79.9571 | -1.7522 | -0.1401 | 0 | 0 | 0 | 0 | 0 | 34.8001 | FALSE | 0.0020275 | 0.0036359
598 | 599 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | Br-PFHxS | 0 | Qualifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0.00189 | 4 | 0.0004725 | 0 | 7.16 | 7.35 | 0 | Baseline | 0 | 62108.3868 | 0 | 0 | 15176.9615 | 0 | 0 | 4.0923 | 85.6281 | 15180.8482 | 0 | 0.99425 | 0 | 7.367 | 0 | 0 | 0 | 7.2248 | 0 | 7.5629 | 0 | 0.3381 | 0 | 0.064349 | 0 | 196.0244 | 0 | 0.00024807 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 225.8193 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0004725 | 0
600 | 601 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFHpS | 0 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7.98 | 7.35 | 0 | Baseline | 0 | 62108.3868 | 0 | 0 | 15176.9615 | 0 | 0 | 4.0923 | 20.7971 | 15180.8482 | 0 | 0.99425 | 0 | 7.367 | 0 | 0 | 0 | 7.2248 | 0 | 7.5629 | 0 | 0.3381 | 0 | 0.064349 | 0 | 196.0244 | 0 | 0.00024807 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
602 | 603 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | L-PFOS | 0 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0.00788 | 4 | 0.00197 | 0 | 8.46 | 8.45 | 0 | Baseline | 0 | 39946.3381 | 0 | 0 | 8795.7032 | 0 | 0 | 4.5416 | 17.6604 | 8802.003 | 0 | 0.99246 | 0 | 8.4906 | 0 | 0 | 0 | 8.3125 | 0 | 8.6938 | 0 | 0.38137 | 0 | 0.070423 | 0 | 189.8646 | 0 | 0.0015213 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.00197 | 0
604 | 605 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | Br-PFOS | 0 | Qualifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0.00211 | 4 | 0.0005275 | 0 | 8.3 | 8.45 | 0 | Baseline | 0 | 39946.3381 | 0 | 0 | 8795.7032 | 0 | 0 | 4.5416 | 17.6604 | 8802.003 | 0 | 0.99246 | 0 | 8.4906 | 0 | 0 | 0 | 8.3125 | 0 | 8.6938 | 0 | 0.38137 | 0 | 0.070423 | 0 | 189.8646 | 0 | 0.0015213 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0005275 | 0
606 | 607 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFNS | 0 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.85 | 8.45 | 0 | Baseline | 0 | 39946.3381 | 0 | 0 | 8795.7032 | 0 | 0 | 4.5416 | 37.8346 | 8802.003 | 0 | 0.99246 | 0 | 8.4906 | 0 | 0 | 0 | 8.3125 | 0 | 8.6938 | 0 | 0.38137 | 0 | 0.070423 | 0 | 189.8646 | 0 | 0.0015213 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
608 | 609 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFDS | 0 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.19 | 8.45 | 0 | Baseline | 0 | 39946.3381 | 0 | 0 | 8795.7032 | 0 | 0 | 4.5416 | 12.9069 | 8802.003 | 0 | 0.99246 | 0 | 8.4906 | 0 | 0 | 0 | 8.3125 | 0 | 8.6938 | 0 | 0.38137 | 0 | 0.070423 | 0 | 189.8646 | 0 | 0.0015213 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
612 | 613 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FBSA | 0 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 4.6 | 8.86 | 0 | Baseline | 0 | 72686.8787 | 0 | 0 | 15874.5221 | 0 | 0 | 4.5788 | 29.8558 | 15877.1621 | 0 | 0.99826 | 0 | 9.0235 | 0 | 0 | 0 | 8.8139 | 0 | 9.2986 | 0 | 0.48468 | 0 | 0.070645 | 0 | 124.3625 | 0 | 0.00036746 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
614 | 615 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FPeSA | 0.00743992507804104 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7 | 8.86 | Manual | Baseline | 221.4338 | 72686.8787 | 0.0030464 | 55.8626 | 15874.5221 | 0.003519 | 3.9639 | 4.5788 | 56.5828 | 15877.1621 | 0.78704 | 0.99826 | 7.1005 | 9.0235 | 1.4354 | 0.10048 | 7.045 | 8.8139 | 7.1948 | 9.2986 | 0.14983 | 0.48468 | 0.074566 | 0.070645 | 32.2234 | 124.3625 | 0.01018 | 0.00036746 | TRUE | 0.78689 | TRUE | [M-H]- | 74.3993 | 7.0475 | 7.1876 | 0.14008 | 7.0527 | 7.18 | 0.12729 | -3.7955 | 2.1616 | 3.6718 | 3.9576 | 0 | 18 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 221.4338 | No Control Sample | 0 | off | off | 221.4338 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9594 | 10.4709 | 3.6434 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.6504 | FALSE | 0.0025 | 0.0030464
616 | 617 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FHxSA | 0.0123845064511062 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7.86 | 8.86 | Manual | Baseline | 183.9388 | 72686.8787 | 0.0025306 | 54.484 | 15874.5221 | 0.0034322 | 3.376 | 4.5788 | 55.2864 | 15877.1621 | 0.67297 | 0.99826 | 8.0054 | 9.0235 | 1.85 | 0.14541 | 7.9503 | 8.8139 | 8.0659 | 9.2986 | 0.11557 | 0.48468 | 0.053466 | 0.070645 | 22.9121 | 124.3625 | 0.010198 | 0.00036746 | TRUE | 0.88717 | TRUE | [M-H]- | 123.8451 | 7.953 | 8.0608 | 0.10781 | 7.9609 | 8.055 | 0.094104 | -4.8078 | 0.876 | 0.75462 | 8.6269 | 0 | 10 | 6 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 183.9388 | No Control Sample | 0 | off | off | 183.9388 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9541 | 3.8933 | 1.5494 | 77.9661 | 9.1856 | 0.71616 | 0 | 0 | 0 | 0 | 0 | 16.7007 | FALSE | 0.0025 | 0.0025306
618 | 619 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FHpSA | 0.0160414485625542 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.58 | 8.86 | Baseline | Baseline | 493.4872 | 72686.8787 | 0.0067892 | 102.3805 | 15874.5221 | 0.0064494 | 4.8201 | 4.5788 | 102.3805 | 15877.1621 | 1 | 0.99826 | 8.5728 | 9.0235 | 0.084148 | 0.0072199 | 8.3394 | 8.8139 | 8.7821 | 9.2986 | 0.4427 | 0.48468 | 0.07238 | 0.070645 | 94.2754 | 124.3625 | 0 | 0.00036746 | FALSE | 0.95005 | TRUE | [M-H]- | 160.4145 | 8.479 | 8.653 | 0.17403 | 8.4992 | 8.6429 | 0.14375 | 0 | 1.0054 | 1.1674 | 2.8938 | 0 | 25 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 493.4872 | No Control Sample | 0 | off | off | 493.4872 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9505 | 2.5605 | 1.147 | 77.9648 | -7.289 | -0.56829 | 0 | 0 | 0 | 0 | 0 | 8.8136 | FALSE | 0.0025 | 0.0067892
620 | 621 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FOSA | 0.0199977356036385 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.88 | 8.86 | Manual | Baseline | 330.7366 | 72686.8787 | 0.0045502 | 71.4448 | 15874.5221 | 0.0045006 | 4.6293 | 4.5788 | 72.1702 | 15877.1621 | 0.80488 | 0.99826 | 9.0191 | 9.0235 | 1.5663 | 0.13909 | 8.9204 | 8.8139 | 9.0953 | 9.2986 | 0.17493 | 0.48468 | 0.073756 | 0.070645 | 25.7229 | 124.3625 | 0 | 0.00036746 | TRUE | 0.99951 | TRUE | [M-H]- | 199.9774 | 8.9337 | 9.0862 | 0.15257 | 8.9415 | 9.0779 | 0.13638 | 0 | 0.9102 | 0.79517 | 3.7528 | 0 | 11 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 330.7366 | No Control Sample | 0 | off | off | 330.7366 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9472 | 2.0569 | 1.0242 | 77.9656 | 2.193 | 0.17098 | 0 | 0 | 0 | 0 | 0 | 9.8952 | FALSE | 0.0025 | 0.0045502
639 | 640 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 4:2 FTS | 0 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 5.94 | 5.92 | 0 | Baseline | 0 | 18479.0831 | 0 | 0 | 3170.9578 | 0 | 0 | 5.8276 | 35.4926 | 3170.9578 | 0 | 0.99599 | 0 | 5.9522 | 0 | 0 | 0 | 5.7496 | 0 | 6.1494 | 0 | 0.39975 | 0 | 0.092177 | 0 | 203.8611 | 0 | 7.6817e-11 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
641 | 642 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 6:2 FTS | 0.0110346599021657 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7.92 | 7.92 | Manual | Baseline | 113.2436 | 16914.6219 | 0.006695 | 26.4713 | 4191.0725 | 0.0063161 | 4.278 | 4.0359 | 26.8103 | 4193.0057 | 0.96702 | 0.995 | 7.9326 | 7.9423 | 0.15869 | 0.012568 | 7.8474 | 7.8286 | 8.0091 | 8.0744 | 0.16167 | 0.24572 | 0.070829 | 0.061777 | 60.2218 | 108.0431 | 0.015268 | 0.00085934 | TRUE | 0.99877 | TRUE | [M-H]- | 110.3466 | 7.8532 | 8.0007 | 0.1475 | 7.858 | 7.997 | 0.13892 | 2.5 | 0.97138 | 0.95307 | 3.6323 | 0 | 15 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 113.2436 | No Control Sample | 0 | off | off | 113.2436 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9691 | 2.9001 | 1.2382 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 302.6834 | FALSE | 0.0025 | 0.006695
643 | 644 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 8:2 FTS | 0 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.88 | 8.89 | 0 | Baseline | 0 | 19268.3378 | 0 | 0 | 3913.5426 | 0 | 0 | 4.9235 | 46.7098 | 3919.1333 | 0 | 0.99086 | 0 | 8.8997 | 0 | 0 | 0 | 8.7537 | 0 | 9.0292 | 0 | 0.27545 | 0 | 0.077112 | 0 | 172.451 | 0 | 0.0026716 | TRUE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
645 | 646 | PFAS CS1 0.01ng/ml | 5 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 6) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 10:2 FTS | 0 | Quantifiers | 10/26/2023 20:14:03 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.5 | 8.89 | 0 | Baseline | 0 | 19268.3378 | 0 | 0 | 3913.5426 | 0 | 0 | 4.9235 | 19.9471 | 3919.1333 | 0 | 0.99086 | 0 | 8.8997 | 0 | 0 | 0 | 8.7537 | 0 | 9.0292 | 0 | 0.27545 | 0 | 0.077112 | 0 | 172.451 | 0 | 0.0026716 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
706 | 707 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFBA | 0 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 1.49 | 1.48 | 0 | Baseline | 0 | 106321.1967 | 0 | 0 | 15284.3754 | 0 | 0 | 6.9562 | 108.2302 | 15337.2583 | 0 | 0.98237 | 0 | 1.4705 | 0 | 0 | 0 | 1.339 | 0 | 1.684 | 0 | 0.34497 | 0 | 0.10807 | 0 | 234.9661 | 0 | 0.0060809 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
708 | 709 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFPeA | 0 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 3.18 | 3.21 | 0 | Baseline | 0 | 82574.9492 | 0 | 0 | 8760.4641 | 0 | 0 | 9.4259 | 78.9945 | 8785.2438 | 0 | 0.98575 | 0 | 3.2029 | 0 | 0 | 0 | 2.9646 | 0 | 3.468 | 0 | 0.50338 | 0 | 0.1426 | 0 | 598.7937 | 0 | 0.0044228 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
710 | 711 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFHxA | 0.025628528703879 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy, Integration Quality, Asymmetry, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.45124 | 6.13 | 6.12 | Baseline | Baseline | 919.8289 | 154218.1035 | 0.0059645 | 156.0187 | 28031.3529 | 0.0055659 | 5.8956 | 5.5016 | 161.1211 | 28069.4266 | 0.74727 | 0.99133 | 6.142 | 6.1411 | 0.1959 | 0.012008 | 6.0228 | 5.955 | 6.2298 | 6.3123 | 0.20703 | 0.35728 | 0.097657 | 0.085698 | 32.311 | 242.6437 | 0.047079 | 0.0025594 | FALSE | 1.0002 | TRUE | [M-H]- | 256.2853 | 6.0466 | 6.2207 | 0.17406 | 6.0549 | 6.2139 | 0.15894 | 35.4783 | 0.74148 | 0.45768 | 0.90141 | 0 | 23 | 12 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 919.8289 | No Control Sample | 203843.5852 | off | off | 919.8289 | 203843.5852 | off | 100 | off | 0.01 | off | 4.5124e-05 | off | 312.9744 | 5.1608 | 1.6152 | 268.9825 | 0.25869 | 0.069583 | 0 | 0 | 0 | 0 | 0 | 11049.1474 | FALSE | 0.0025 | 0.0059645
712 | 713 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFHpA | 0.0278114625687662 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.53396 | 7.26 | 7.27 | Baseline | Baseline | 1219.1839 | 175193.0455 | 0.0069591 | 293.7187 | 42988.154 | 0.0068326 | 4.1509 | 4.0754 | 297.197 | 42994.7731 | 0.79411 | 0.99819 | 7.2927 | 7.2946 | 0.45031 | 0.032692 | 7.1943 | 7.1399 | 7.387 | 7.4839 | 0.19267 | 0.34398 | 0.06492 | 0.062762 | 57.5426 | 250.3242 | 0.018789 | 3.552e-05 | FALSE | 0.99973 | TRUE | [M-H]- | 278.1146 | 7.221 | 7.3658 | 0.14481 | 7.2348 | 7.354 | 0.11919 | 28.6436 | 0.9871 | 1.003 | 1.2527 | 0 | 23 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1219.1839 | No Control Sample | 228328.7703 | off | off | 1219.1839 | 228328.7703 | off | 100 | off | 0.01 | off | 5.3396e-05 | off | 362.9725 | 7.9541 | 2.8871 | 318.983 | 12.32 | 3.9298 | 0 | 0 | 0 | 0 | 0 | 39.2237 | FALSE | 0.0025 | 0.0069591
714 | 715 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFOA | 0.0259571920850602 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.56199 | 7.96 | 7.94 | Baseline | Baseline | 1076.0765 | 170937.2545 | 0.0062952 | 227.6739 | 39792.5278 | 0.0057215 | 4.7264 | 4.2957 | 237.6736 | 39803.3959 | 0.53864 | 0.99707 | 7.9856 | 7.9825 | 0.32196 | 0.025628 | 7.8744 | 7.8264 | 8.1115 | 8.1666 | 0.23708 | 0.3402 | 0.072508 | 0.066357 | 25.8502 | 98.2033 | 0.041007 | 0.00060369 | FALSE | 1.0004 | TRUE | [M-H]- | 259.5719 | 7.8975 | 8.0607 | 0.16321 | 7.9109 | 8.0466 | 0.13574 | -39.379 | 0.93904 | 0.84655 | 1.3911 | 0 | 18 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1076.0765 | No Control Sample | 191475.4056 | off | off | 1076.0765 | 191475.4056 | off | 100 | off | 0.01 | off | 5.6199e-05 | off | 412.9689 | 5.9096 | 2.4405 | 368.9749 | -3.2056 | -1.1828 | 0 | 0 | 0 | 0 | 0 | 32.1145 | FALSE | 0.0025 | 0.0062952
716 | 717 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFNA | 0.0109021571148673 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.23055 | 8.46 | 8.47 | Manual | Baseline | 366.9069 | 144870.5568 | 0.0025327 | 71.5575 | 30673.1296 | 0.0023329 | 5.1274 | 4.723 | 72.4389 | 30680.5491 | 0.71785 | 0.99795 | 8.491 | 8.4944 | 0.36631 | 0.03099 | 8.3965 | 8.2957 | 8.5686 | 8.7098 | 0.1721 | 0.41403 | 0.083519 | 0.073657 | 18.7451 | 168.7968 | 0.005087 | 0.00051681 | TRUE | 0.9996 | TRUE | [M-H]- | 109.0216 | 8.4051 | 8.5617 | 0.15661 | 8.4149 | 8.5564 | 0.1415 | -2.1151 | 0.85956 | 0.74134 | 1.7209 | 0 | 11 | 7 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Fail | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 366.9069 | No Control Sample | 159143.3885 | off | off | 366.9069 | 159143.3885 | off | 100 | off | 0.01 | off | 2.3055e-05 | off | 462.9645 | 2.8051 | 1.2987 | 418.977 | 9.0698 | 3.8 | 0 | 0 | 0 | 0 | 0 | 14.6459 | FALSE | 0.0025 | 0.0025327
718 | 719 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFDA | 0.0150263403626164 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.87 | 8.86 | Baseline | Baseline | 502.0781 | 134597.911 | 0.0037302 | 117.4776 | 27611.0729 | 0.0042547 | 4.2738 | 4.8748 | 117.524 | 27628.6878 | 0.95528 | 0.99631 | 8.8911 | 8.9073 | 0.2381 | 0.021119 | 8.7602 | 8.7306 | 9.0158 | 9.0427 | 0.25552 | 0.31203 | 0.067541 | 0.075897 | 53.3633 | 199.6694 | 0.00080083 | 0.0010005 | FALSE | 0.99818 | TRUE | [M-H]- | 150.2634 | 8.8249 | 8.9661 | 0.14121 | 8.8309 | 8.9568 | 0.12588 | 0.36819 | 1.1589 | 1.2897 | 0.8265 | 0 | 13 | 5 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 502.0781 | No Control Sample | 0 | off | off | 502.0781 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9639 | 7.5423 | 3.8689 | 468.9708 | 1.6867 | 0.79102 | 0 | 0 | 0 | 0 | 0 | 5807.0095 | FALSE | 0.0025 | 0.0037302
720 | 721 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFUdA | 0.0110212686884649 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.21 | 9.2 | Manual | Baseline | 256.4652 | 104930.2366 | 0.0024441 | 46.9197 | 23280.103 | 0.0020154 | 5.466 | 4.5073 | 47.771 | 23283.9102 | 0.76328 | 0.99781 | 9.2438 | 9.239 | 0.36686 | 0.033788 | 9.1727 | 9.1004 | 9.3804 | 9.4515 | 0.2078 | 0.35117 | 0.073103 | 0.070997 | 18.3206 | 215.4474 | 0.010233 | 0.00039048 | TRUE | 1.0005 | TRUE | [M-H]- | 110.2127 | 9.1817 | 9.3729 | 0.19121 | 9.1904 | 9.3653 | 0.17494 | 2.3105 | 2.0039 | 3.4843 | 1.8936 | 0 | 13 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Fail | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 256.4652 | No Control Sample | 0 | off | off | 256.4652 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9569 | 0.037131 | 0.020903 | 518.9789 | 22.2619 | 11.5532 | 0 | 0 | 0 | 0 | 0 | 50.606 | FALSE | 0.0025 | 0.0024441
722 | 723 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFDoA | 0.00996958785350627 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.47 | 9.47 | Manual | Baseline | 213.93 | 82921.4798 | 0.0025799 | 67.7826 | 19979.4713 | 0.0033926 | 3.1561 | 4.1503 | 69.393 | 20003.1933 | 0.53115 | 0.99349 | 9.5252 | 9.5086 | 0.58272 | 0.055184 | 9.4357 | 9.3866 | 9.5806 | 9.6163 | 0.14493 | 0.2297 | 0.04639 | 0.065125 | 22.8186 | 258.1962 | 0.0051445 | 0.0020872 | TRUE | 1.0017 | TRUE | [M-H]- | 99.6959 | 9.4471 | 9.5739 | 0.12687 | 9.46 | 9.5673 | 0.10733 | 2.4061 | 0.81066 | 0.6423 | 1.1933 | 0 | 11 | 5 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 213.93 | No Control Sample | 0 | off | off | 213.93 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9611 | 12.1331 | 7.4371 | 568.9645 | 1.9376 | 1.1025 | 0 | 0 | 0 | 0 | 0 | 26917.4847 | FALSE | 0.0025 | 0.0025799
724 | 725 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFTrDA | 0.0104707836471025 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.68 | 9.47 | Manual | Baseline | 163.7793 | 82921.4798 | 0.0019751 | 32.9145 | 19979.4713 | 0.0016474 | 4.9759 | 4.1503 | 33.4182 | 20003.1933 | 0.6848 | 0.99349 | 9.7318 | 9.5086 | 0.53527 | 0.051815 | 9.6257 | 9.3866 | 9.7891 | 9.6163 | 0.16346 | 0.2297 | 0.077137 | 0.065125 | 18.6501 | 258.1962 | 0 | 0.0020872 | TRUE | 1.0235 | TRUE | [M-H]- | 104.7078 | 9.6313 | 9.7842 | 0.15294 | 9.6377 | 9.7819 | 0.14425 | 0 | 0.87586 | 0.78291 | 0 | 0 | 16 | 9 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 163.7793 | No Control Sample | 0 | off | off | 163.7793 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9544 | 5.9746 | 3.9609 | 619.2686 | 495.7815 | 306.8698 | 0 | 0 | 0 | 0 | 0 | 14.1366 | FALSE | 0.0025 | 0.0019751
726 | 727 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFTeDA | 0.0164198588624939 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.87 | 9.86 | Manual | Baseline | 97.569 | 16759.8033 | 0.0058216 | 39.2325 | 5018.0467 | 0.0078183 | 2.4869 | 3.3399 | 52.3484 | 5028.0244 | 0.23321 | 0.98807 | 9.8993 | 9.9001 | 0.29732 | 0.029345 | 9.8633 | 9.7867 | 9.9477 | 9.9966 | 0.084373 | 0.20995 | 0.037973 | 0.051995 | 26.2395 | 738.0837 | 0.0067051 | 0.003737 | TRUE | 0.99992 | TRUE | [M-H]- | 164.1986 | 9.8657 | 9.9546 | 0.088978 | 9.8679 | 9.9543 | 0.086397 | -3.1178 | 1.5024 | 2.1577 | 0 | 0 | 11 | 6 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 97.569 | No Control Sample | 0 | off | off | 97.569 | 0 | off | 0 | off | 0 | off | 0 | off | 713.1909 | 341.7843 | 243.6742 | 668.9569 | -1.5605 | -1.0439 | 0 | 0 | 0 | 0 | 0 | inf | FALSE | 0.0025 | 0.0058216
733 | 734 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFBS | 0 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 3.5 | 3.5 | 0 | Baseline | 0 | 41938.1957 | 0 | 0 | 3941.0256 | 0 | 0 | 10.6414 | 35.775 | 4004.9452 | 0 | 0.94961 | 0 | 3.8432 | 0 | 0 | 0 | 3.6393 | 0 | 4.0596 | 0 | 0.42032 | 0 | 0.17253 | 0 | 768.3159 | 0 | 0.010683 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
735 | 736 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFPeS | 0 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 6.38 | 3.5 | 0 | Baseline | 0 | 41938.1957 | 0 | 0 | 3941.0256 | 0 | 0 | 10.6414 | 42.3172 | 4004.9452 | 0 | 0.94961 | 0 | 3.8432 | 0 | 0 | 0 | 3.6393 | 0 | 4.0596 | 0 | 0.42032 | 0 | 0.17253 | 0 | 768.3159 | 0 | 0.010683 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
737 | 738 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | L-PFHxS | 0.0190246172816249 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.00811 | 4 | 0.0020275 | 0 | 7.34 | 7.35 | Manual | Baseline | 225.3473 | 63412.1961 | 0.0035537 | 59.2464 | 15602.0351 | 0.0037974 | 3.8036 | 4.0644 | 60.7031 | 15631.5879 | 0.69548 | 0.98771 | 7.3832 | 7.3689 | 0.58801 | 0.04316 | 7.2916 | 7.2634 | 7.4365 | 7.507 | 0.145 | 0.2436 | 0.052589 | 0.063481 | 31.9397 | 165.8248 | 0.010297 | 0.0021115 | TRUE | 1.0019 | TRUE | [M-H]- | 234.5822 | 7.2977 | 7.4352 | 0.13749 | 7.3149 | 7.431 | 0.1161 | -4.2075 | 0.83205 | 0.77502 | 10.0515 | 0 | 19 | 8 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 225.3473 | No Control Sample | 0 | off | off | 225.3473 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9409 | 10.6642 | 4.2544 | 79.9566 | -7.0152 | -0.56092 | 0 | 0 | 0 | 0 | 0 | 1118.7845 | FALSE | 0.0020275 | 0.0035537
739 | 740 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | Br-PFHxS | 0 | Qualifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0.00189 | 4 | 0.0004725 | 0 | 7.16 | 7.35 | 0 | Baseline | 0 | 63412.1961 | 0 | 0 | 15602.0351 | 0 | 0 | 4.0644 | 60.7031 | 15631.5879 | 0 | 0.98771 | 0 | 7.3689 | 0 | 0 | 0 | 7.2634 | 0 | 7.507 | 0 | 0.2436 | 0 | 0.063481 | 0 | 165.8248 | 0 | 0.0021115 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 225.3473 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0004725 | 0
741 | 742 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFHpS | 0 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7.98 | 7.35 | 0 | Baseline | 0 | 63412.1961 | 0 | 0 | 15602.0351 | 0 | 0 | 4.0644 | 19.9973 | 15631.5879 | 0 | 0.98771 | 0 | 7.3689 | 0 | 0 | 0 | 7.2634 | 0 | 7.507 | 0 | 0.2436 | 0 | 0.063481 | 0 | 165.8248 | 0 | 0.0021115 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
743 | 744 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | L-PFOS | 0 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0.00788 | 4 | 0.00197 | 0 | 8.46 | 8.45 | 0 | Baseline | 0 | 40851.8543 | 0 | 0 | 8903.0781 | 0 | 0 | 4.5885 | 50.1354 | 8911.7316 | 0 | 0.98989 | 0 | 8.4937 | 0 | 0 | 0 | 8.3493 | 0 | 8.6254 | 0 | 0.27608 | 0 | 0.071921 | 0 | 132.6522 | 0 | 0.0018963 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.00197 | 0
745 | 746 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | Br-PFOS | 0 | Qualifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0.00211 | 4 | 0.0005275 | 0 | 8.3 | 8.45 | 0 | Baseline | 0 | 40851.8543 | 0 | 0 | 8903.0781 | 0 | 0 | 4.5885 | 50.1354 | 8911.7316 | 0 | 0.98989 | 0 | 8.4937 | 0 | 0 | 0 | 8.3493 | 0 | 8.6254 | 0 | 0.27608 | 0 | 0.071921 | 0 | 132.6522 | 0 | 0.0018963 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0005275 | 0
747 | 748 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFNS | 0 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.85 | 8.45 | 0 | Baseline | 0 | 40851.8543 | 0 | 0 | 8903.0781 | 0 | 0 | 4.5885 | 25.105 | 8911.7316 | 0 | 0.98989 | 0 | 8.4937 | 0 | 0 | 0 | 8.3493 | 0 | 8.6254 | 0 | 0.27608 | 0 | 0.071921 | 0 | 132.6522 | 0 | 0.0018963 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
749 | 750 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFDS | 0 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.19 | 8.45 | 0 | Baseline | 0 | 40851.8543 | 0 | 0 | 8903.0781 | 0 | 0 | 4.5885 | 17.7463 | 8911.7316 | 0 | 0.98989 | 0 | 8.4937 | 0 | 0 | 0 | 8.3493 | 0 | 8.6254 | 0 | 0.27608 | 0 | 0.071921 | 0 | 132.6522 | 0 | 0.0018963 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
753 | 754 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FBSA | 0 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 4.6 | 8.86 | 0 | Baseline | 0 | 72909.1064 | 0 | 0 | 15893.7253 | 0 | 0 | 4.5873 | 31.2042 | 15919.4385 | 0 | 0.99017 | 0 | 9.0264 | 0 | 0 | 0 | 8.8566 | 0 | 9.1612 | 0 | 0.30458 | 0 | 0.071015 | 0 | 172.6057 | 0 | 0.0027659 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
755 | 756 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FPeSA | 0.0103117316951285 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7 | 8.86 | Manual | Baseline | 307.8454 | 72909.1064 | 0.0042223 | 57.7815 | 15893.7253 | 0.0036355 | 5.3278 | 4.5873 | 58.2046 | 15919.4385 | 0.91214 | 0.99017 | 7.1134 | 9.0264 | 1.6197 | 0.11338 | 7.0074 | 8.8566 | 7.1922 | 9.1612 | 0.18481 | 0.30458 | 0.08181 | 0.071015 | 38.3042 | 172.6057 | 0.0050619 | 0.0027659 | TRUE | 0.78807 | TRUE | [M-H]- | 103.1173 | 7.011 | 7.1843 | 0.17327 | 7.0189 | 7.1807 | 0.16189 | -1.5827 | 0.87823 | 0.78264 | 2.0039 | 0 | 21 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 307.8454 | No Control Sample | 0 | off | off | 307.8454 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9585 | 7.8803 | 2.742 | 77.9638 | -19.5678 | -1.5256 | 0 | 0 | 0 | 0 | 0 | 18.2122 | FALSE | 0.0025 | 0.0042223
757 | 758 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FHxSA | 0.0186354952155339 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7.86 | 8.86 | Manual | Baseline | 277.6268 | 72909.1064 | 0.0038078 | 60.3963 | 15893.7253 | 0.0038 | 4.5968 | 4.5873 | 60.7003 | 15919.4385 | 0.79363 | 0.99017 | 7.9884 | 9.0264 | 1.6334 | 0.12839 | 7.9245 | 8.8566 | 8.0714 | 9.1612 | 0.14688 | 0.30458 | 0.076852 | 0.071015 | 22.7935 | 172.6057 | 0.010101 | 0.0027659 | TRUE | 0.885 | TRUE | [M-H]- | 186.355 | 7.9279 | 8.0608 | 0.13294 | 7.9338 | 8.0537 | 0.11985 | -4.1535 | 0.9451 | 0.86132 | 4.5747 | 0 | 13 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 277.6268 | No Control Sample | 0 | off | off | 277.6268 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9536 | 2.4519 | 0.97573 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 19.1494 | FALSE | 0.0025 | 0.0038078
759 | 760 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FHpSA | 0.00785249042620452 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.58 | 8.86 | Manual | Baseline | 242.3067 | 72909.1064 | 0.0033234 | 45.0368 | 15893.7253 | 0.0028336 | 5.3802 | 4.5873 | 45.1828 | 15919.4385 | 0.99268 | 0.99017 | 8.5864 | 9.0264 | 0.074816 | 0.0064192 | 8.5008 | 8.8566 | 8.6818 | 9.1612 | 0.18097 | 0.30458 | 0.088154 | 0.071015 | 176.5439 | 172.6057 | 0 | 0.0027659 | TRUE | 0.95126 | TRUE | [M-H]- | 78.5249 | 8.517 | 8.664 | 0.14707 | 8.5219 | 8.6575 | 0.13569 | 0 | 1.1039 | 1.1986 | 6.6132 | 0 | 12 | 7 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 242.3067 | No Control Sample | 0 | off | off | 242.3067 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9516 | 4.9222 | 2.2049 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 12.9772 | FALSE | 0.0025 | 0.0033234
761 | 762 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FOSA | 0.0253563653633761 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.88 | 8.86 | Baseline | Baseline | 420.6436 | 72909.1064 | 0.0057694 | 100.9415 | 15893.7253 | 0.006351 | 4.1672 | 4.5873 | 100.9461 | 15919.4385 | 0.87147 | 0.99017 | 9.0367 | 9.0264 | 1.7647 | 0.1567 | 8.8885 | 8.8566 | 9.153 | 9.1612 | 0.2645 | 0.30458 | 0.064713 | 0.071015 | 32.0487 | 172.6057 | 8.0438e-05 | 0.0027659 | FALSE | 1.0011 | TRUE | [M-H]- | 253.5637 | 8.9617 | 9.1086 | 0.14685 | 8.975 | 9.1 | 0.12499 | 0.030698 | 0.97486 | 1.0131 | 3.7793 | 0 | 14 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 420.6436 | No Control Sample | 0 | off | off | 420.6436 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9504 | 8.4377 | 4.2015 | 77.9643 | -13.8602 | -1.0806 | 0 | 0 | 0 | 0 | 0 | 36.602 | FALSE | 0.0025 | 0.0057694
780 | 781 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 4:2 FTS | 0 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 5.94 | 5.92 | 0 | Baseline | 0 | 17866.7775 | 0 | 0 | 3012.7491 | 0 | 0 | 5.9304 | 24.1138 | 3016.0794 | 0 | 0.99353 | 0 | 5.95 | 0 | 0 | 0 | 5.7829 | 0 | 6.122 | 0 | 0.3391 | 0 | 0.093571 | 0 | 150.7662 | 0 | 0.0023214 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
782 | 783 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 6:2 FTS | 0.0157177322918391 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7.92 | 7.92 | Manual | Baseline | 160.4649 | 16826.6453 | 0.0095364 | 38.2164 | 3972.3361 | 0.0096206 | 4.1989 | 4.236 | 38.7234 | 3972.3361 | 0.96082 | 0.99887 | 7.9515 | 7.9443 | 0.39797 | 0.031519 | 7.8763 | 7.7183 | 8.0714 | 8.1227 | 0.19515 | 0.40445 | 0.044479 | 0.065727 | 45.6501 | 95.7169 | 0.0050918 | 9.2233e-11 | TRUE | 1.0009 | TRUE | [M-H]- | 157.1773 | 7.8824 | 8.0633 | 0.18094 | 7.8879 | 8.0553 | 0.16739 | -0.99711 | 1.2757 | 1.5584 | 3.8829 | 0 | 17 | 5 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 160.4649 | No Control Sample | 0 | off | off | 160.4649 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9707 | 6.4489 | 2.7535 | 406.9646 | 7.6054 | 3.0951 | 0 | 0 | 0 | 0 | 0 | 282.5514 | FALSE | 0.0025 | 0.0095364
784 | 785 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 8:2 FTS | 0 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.88 | 8.89 | 0 | Baseline | 0 | 19923.9469 | 0 | 0 | 4173.8894 | 0 | 0 | 4.7735 | 22.1437 | 4173.8906 | 0 | 0.99883 | 0 | 8.9011 | 0 | 0 | 0 | 8.7163 | 0 | 9.1056 | 0 | 0.38937 | 0 | 0.073918 | 0 | 211.7657 | 0 | 6.483e-07 | TRUE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
786 | 787 | PFAS CS1 0.01ng/ml | 6 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 7) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 10:2 FTS | 0 | Quantifiers | 10/26/2023 20:30:40 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.5 | 8.89 | 0 | Baseline | 0 | 19923.9469 | 0 | 0 | 4173.8894 | 0 | 0 | 4.7735 | 46.2807 | 4173.8906 | 0 | 0.99883 | 0 | 8.9011 | 0 | 0 | 0 | 8.7163 | 0 | 9.1056 | 0 | 0.38937 | 0 | 0.073918 | 0 | 211.7657 | 0 | 6.483e-07 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
847 | 848 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFBA | 0 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 1.49 | 1.48 | 0 | Baseline | 0 | 106593.9063 | 0 | 0 | 15106.1195 | 0 | 0 | 7.0563 | 68.5915 | 15170.6742 | 0 | 0.97766 | 0 | 1.4557 | 0 | 0 | 0 | 1.3042 | 0 | 1.6586 | 0 | 0.35442 | 0 | 0.10983 | 0 | 241.4629 | 0 | 0.010874 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
849 | 850 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFPeA | 0 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 3.18 | 3.21 | 0 | Manual | 0 | 82633.3049 | 0 | 0 | 8359.2821 | 0 | 0 | 9.8852 | 48.2541 | 8376.5889 | 0 | 0.99 | 0 | 3.1985 | 0 | 0 | 0 | 2.9688 | 0 | 3.5133 | 0 | 0.54455 | 0 | 0.15778 | 0 | 608.7428 | 0 | 0.0013158 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
851 | 852 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFHxA | 0.0221649710557512 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.40817 | 6.13 | 6.12 | Manual | Baseline | 832.0229 | 161294.6955 | 0.0051584 | 173.7674 | 29166.481 | 0.0059578 | 4.7881 | 5.5301 | 178.8366 | 29188.3456 | 0.68029 | 0.99429 | 6.1302 | 6.1381 | 0.0040589 | 0.00024881 | 6.0693 | 5.9372 | 6.2222 | 6.3288 | 0.15286 | 0.39152 | 0.097029 | 0.085036 | 35.2747 | 264.4304 | 0.0051717 | 0.0014623 | TRUE | 0.99873 | TRUE | [M-H]- | 221.6497 | 6.0717 | 6.2124 | 0.14072 | 6.0766 | 6.2072 | 0.13069 | -5.8791 | 1.339 | 1.7402 | 0.81585 | 0 | 18 | 13 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 832.0229 | No Control Sample | 203843.5852 | off | off | 832.0229 | 203843.5852 | off | 100 | off | 0.01 | off | 4.0817e-05 | off | 312.9754 | 8.2009 | 2.5667 | 268.9857 | 12.239 | 3.2921 | 0 | 0 | 0 | 0 | 0 | 9447.0071 | FALSE | 0.0025 | 0.0051584
853 | 854 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFHpA | 0.0213885867367985 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.43052 | 7.26 | 7.27 | Baseline | Baseline | 983.0057 | 183673.0303 | 0.0053519 | 243.5539 | 44544.7894 | 0.0054676 | 4.0361 | 4.1233 | 251.0893 | 44555.9708 | 0.71718 | 0.99786 | 7.288 | 7.2856 | 0.38499 | 0.02795 | 7.1913 | 7.1464 | 7.359 | 7.4723 | 0.1677 | 0.32583 | 0.067695 | 0.063451 | 47.7284 | 194.9919 | 0.020394 | 0.00057975 | FALSE | 1.0003 | TRUE | [M-H]- | 213.8859 | 7.2066 | 7.345 | 0.13834 | 7.2197 | 7.3411 | 0.12147 | -29.6181 | 0.81791 | 0.69804 | 1.4823 | 0 | 20 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 983.0057 | No Control Sample | 228328.7703 | off | off | 983.0057 | 228328.7703 | off | 100 | off | 0.01 | off | 4.3052e-05 | off | 362.9719 | 6.3998 | 2.3229 | 318.98 | 2.7259 | 0.86952 | 0 | 0 | 0 | 0 | 0 | 35.2851 | FALSE | 0.0025 | 0.0053519
855 | 856 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFOA | 0.0207689956805275 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.46614 | 7.96 | 7.94 | Baseline | Baseline | 892.5449 | 177200.8763 | 0.0050369 | 181.9344 | 44416.2164 | 0.0040961 | 4.9059 | 3.9896 | 189.5583 | 44435.2117 | 0.51132 | 0.99713 | 7.9796 | 7.9735 | 0.24565 | 0.019554 | 7.8728 | 7.8248 | 8.057 | 8.1288 | 0.18418 | 0.30402 | 0.081139 | 0.059761 | 24.6272 | 140.9458 | 0.0052357 | 0.00079947 | FALSE | 1.0008 | TRUE | [M-H]- | 207.69 | 7.8962 | 8.0498 | 0.15359 | 7.9051 | 8.0428 | 0.13766 | -5.1718 | 0.78051 | 0.53891 | 1.5023 | 0 | 15 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 892.5449 | No Control Sample | 191475.4056 | off | off | 892.5449 | 191475.4056 | off | 100 | off | 0.01 | off | 4.6614e-05 | off | 412.9702 | 9.0495 | 3.7372 | 368.9827 | 17.9021 | 6.6055 | 0 | 0 | 0 | 0 | 0 | 35.0004 | FALSE | 0.0025 | 0.0050369
857 | 858 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFNA | 0.0172224433639912 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.375 | 8.46 | 8.47 | Baseline | Baseline | 596.7937 | 149164.7354 | 0.0040009 | 120.6899 | 31366.4908 | 0.0038477 | 4.9449 | 4.7555 | 126.4117 | 31377.0464 | 0.75964 | 0.99726 | 8.4959 | 8.485 | 0.42465 | 0.035925 | 8.3864 | 8.2737 | 8.5788 | 8.6693 | 0.19237 | 0.3956 | 0.080596 | 0.074304 | 30.5864 | 125.6532 | 0.0060067 | 0.00063696 | FALSE | 1.0013 | TRUE | [M-H]- | 172.2244 | 8.4094 | 8.5583 | 0.14898 | 8.4173 | 8.5542 | 0.13693 | -3.7686 | 0.92193 | 0.87958 | 0.5679 | 0 | 11 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Fail | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 596.7937 | No Control Sample | 159143.3885 | off | off | 596.7937 | 159143.3885 | off | 100 | off | 0.01 | off | 3.75e-05 | off | 462.965 | 3.879 | 1.7959 | 418.9745 | 3.1208 | 1.3075 | 0 | 0 | 0 | 0 | 0 | 29.0698 | FALSE | 0.0025 | 0.0040009
859 | 860 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFDA | 0.0103494302146139 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.87 | 8.86 | Manual | Baseline | 361.3966 | 140665.593 | 0.0025692 | 65.2878 | 30484.0276 | 0.0021417 | 5.5354 | 4.6144 | 66.6693 | 30489.8455 | 0.78458 | 0.99833 | 8.8979 | 8.8937 | 0.31459 | 0.027904 | 8.7975 | 8.7084 | 8.9861 | 9.0943 | 0.18852 | 0.38595 | 0.092643 | 0.071742 | 17.8694 | 205.1147 | 0.010263 | 0.00039446 | TRUE | 1.0005 | TRUE | [M-H]- | 103.4943 | 8.8056 | 8.9785 | 0.17293 | 8.8134 | 8.9682 | 0.15472 | 3.5542 | 0.88354 | 0.71883 | 0.76396 | 0 | 11 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Fail | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 361.3966 | No Control Sample | 0 | off | off | 361.3966 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9692 | 17.8113 | 9.1365 | 468.9703 | 0.64525 | 0.30261 | 0 | 0 | 0 | 0 | 0 | 7268.6739 | FALSE | 0.0025 | 0.0025692
861 | 862 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFUdA | 0.0084030390754037 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.21 | 9.2 | Manual | Baseline | 206.3984 | 110757.6381 | 0.0018635 | 51.4084 | 24320.3086 | 0.0021138 | 4.0149 | 4.5541 | 52.3529 | 24337.0236 | 0.64856 | 0.99531 | 9.2316 | 9.2321 | 0.23461 | 0.021608 | 9.1658 | 9.0952 | 9.3084 | 9.3996 | 0.14253 | 0.30438 | 0.063867 | 0.071409 | 18.4843 | 177.6527 | 0.0051174 | 0.0014549 | TRUE | 0.99995 | TRUE | [M-H]- | 84.0304 | 9.171 | 9.301 | 0.12995 | 9.1772 | 9.2927 | 0.1155 | -1.8458 | 1.2202 | 1.452 | 1.7242 | 0 | 10 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Fail | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 206.3984 | No Control Sample | 0 | off | off | 206.3984 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9519 | -8.8452 | -4.9794 | 518.9692 | 3.5844 | 1.8602 | 0 | 0 | 0 | 0 | 0 | 98.9165 | FALSE | 0.0025 | 0.0018635
863 | 864 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFDoA | 0.0138141891172658 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.47 | 9.47 | Manual | Baseline | 306.7771 | 85816.3739 | 0.0035748 | 60.4355 | 20994.2059 | 0.0028787 | 5.0761 | 4.0876 | 61.3415 | 21003.2275 | 0.81706 | 0.9958 | 9.5112 | 9.4961 | 0.43533 | 0.041226 | 9.4127 | 9.3487 | 9.5958 | 9.6298 | 0.18306 | 0.28117 | 0.079423 | 0.064252 | 33.7847 | 323.0052 | 0.010201 | 0.00077957 | TRUE | 1.0016 | TRUE | [M-H]- | 138.1419 | 9.4199 | 9.5892 | 0.16933 | 9.4286 | 9.5816 | 0.15297 | -3.3677 | 0.94133 | 0.8846 | 0.77563 | 0 | 13 | 7 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 306.7771 | No Control Sample | 0 | off | off | 306.7771 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9628 | 14.924 | 9.1477 | 568.9735 | 17.8339 | 10.1468 | 0 | 0 | 0 | 0 | 0 | 11105.9228 | FALSE | 0.0025 | 0.0035748
865 | 866 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFTrDA | 0.0117119286225811 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.68 | 9.47 | Manual | Baseline | 189.5882 | 85816.3739 | 0.0022092 | 44.8628 | 20994.2059 | 0.0021369 | 4.226 | 4.0876 | 45.6338 | 21003.2275 | 0.67048 | 0.9958 | 9.705 | 9.4961 | 0.25822 | 0.024995 | 9.6512 | 9.3487 | 9.7865 | 9.6298 | 0.13533 | 0.28117 | 0.064109 | 0.064252 | 26.3437 | 323.0052 | 0.010223 | 0.00077957 | TRUE | 1.022 | TRUE | [M-H]- | 117.1193 | 9.6559 | 9.7814 | 0.12547 | 9.6588 | 9.7784 | 0.11958 | -3.3889 | 1.6334 | 2.3645 | 0 | 0 | 14 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 189.5882 | No Control Sample | 0 | off | off | 189.5882 | 0 | off | 0 | off | 0 | off | 0 | off | 663.353 | 607.2513 | 402.5775 | 618.9605 | -1.8631 | -1.1532 | 0 | 0 | 0 | 0 | 0 | inf | FALSE | 0.0025 | 0.0022092
867 | 868 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFTeDA | 0.00475920905496636 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.87 | 9.86 | Manual | Baseline | 31.6191 | 18738.7578 | 0.0016874 | 13.0427 | 5416.6883 | 0.0024079 | 2.4243 | 3.4594 | 72.998 | 5416.7335 | 0.092952 | 0.99657 | 9.8913 | 9.8889 | 0.21567 | 0.021286 | 9.8391 | 9.7655 | 9.9349 | 10.0173 | 0.09582 | 0.2518 | 0.03632 | 0.053567 | 7.7761 | 582.569 | 0.0048052 | 1.6386e-05 | TRUE | 1.0002 | TRUE | [M-H]- | 47.5921 | 9.8428 | 9.9325 | 0.089744 | 9.8504 | 9.9299 | 0.07947 | -0.65408 | 0.87609 | 0.82325 | 0 | 0 | 12 | 5 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 31.6191 | No Control Sample | 0 | off | off | 31.6191 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9448 | -3.4163 | -2.4357 | 668.9691 | 16.7607 | 11.2122 | 0 | 0 | 0 | 0 | 0 | 9661.1277 | FALSE | 0.0025 | 0.0016874
874 | 875 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFBS | 0 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 3.5 | 3.5 | 0 | Baseline | 0 | 42649.6836 | 0 | 0 | 3684.5661 | 0 | 0 | 11.5752 | 39.72 | 3725.2339 | 0 | 0.96123 | 0 | 3.8406 | 0 | 0 | 0 | 3.6071 | 0 | 4.093 | 0 | 0.48588 | 0 | 0.17885 | 0 | 679.6694 | 0 | 0.0072642 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
876 | 877 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFPeS | 0 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 6.38 | 3.5 | 0 | Baseline | 0 | 42649.6836 | 0 | 0 | 3684.5661 | 0 | 0 | 11.5752 | 63.6565 | 3725.2339 | 0 | 0.96123 | 0 | 3.8406 | 0 | 0 | 0 | 3.6071 | 0 | 4.093 | 0 | 0.48588 | 0 | 0.17885 | 0 | 679.6694 | 0 | 0.0072642 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
878 | 879 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | L-PFHxS | 0.0237736123499983 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.00811 | 4 | 0.0020275 | 0 | 7.34 | 7.35 | Manual | Baseline | 285.9136 | 64383.7242 | 0.0044408 | 95.0154 | 15418.7574 | 0.0061623 | 3.0091 | 4.1757 | 95.8465 | 15424.3258 | 0.68415 | 0.99411 | 7.3474 | 7.3615 | 0.10053 | 0.0073786 | 7.3132 | 7.2327 | 7.4253 | 7.5474 | 0.11214 | 0.3147 | 0.042822 | 0.065126 | 55.1358 | 219.4782 | 0.0050691 | 0.00086134 | TRUE | 0.99808 | TRUE | [M-H]- | 293.1395 | 7.3145 | 7.417 | 0.10244 | 7.319 | 7.4127 | 0.093698 | -4.2949 | 1.74 | 2.7582 | 8.9391 | 0 | 15 | 7 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 285.9136 | No Control Sample | 0 | off | off | 285.9136 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9396 | 7.4096 | 2.956 | 79.9564 | -10.6279 | -0.84978 | 0 | 0 | 0 | 0 | 0 | 1655.667 | FALSE | 0.0020275 | 0.0044408
880 | 881 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | Br-PFHxS | 0 | Qualifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0.00189 | 4 | 0.0004725 | 0 | 7.16 | 7.35 | 0 | Baseline | 0 | 64383.7242 | 0 | 0 | 15418.7574 | 0 | 0 | 4.1757 | 95.8465 | 15424.3258 | 0 | 0.99411 | 0 | 7.3615 | 0 | 0 | 0 | 7.2327 | 0 | 7.5474 | 0 | 0.3147 | 0 | 0.065126 | 0 | 219.4782 | 0 | 0.00086134 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 285.9136 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0004725 | 0
882 | 883 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFHpS | 0 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7.98 | 7.35 | 0 | Baseline | 0 | 64383.7242 | 0 | 0 | 15418.7574 | 0 | 0 | 4.1757 | 12.8209 | 15424.3258 | 0 | 0.99411 | 0 | 7.3615 | 0 | 0 | 0 | 7.2327 | 0 | 7.5474 | 0 | 0.3147 | 0 | 0.065126 | 0 | 219.4782 | 0 | 0.00086134 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
884 | 885 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | L-PFOS | 0 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0.00788 | 4 | 0.00197 | 0 | 8.46 | 8.45 | 0 | Baseline | 0 | 40097.9594 | 0 | 0 | 8339.8523 | 0 | 0 | 4.808 | 41.5969 | 8341.8647 | 0 | 0.99533 | 0 | 8.4822 | 0 | 0 | 0 | 8.3246 | 0 | 8.6732 | 0 | 0.34858 | 0 | 0.074913 | 0 | 111.255 | 0 | 0.0005285 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.00197 | 0
886 | 887 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | Br-PFOS | 0 | Qualifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0.00211 | 4 | 0.0005275 | 0 | 8.3 | 8.45 | 0 | Baseline | 0 | 40097.9594 | 0 | 0 | 8339.8523 | 0 | 0 | 4.808 | 41.5969 | 8341.8647 | 0 | 0.99533 | 0 | 8.4822 | 0 | 0 | 0 | 8.3246 | 0 | 8.6732 | 0 | 0.34858 | 0 | 0.074913 | 0 | 111.255 | 0 | 0.0005285 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0005275 | 0
888 | 889 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFNS | 0 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.85 | 8.45 | 0 | Baseline | 0 | 40097.9594 | 0 | 0 | 8339.8523 | 0 | 0 | 4.808 | 17.6577 | 8341.8647 | 0 | 0.99533 | 0 | 8.4822 | 0 | 0 | 0 | 8.3246 | 0 | 8.6732 | 0 | 0.34858 | 0 | 0.074913 | 0 | 111.255 | 0 | 0.0005285 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
890 | 891 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFDS | 0 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.19 | 8.45 | 0 | Baseline | 0 | 40097.9594 | 0 | 0 | 8339.8523 | 0 | 0 | 4.808 | 9.6788 | 8341.8647 | 0 | 0.99533 | 0 | 8.4822 | 0 | 0 | 0 | 8.3246 | 0 | 8.6732 | 0 | 0.34858 | 0 | 0.074913 | 0 | 111.255 | 0 | 0.0005285 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
894 | 895 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FBSA | 0 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 4.6 | 8.86 | 0 | Baseline | 0 | 72800.0694 | 0 | 0 | 15819.1495 | 0 | 0 | 4.602 | 20.8884 | 15820.4569 | 0 | 0.99719 | 0 | 9.0193 | 0 | 0 | 0 | 8.7906 | 0 | 9.2342 | 0 | 0.44362 | 0 | 0.071062 | 0 | 117.2646 | 0 | 0.00014807 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
896 | 897 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FPeSA | 0.0119535618312977 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy, Asymmetry, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7 | 8.86 | Manual | Baseline | 356.3268 | 72800.0694 | 0.0048946 | 66.4896 | 15819.1495 | 0.0042031 | 5.3591 | 4.602 | 67.1646 | 15820.4569 | 0.98002 | 0.99719 | 7.1033 | 9.0193 | 1.476 | 0.10332 | 7.0214 | 8.7906 | 7.1848 | 9.2342 | 0.16338 | 0.44362 | 0.089895 | 0.071062 | 2374.7498 | 117.2646 | 0 | 0.00014807 | TRUE | 0.78757 | TRUE | [M-H]- | 119.5356 | 7.026 | 7.1772 | 0.15118 | 7.0312 | 7.1719 | 0.14065 | 0 | 0.75834 | 0.48874 | 2.723 | 0 | 20 | 11 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 356.3268 | No Control Sample | 0 | off | off | 356.3268 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9578 | 5.7587 | 2.0038 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.6504 | FALSE | 0.0025 | 0.0048946
898 | 899 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FHxSA | 0.0152926719570651 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7.86 | 8.86 | Manual | Baseline | 227.4856 | 72800.0694 | 0.0031248 | 43.8613 | 15819.1495 | 0.0027727 | 5.1865 | 4.602 | 44.7286 | 15820.4569 | 0.67677 | 0.99719 | 7.986 | 9.0193 | 1.603 | 0.126 | 7.9023 | 8.7906 | 8.1161 | 9.2342 | 0.2138 | 0.44362 | 0.072415 | 0.071062 | 23.6755 | 117.2646 | 0.015373 | 0.00014807 | TRUE | 0.88544 | TRUE | [M-H]- | 152.9267 | 7.9068 | 8.115 | 0.20825 | 7.9132 | 8.1061 | 0.19289 | -3.1539 | 1.3552 | 1.7385 | 6.6766 | 0 | 17 | 8 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 227.4856 | No Control Sample | 0 | off | off | 227.4856 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9543 | 4.2869 | 1.706 | 77.965 | -5.0915 | -0.39696 | 0 | 0 | 0 | 0 | 0 | 5.1452 | FALSE | 0.0025 | 0.0031248
900 | 901 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FHpSA | 0.0159564608901747 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.58 | 8.86 | Baseline | Baseline | 491.6371 | 72800.0694 | 0.0067533 | 104.0805 | 15819.1495 | 0.0065794 | 4.7236 | 4.602 | 104.0805 | 15820.4569 | 1 | 0.99719 | 8.5813 | 9.0193 | 0.015354 | 0.0013174 | 8.3703 | 8.7906 | 8.7588 | 9.2342 | 0.3885 | 0.44362 | 0.076553 | 0.071062 | 171.676 | 117.2646 | 0 | 0.00014807 | FALSE | 0.95144 | TRUE | [M-H]- | 159.5646 | 8.4935 | 8.6439 | 0.15034 | 8.4993 | 8.6378 | 0.13857 | 0 | 0.86434 | 0.70555 | 2.4426 | 0 | 22 | 7 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 491.6371 | No Control Sample | 0 | off | off | 491.6371 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9512 | 3.9708 | 1.7787 | 77.9647 | -8.9406 | -0.69705 | 0 | 0 | 0 | 0 | 0 | 10.3505 | FALSE | 0.0025 | 0.0067533
902 | 903 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FOSA | 0.015285010034667 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.88 | 8.86 | Manual | Baseline | 253.1879 | 72800.0694 | 0.0034779 | 47.8112 | 15819.1495 | 0.0030224 | 5.2956 | 4.602 | 48.3995 | 15820.4569 | 0.83312 | 0.99719 | 9.0231 | 9.0193 | 1.6118 | 0.14313 | 8.9336 | 8.7906 | 9.1237 | 9.2342 | 0.19013 | 0.44362 | 0.091873 | 0.071062 | 24.3433 | 117.2646 | 0.005087 | 0.00014807 | TRUE | 1.0004 | TRUE | [M-H]- | 152.8501 | 8.9489 | 9.1096 | 0.16071 | 8.9538 | 9.1043 | 0.15053 | 1.2792 | 1.2499 | 1.5332 | 7.1594 | 0 | 11 | 7 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 253.1879 | No Control Sample | 0 | off | off | 253.1879 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9502 | 8.0041 | 3.9856 | 77.9641 | -16.3481 | -1.2746 | 0 | 0 | 0 | 0 | 0 | 5.3241 | FALSE | 0.0025 | 0.0034779
921 | 922 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 4:2 FTS | 0 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 5.94 | 5.92 | 0 | Baseline | 0 | 19739.229 | 0 | 0 | 3486.0849 | 0 | 0 | 5.6623 | 28.9583 | 3493.0157 | 0 | 0.98877 | 0 | 5.9438 | 0 | 0 | 0 | 5.7957 | 0 | 6.1013 | 0 | 0.30562 | 0 | 0.087658 | 0 | 182.4283 | 0 | 0.00077029 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
923 | 924 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 6:2 FTS | 0.01276963224549 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7.92 | 7.92 | Manual | Baseline | 125.6267 | 16214.7795 | 0.0077477 | 32.4831 | 3829.0797 | 0.0084833 | 3.8675 | 4.2346 | 32.8128 | 3829.0797 | 0.97523 | 1 | 7.9552 | 7.9363 | 0.44504 | 0.035247 | 7.8829 | 7.715 | 8.018 | 8.1645 | 0.13513 | 0.44952 | 0.064239 | 0.065787 | 1335.0204 | 129.9891 | 0 | 0 | TRUE | 1.0024 | TRUE | [M-H]- | 127.6963 | 7.893 | 8.0105 | 0.11753 | 7.8976 | 8.006 | 0.10845 | 0 | 1.0037 | 1.0078 | 6.6256 | 0 | 13 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 125.6267 | No Control Sample | 0 | off | off | 125.6267 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9731 | 12.1358 | 5.1816 | 406.9645 | 7.45 | 3.0318 | 0 | 0 | 0 | 0 | 0 | 20226.1272 | FALSE | 0.0025 | 0.0077477
925 | 926 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 8:2 FTS | 0 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.88 | 8.89 | 0 | Baseline | 0 | 19782.1836 | 0 | 0 | 4130.2903 | 0 | 0 | 4.7895 | 42.674 | 4130.3402 | 0 | 0.99758 | 0 | 8.8936 | 0 | 0 | 0 | 8.7131 | 0 | 9.0624 | 0 | 0.34927 | 0 | 0.074543 | 0 | 125.2144 | 0 | 2.2822e-05 | TRUE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
927 | 928 | PFAS CS1 0.01ng/ml | 7 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 8) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 10:2 FTS | 0 | Quantifiers | 10/26/2023 20:47:16 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.5 | 8.89 | 0 | Baseline | 0 | 19782.1836 | 0 | 0 | 4130.2903 | 0 | 0 | 4.7895 | 21.0337 | 4130.3402 | 0 | 0.99758 | 0 | 8.8936 | 0 | 0 | 0 | 8.7131 | 0 | 9.0624 | 0 | 0.34927 | 0 | 0.074543 | 0 | 125.2144 | 0 | 2.2822e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
988 | 989 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFBA | 0 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 1.49 | 1.48 | 0 | Baseline | 0 | 111108.3154 | 0 | 0 | 15944.3287 | 0 | 0 | 6.9685 | 101.0372 | 15965.7843 | 0 | 0.99315 | 0 | 1.4717 | 0 | 0 | 0 | 1.3382 | 0 | 1.7136 | 0 | 0.3754 | 0 | 0.11273 | 0 | 210.4551 | 0 | 0.00023073 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
990 | 991 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFPeA | 0.0199284274164964 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy, Integration Quality | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 3.18 | 3.21 | Manual | Manual | 363.1232 | 81476.6115 | 0.0044568 | 60.5974 | 9190.6695 | 0.0065934 | 5.9924 | 8.8651 | 66.9732 | 9206.9376 | 0.35757 | 0.98944 | 3.1811 | 3.1978 | 0.035198 | 0.0011193 | 3.0564 | 2.971 | 3.3525 | 3.4686 | 0.29608 | 0.49758 | 0.11569 | 0.14569 | 13.6166 | 619.8319 | 0.016662 | 0.0016696 | TRUE | 0.99479 | TRUE | [M-H]- | 199.2843 | 3.0594 | 3.3468 | 0.28743 | 3.0605 | 3.3435 | 0.28302 | -3.4101 | 1.4202 | 1.8272 | 0 | 0 | 53 | 22 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 363.1232 | No Control Sample | 0 | off | off | 363.1232 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9783 | 8.7656 | 2.3051 | 218.9843 | -6.869 | -1.5042 | 0 | 0 | 0 | 0 | 0 | 64.6518 | FALSE | 0.0025 | 0.0044568
992 | 993 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFHxA | 0.0218425667892326 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.40288 | 6.13 | 6.12 | Baseline | Baseline | 821.2362 | 161553.4971 | 0.0050834 | 139.0104 | 28003.2852 | 0.0049641 | 5.9077 | 5.7691 | 140.0234 | 28017.5316 | 0.70151 | 0.99572 | 6.1579 | 6.1451 | 0.45553 | 0.027924 | 6.06 | 5.946 | 6.2386 | 6.3494 | 0.17862 | 0.40333 | 0.098572 | 0.08743 | 31.9295 | 249.9096 | 0.0075191 | 0.0010091 | FALSE | 1.0021 | TRUE | [M-H]- | 218.4257 | 6.0717 | 6.229 | 0.15738 | 6.0756 | 6.2252 | 0.14958 | 5.8519 | 0.85757 | 0.70379 | 0.95418 | 0 | 20 | 12 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 821.2362 | No Control Sample | 203843.5852 | off | off | 821.2362 | 203843.5852 | off | 100 | off | 0.01 | off | 4.0288e-05 | off | 312.9774 | 14.667 | 4.5904 | 268.9817 | -2.4619 | -0.66221 | 0 | 0 | 0 | 0 | 0 | 8331.1178 | FALSE | 0.0025 | 0.0050834
994 | 995 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFHpA | 0.0252005946445188 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.49596 | 7.26 | 7.27 | Manual | Baseline | 1132.4157 | 179583.5542 | 0.0063058 | 262.0033 | 42382.3064 | 0.0061819 | 4.3221 | 4.2372 | 266.5622 | 42400.5907 | 0.77215 | 0.99612 | 7.2979 | 7.3036 | 0.5223 | 0.037919 | 7.2091 | 7.1488 | 7.4188 | 7.485 | 0.20969 | 0.3362 | 0.062852 | 0.066323 | 48.8168 | 246.7425 | 0.0051126 | 0.00091739 | TRUE | 0.99922 | TRUE | [M-H]- | 252.0059 | 7.2263 | 7.4055 | 0.17914 | 7.2423 | 7.3817 | 0.1394 | 6.3882 | 1.321 | 1.6903 | 1.4526 | 0 | 26 | 10 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1132.4157 | No Control Sample | 228328.7703 | off | off | 1132.4157 | 228328.7703 | off | 100 | off | 0.01 | off | 4.9596e-05 | off | 362.972 | 6.6084 | 2.3986 | 318.9822 | 9.6266 | 3.0707 | 0 | 0 | 0 | 0 | 0 | 23.9312 | FALSE | 0.0025 | 0.0063058
996 | 997 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFOA | 0.0236944937450979 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.52249 | 7.96 | 7.94 | Baseline | Baseline | 1000.4463 | 174099.5457 | 0.0057464 | 257.5407 | 41822.1396 | 0.006158 | 3.8846 | 4.1629 | 259.2129 | 41837.8733 | 0.66431 | 0.99652 | 7.994 | 7.9934 | 0.4277 | 0.034045 | 7.9024 | 7.8543 | 8.0806 | 8.1357 | 0.17818 | 0.28143 | 0.060172 | 0.064949 | 41.3004 | 146.9767 | 0.0029571 | 0.00075541 | FALSE | 1.0001 | TRUE | [M-H]- | 236.9449 | 7.9254 | 8.0604 | 0.135 | 7.938 | 8.0534 | 0.11546 | -4.2741 | 0.93686 | 0.94268 | 1.1876 | 0 | 14 | 6 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1000.4463 | No Control Sample | 191475.4056 | off | off | 1000.4463 | 191475.4056 | off | 100 | off | 0.01 | off | 5.2249e-05 | off | 412.9732 | 16.3902 | 6.7686 | 368.9769 | 2.2754 | 0.83958 | 0 | 0 | 0 | 0 | 0 | 23.0208 | FALSE | 0.0025 | 0.0057464
998 | 999 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFNA | 0.0143810259442389 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0.32471 | 8.46 | 8.47 | Manual | Baseline | 516.7544 | 154678.8832 | 0.0033408 | 100.6243 | 33656.9985 | 0.0029897 | 5.1355 | 4.5957 | 103.463 | 33662.0499 | 0.68668 | 0.99797 | 8.5106 | 8.5066 | 0.59754 | 0.050552 | 8.4346 | 8.329 | 8.5993 | 8.7499 | 0.16465 | 0.42095 | 0.087175 | 0.070973 | 18.3095 | 116.3555 | 0.020668 | 0.00035073 | TRUE | 1.0005 | TRUE | [M-H]- | 143.8103 | 8.4366 | 8.5877 | 0.15111 | 8.4414 | 8.5809 | 0.13945 | -12.6308 | 1.23 | 1.4652 | 0.97892 | 0 | 11 | 7 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Fail | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 516.7544 | No Control Sample | 159143.3885 | off | off | 516.7544 | 159143.3885 | off | 100 | off | 0.01 | off | 3.2471e-05 | off | 462.9633 | 0.13744 | 0.063628 | 418.9752 | 4.8862 | 2.0472 | 0 | 0 | 0 | 0 | 0 | 9.8103 | FALSE | 0.0025 | 0.0033408
1000 | 1001 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFDA | 0.0161947271213573 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.87 | 8.86 | Baseline | Baseline | 520.4515 | 129457.4004 | 0.0040203 | 106.9014 | 26426.2928 | 0.0040453 | 4.8685 | 4.8988 | 106.906 | 26426.8141 | 0.95773 | 0.99969 | 8.9249 | 8.9189 | 0.61841 | 0.054853 | 8.8041 | 8.7133 | 9.0901 | 9.1767 | 0.28598 | 0.46338 | 0.075059 | 0.076838 | 44.8784 | 167.0238 | 7.6662e-05 | 4.4316e-05 | FALSE | 1.0007 | TRUE | [M-H]- | 161.9473 | 8.8499 | 9.0133 | 0.16335 | 8.8578 | 9.0045 | 0.14666 | -0.028657 | 1.0224 | 1.0377 | 0.9197 | 0 | 15 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 520.4515 | No Control Sample | 0 | off | off | 520.4515 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9618 | 3.5211 | 1.8062 | 468.9713 | 2.7108 | 1.2713 | 0 | 0 | 0 | 0 | 0 | 5960.5486 | FALSE | 0.0025 | 0.0040203
1002 | 1003 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFUdA | 0.00561975015393588 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.21 | 9.2 | Manual | Baseline | 119.9142 | 96218.2584 | 0.0012463 | 31.2176 | 20986.5155 | 0.0014875 | 3.8412 | 4.5848 | 32.1109 | 20990.3128 | 0.52268 | 0.99806 | 9.2516 | 9.2566 | 0.45132 | 0.041567 | 9.1797 | 9.1183 | 9.321 | 9.4346 | 0.14127 | 0.31627 | 0.062878 | 0.072465 | 14.0065 | 165.9792 | 0.015507 | 0.0003635 | TRUE | 0.99946 | TRUE | [M-H]- | 56.1975 | 9.1913 | 9.3084 | 0.11704 | 9.1956 | 9.3051 | 0.10958 | -3.4267 | 1.0215 | 1.0649 | 1.9877 | 0 | 9 | 5 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Fail | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 119.9142 | No Control Sample | 0 | off | off | 119.9142 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9584 | 2.849 | 1.6038 | 518.9743 | 13.4672 | 6.9891 | 0 | 0 | 0 | 0 | 0 | 83.6655 | FALSE | 0.0025 | 0.0012463
1004 | 1005 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFDoA | 0.00985345130796701 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.47 | 9.47 | Manual | Baseline | 212.5465 | 83356.2619 | 0.0025499 | 49.0739 | 20225.5188 | 0.0024263 | 4.3312 | 4.1213 | 49.9668 | 20231.2061 | 0.73727 | 0.9972 | 9.5277 | 9.5226 | 0.60972 | 0.057741 | 9.456 | 9.3862 | 9.6118 | 9.6722 | 0.15578 | 0.28607 | 0.068152 | 0.063832 | 22.6666 | 297.4991 | 0.0051166 | 0.00061179 | TRUE | 1.0005 | TRUE | [M-H]- | 98.5345 | 9.4646 | 9.6057 | 0.14117 | 9.4689 | 9.5985 | 0.12956 | -1.6119 | 1.1659 | 1.306 | 1.0019 | 0 | 13 | 7 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 212.5465 | No Control Sample | 0 | off | off | 212.5465 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9525 | -1.8664 | -1.144 | 568.9683 | 8.6086 | 4.898 | 0 | 0 | 0 | 0 | 0 | 34797.9752 | FALSE | 0.0025 | 0.0025499
1006 | 1007 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFTrDA | 0.00913517552730605 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.68 | 9.47 | Manual | Baseline | 143.6375 | 83356.2619 | 0.0017232 | 48.7667 | 20225.5188 | 0.0024111 | 2.9454 | 4.1213 | 49.9145 | 20231.2061 | 0.69889 | 0.9972 | 9.7313 | 9.5226 | 0.53047 | 0.051349 | 9.6829 | 9.3862 | 9.7941 | 9.6722 | 0.11117 | 0.28607 | 0.041335 | 0.063832 | 30.6168 | 297.4991 | 0.0051434 | 0.00061179 | TRUE | 1.0219 | TRUE | [M-H]- | 91.3518 | 9.6841 | 9.7915 | 0.10735 | 9.6878 | 9.7882 | 0.10039 | -2.2562 | 1.1687 | 1.3775 | 0 | 0 | 12 | 5 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 143.6375 | No Control Sample | 0 | off | off | 143.6375 | 0 | off | 0 | off | 0 | off | 0 | off | 663.1672 | 326.9473 | 216.7499 | 618.9599 | -2.8567 | -1.7682 | 0 | 0 | 0 | 0 | 0 | inf | FALSE | 0.0025 | 0.0017232
1008 | 1009 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFTeDA | 0.0152115811272209 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.87 | 9.86 | Manual | Baseline | 96.0791 | 17814.8014 | 0.0053932 | 38.9386 | 5553.3986 | 0.0070117 | 2.4674 | 3.2079 | 68.1107 | 5557.0015 | 0.25279 | 0.99306 | 9.9163 | 9.9111 | 0.46929 | 0.046319 | 9.8542 | 9.8266 | 9.9508 | 10.0031 | 0.09661 | 0.1765 | 0.036835 | 0.0506 | 19.4631 | 641.5137 | 0 | 0.0013045 | TRUE | 1.0005 | TRUE | [M-H]- | 152.1158 | 9.8598 | 9.9487 | 0.088898 | 9.865 | 9.9453 | 0.080251 | 0 | 0.80213 | 0.59884 | 0 | 0 | 12 | 5 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 96.0791 | No Control Sample | 0 | off | off | 96.0791 | 0 | off | 0 | off | 0 | off | 0 | off | 713.1885 | 338.3181 | 241.2029 | 668.9693 | 17.0764 | 11.4234 | 0 | 0 | 0 | 0 | 0 | inf | FALSE | 0.0025 | 0.0053932
1015 | 1016 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFBS | 0 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 3.5 | 3.5 | 0 | Baseline | 0 | 42038.6073 | 0 | 0 | 4186.7765 | 0 | 0 | 10.0408 | 26.805 | 4222.1587 | 0 | 0.96854 | 0 | 3.8343 | 0 | 0 | 0 | 3.6019 | 0 | 4.1016 | 0 | 0.49963 | 0 | 0.15234 | 0 | 794.3666 | 0 | 0.0073254 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
1017 | 1018 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFPeS | 0 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 6.38 | 3.5 | 0 | Baseline | 0 | 42038.6073 | 0 | 0 | 4186.7765 | 0 | 0 | 10.0408 | 26.8077 | 4222.1587 | 0 | 0.96854 | 0 | 3.8343 | 0 | 0 | 0 | 3.6019 | 0 | 4.1016 | 0 | 0.49963 | 0 | 0.15234 | 0 | 794.3666 | 0 | 0.0073254 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
1019 | 1020 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | L-PFHxS | 0.0158461356192351 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.00811 | 4 | 0.0020275 | 0 | 7.34 | 7.35 | Manual | Baseline | 189.1132 | 63890.2909 | 0.00296 | 65.7507 | 15174.6136 | 0.0043329 | 2.8762 | 4.2103 | 67.4275 | 15205.7292 | 0.67803 | 0.98519 | 7.395 | 7.3783 | 0.74875 | 0.054958 | 7.319 | 7.263 | 7.47 | 7.5226 | 0.15105 | 0.2596 | 0.039866 | 0.065559 | 46.1197 | 181.9378 | 0.010307 | 0.0018305 | TRUE | 1.0023 | TRUE | [M-H]- | 195.3901 | 7.3132 | 7.4666 | 0.15343 | 7.3363 | 7.4327 | 0.0964 | -4.4864 | 0.89989 | 0.55328 | 17.6443 | 0 | 20 | 7 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 189.1132 | No Control Sample | 0 | off | off | 189.1132 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9381 | 3.7925 | 1.513 | 79.9579 | 8.2189 | 0.65716 | 0 | 0 | 0 | 0 | 0 | 1035.8134 | FALSE | 0.0020275 | 0.00296
1021 | 1022 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | Br-PFHxS | 0 | Qualifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0.00189 | 4 | 0.0004725 | 0 | 7.16 | 7.35 | 0 | Baseline | 0 | 63890.2909 | 0 | 0 | 15174.6136 | 0 | 0 | 4.2103 | 67.4275 | 15205.7292 | 0 | 0.98519 | 0 | 7.3783 | 0 | 0 | 0 | 7.263 | 0 | 7.5226 | 0 | 0.2596 | 0 | 0.065559 | 0 | 181.9378 | 0 | 0.0018305 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 189.1132 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0004725 | 0
1023 | 1024 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFHpS | 0 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7.98 | 7.35 | 0 | Baseline | 0 | 63890.2909 | 0 | 0 | 15174.6136 | 0 | 0 | 4.2103 | 20.9771 | 15205.7292 | 0 | 0.98519 | 0 | 7.3783 | 0 | 0 | 0 | 7.263 | 0 | 7.5226 | 0 | 0.2596 | 0 | 0.065559 | 0 | 181.9378 | 0 | 0.0018305 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
1025 | 1026 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | L-PFOS | 0 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0.00788 | 4 | 0.00197 | 0 | 8.46 | 8.45 | 0 | Baseline | 0 | 41655.5941 | 0 | 0 | 9231.487 | 0 | 0 | 4.5123 | 28.4163 | 9234.518 | 0 | 0.99678 | 0 | 8.5076 | 0 | 0 | 0 | 8.3694 | 0 | 8.6963 | 0 | 0.3269 | 0 | 0.069571 | 0 | 137.5985 | 0 | 0.00052189 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.00197 | 0
1027 | 1028 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | Br-PFOS | 0 | Qualifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy | 0 | 0 | 0.00211 | 4 | 0.0005275 | 0 | 8.3 | 8.45 | 0 | Baseline | 0 | 41655.5941 | 0 | 0 | 9231.487 | 0 | 0 | 4.5123 | 28.4163 | 9234.518 | 0 | 0.99678 | 0 | 8.5076 | 0 | 0 | 0 | 8.3694 | 0 | 8.6963 | 0 | 0.3269 | 0 | 0.069571 | 0 | 137.5985 | 0 | 0.00052189 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Pass | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0005275 | 0
1029 | 1030 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFNS | 0 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.85 | 8.45 | 0 | Baseline | 0 | 41655.5941 | 0 | 0 | 9231.487 | 0 | 0 | 4.5123 | 39.5346 | 9234.518 | 0 | 0.99678 | 0 | 8.5076 | 0 | 0 | 0 | 8.3694 | 0 | 8.6963 | 0 | 0.3269 | 0 | 0.069571 | 0 | 137.5985 | 0 | 0.00052189 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
1031 | 1032 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | PFDS | 0 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.19 | 8.45 | 0 | Baseline | 0 | 41655.5941 | 0 | 0 | 9231.487 | 0 | 0 | 4.5123 | 25.6365 | 9234.518 | 0 | 0.99678 | 0 | 8.5076 | 0 | 0 | 0 | 8.3694 | 0 | 8.6963 | 0 | 0.3269 | 0 | 0.069571 | 0 | 137.5985 | 0 | 0.00052189 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
1035 | 1036 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FBSA | 0 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 4.6 | 8.86 | 0 | Baseline | 0 | 72019.1197 | 0 | 0 | 15674.3357 | 0 | 0 | 4.5947 | 30.9383 | 15676.1508 | 0 | 0.99783 | 0 | 9.0416 | 0 | 0 | 0 | 8.8167 | 0 | 9.2586 | 0 | 0.44188 | 0 | 0.071099 | 0 | 112.602 | 0 | 0.00022363 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
1037 | 1038 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FPeSA | 0.013854982246673 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7 | 8.86 | Manual | Baseline | 408.5763 | 72019.1197 | 0.0056732 | 93.9441 | 15674.3357 | 0.0059935 | 4.3491 | 4.5947 | 95.1518 | 15676.1508 | 0.96855 | 0.99783 | 7.1175 | 9.0416 | 1.6788 | 0.11752 | 7.029 | 8.8167 | 7.198 | 9.2586 | 0.16894 | 0.44188 | 0.070136 | 0.071099 | 1147.7781 | 112.602 | 0.0050897 | 0.00022363 | TRUE | 0.7872 | TRUE | [M-H]- | 138.5498 | 7.0341 | 7.1928 | 0.15864 | 7.0371 | 7.1892 | 0.15213 | -2.8303 | 1.0574 | 1.1115 | 2.0161 | 0 | 20 | 9 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 408.5763 | No Control Sample | 0 | off | off | 408.5763 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9585 | 7.8252 | 2.7228 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 25.916 | FALSE | 0.0025 | 0.0056732
1039 | 1040 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FHxSA | 0.0166318530725563 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7.86 | 8.86 | Manual | Baseline | 244.7525 | 72019.1197 | 0.0033984 | 64.1101 | 15674.3357 | 0.0040901 | 3.8177 | 4.5947 | 64.938 | 15676.1508 | 0.68295 | 0.99783 | 8.0117 | 9.0416 | 1.9303 | 0.15172 | 7.9585 | 8.8167 | 8.0826 | 9.2586 | 0.12404 | 0.44188 | 0.063187 | 0.071099 | 19.409 | 112.602 | 0.00509 | 0.00022363 | TRUE | 0.8861 | TRUE | [M-H]- | 166.3185 | 7.9633 | 8.0763 | 0.11307 | 7.9661 | 8.0691 | 0.10304 | -2.6308 | 1.0731 | 1.0667 | 6.5537 | 0 | 11 | 6 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 244.7525 | No Control Sample | 0 | off | off | 244.7525 | 0 | off | 0 | off | 0 | off | 0 | off | 397.954 | 3.6094 | 1.4364 | 77.9639 | -19.5489 | -1.5241 | 0 | 0 | 0 | 0 | 0 | 23.4666 | FALSE | 0.0025 | 0.0033984
1041 | 1042 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FHpSA | 0.00953045031272893 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.58 | 8.86 | Manual | Baseline | 290.4942 | 72019.1197 | 0.0040336 | 76.3972 | 15674.3357 | 0.004874 | 3.8024 | 4.5947 | 77.1728 | 15676.1508 | 0.78431 | 0.99783 | 8.589 | 9.0416 | 0.10537 | 0.0090405 | 8.5195 | 8.8167 | 8.6535 | 9.2586 | 0.13394 | 0.44188 | 0.059228 | 0.071099 | 28.784 | 112.602 | 0 | 0.00022363 | TRUE | 0.94995 | TRUE | [M-H]- | 95.3045 | 8.5335 | 8.6689 | 0.13545 | 8.537 | 8.6689 | 0.13194 | 0 | 1.3142 | 1.7474 | 4.1875 | 0 | 10 | 5 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 290.4942 | No Control Sample | 0 | off | off | 290.4942 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9502 | 1.8589 | 0.83268 | 77.9695 | 52.0534 | 4.0584 | 0 | 0 | 0 | 0 | 0 | 7.6024 | FALSE | 0.0025 | 0.0040336
1043 | 1044 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | FOSA | 0.0135226389390685 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.88 | 8.86 | Manual | Baseline | 221.5923 | 72019.1197 | 0.0030769 | 43.3545 | 15674.3357 | 0.002766 | 5.1112 | 4.5947 | 43.8452 | 15676.1508 | 0.70206 | 0.99783 | 9.0673 | 9.0416 | 2.1095 | 0.18733 | 8.9461 | 8.8167 | 9.1542 | 9.2586 | 0.20814 | 0.44188 | 0.073271 | 0.071099 | 20.6403 | 112.602 | 0.010164 | 0.00022363 | TRUE | 1.0028 | TRUE | [M-H]- | 135.2264 | 8.9634 | 9.172 | 0.2085 | 8.976 | 9.172 | 0.19592 | 2.1171 | 0.94514 | 1.0049 | 5.5 | 0 | 13 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 221.5923 | No Control Sample | 0 | off | off | 221.5923 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9456 | -1.2652 | -0.63002 | 77.9653 | -1.7065 | -0.13305 | 0 | 0 | 0 | 0 | 0 | 9.8952 | FALSE | 0.0025 | 0.0030769
1062 | 1063 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 4:2 FTS | 0 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 5.94 | 5.92 | 0 | Baseline | 0 | 19110.192 | 0 | 0 | 3211.0196 | 0 | 0 | 5.9514 | 34.7979 | 3211.1079 | 0 | 0.99812 | 0 | 5.9486 | 0 | 0 | 0 | 5.7625 | 0 | 6.1404 | 0 | 0.37788 | 0 | 0.092384 | 0 | 200.2872 | 0 | 5.6032e-05 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
1064 | 1065 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 6:2 FTS | 0.0131131397159486 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy, Asymmetry, Total Width | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 7.92 | 7.92 | Manual | Baseline | 126.5497 | 15906.0357 | 0.0079561 | 46.1285 | 3607.8723 | 0.012786 | 2.7434 | 4.4087 | 46.5249 | 3607.875 | 0.97741 | 0.99521 | 7.948 | 7.9574 | 0.35298 | 0.027956 | 7.8763 | 7.8042 | 7.9882 | 8.1237 | 0.11192 | 0.31952 | 0.036995 | 0.068518 | 1694.9608 | 112.0463 | 0.0050683 | 1.5004e-06 | TRUE | 0.99881 | TRUE | [M-H]- | 131.1314 | 7.884 | 7.9831 | 0.099167 | 7.8906 | 7.9776 | 0.086974 | 2.089 | 0.70638 | 0.36817 | 4.0919 | 0 | 11 | 5 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 126.5497 | No Control Sample | 0 | off | off | 126.5497 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9718 | 9.1871 | 3.9226 | 406.959 | -6.0937 | -2.4799 | 0 | 0 | 0 | 0 | 0 | 24902.5073 | FALSE | 0.0025 | 0.0079561
1066 | 1067 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 8:2 FTS | 0 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 8.88 | 8.89 | 0 | Baseline | 0 | 20771.2166 | 0 | 0 | 4150.1907 | 0 | 0 | 5.0049 | 0 | 4150.1907 | 0 | 1 | 0 | 8.9161 | 0 | 0 | 0 | 8.6806 | 0 | 9.2429 | 0 | 0.56232 | 0 | 0.078656 | 0 | 137.5772 | 0 | 0 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
1068 | 1069 | PFAS CS1 0.01ng/ml | 8 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 9) | CS1 | 0 | 0 | Standard | 0 | 0 | 0 | 2 | 1 | 10:2 FTS | 0 | Quantifiers | 10/26/2023 21:03:55 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0.01 | 4 | 0.0025 | 0 | 9.5 | 8.89 | 0 | Baseline | 0 | 20771.2166 | 0 | 0 | 4150.1907 | 0 | 0 | 5.0049 | 59.8413 | 4150.1907 | 0 | 1 | 0 | 8.9161 | 0 | 0 | 0 | 8.6806 | 0 | 9.2429 | 0 | 0.56232 | 0 | 0.078656 | 0 | 137.5772 | 0 | 0 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0025 | 0
1129 | 1130 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFBA | 0.0372639542004028 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 1.49 | 1.48 | Manual | Manual | 684.6797 | 102432.3072 | 0.0066842 | 119.9563 | 14898.2334 | 0.0080517 | 5.7077 | 6.8755 | 123.5535 | 14918.2875 | 0.4012 | 0.99154 | 1.4699 | 1.4712 | 1.3515 | 0.020137 | 1.3723 | 1.352 | 1.5605 | 1.7087 | 0.18819 | 0.35667 | 0.10524 | 0.10204 | 14.2786 | 212.8643 | 0.01036 | 0.00098615 | TRUE | 0.99912 | TRUE | [M-H]- | 149.0558 | 1.3739 | 1.5619 | 0.188 | 1.3767 | 1.5592 | 0.18256 | -6.6037 | 0.97006 | 0.94021 | 1.1514 | 0 | 39 | 22 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Fail | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 684.6797 | No Control Sample | 0 | off | off | 684.6797 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9824 | 14.8751 | 3.1681 | 168.9865 | -15.0736 | -2.5473 | 0 | 0 | 0 | 0 | 0 | 211.7499 | FALSE | 0.00625 | 0.0066842
1131 | 1132 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFPeA | 0.0305714891835996 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy, Integration Quality | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 3.18 | 3.21 | Manual | Baseline | 504.0222 | 73719.9563 | 0.006837 | 94.9731 | 8472.3016 | 0.01121 | 5.307 | 8.7013 | 102.9229 | 8584.2757 | 0.40728 | 0.9545 | 3.2033 | 3.2098 | 0.73213 | 0.023282 | 3.1032 | 3.0324 | 3.3707 | 3.4231 | 0.26759 | 0.39073 | 0.0838 | 0.13863 | 18.733 | 665.9284 | 0.021783 | 0.0015491 | TRUE | 0.99797 | TRUE | [M-H]- | 122.286 | 3.1051 | 3.3669 | 0.26186 | 3.1075 | 3.3605 | 0.25302 | 7.7314 | 1.4583 | 1.8965 | 1.4899 | 0 | 48 | 17 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 504.0222 | No Control Sample | 0 | off | off | 504.0222 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9787 | 10.2214 | 2.688 | 218.9846 | -5.4963 | -1.2036 | 0 | 0 | 0 | 0 | 0 | 5.484 | FALSE | 0.00625 | 0.006837
1133 | 1134 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFHxA | 0.0395589240074458 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0.65655 | 6.13 | 6.12 | Manual | Baseline | 1338.3298 | 145368.61 | 0.0092065 | 247.7178 | 24032.1426 | 0.010308 | 5.4026 | 6.0489 | 253.2923 | 24047.9526 | 0.76519 | 0.99508 | 6.1664 | 6.1602 | 0.59399 | 0.036411 | 6.024 | 5.962 | 6.2449 | 6.3467 | 0.22094 | 0.38467 | 0.079907 | 0.094047 | 46.4395 | 228.3407 | 0.0051382 | 0.0011871 | TRUE | 1.001 | TRUE | [M-H]- | 158.2357 | 6.0326 | 6.241 | 0.20835 | 6.0632 | 6.2325 | 0.16931 | -5.7609 | 0.81124 | 0.7299 | 1.2992 | 0 | 25 | 10 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1338.3298 | No Control Sample | 203843.5852 | off | off | 1338.3298 | 203843.5852 | off | 40 | off | 0.025 | off | 0.00016414 | off | 312.9742 | 4.4282 | 1.3859 | 268.98 | -8.9823 | -2.4161 | 0 | 0 | 0 | 0 | 0 | 3751.6668 | FALSE | 0.00625 | 0.0092065
1135 | 1136 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFHpA | 0.036851999716905 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0.68179 | 7.26 | 7.27 | Baseline | Baseline | 1556.7293 | 168819.848 | 0.0092212 | 432.172 | 38732.2729 | 0.011158 | 3.6021 | 4.3586 | 452.45 | 38740.2698 | 0.73343 | 0.99631 | 7.3036 | 7.3066 | 0.60031 | 0.043583 | 7.2179 | 7.1635 | 7.3664 | 7.4881 | 0.14857 | 0.32462 | 0.056854 | 0.068113 | 62.3885 | 227.4304 | 0.054781 | 0.00044406 | FALSE | 0.99959 | TRUE | [M-H]- | 147.408 | 7.2431 | 7.3586 | 0.11553 | 7.2483 | 7.3536 | 0.10522 | 159.3562 | 0.912 | 0.81066 | 1.7108 | 0 | 18 | 8 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1556.7293 | No Control Sample | 228328.7703 | off | off | 1556.7293 | 228328.7703 | off | 40 | off | 0.025 | off | 0.00017045 | off | 362.9734 | 10.5266 | 3.8209 | 318.9805 | 4.4042 | 1.4049 | 0 | 0 | 0 | 0 | 0 | 20.8219 | FALSE | 0.00625 | 0.0092212
1137 | 1138 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFOA | 0.0399077529516515 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0.83179 | 7.96 | 7.94 | Baseline | Baseline | 1592.6813 | 164559.4754 | 0.0096785 | 403.2807 | 40349.8064 | 0.0099946 | 3.9493 | 4.0783 | 414.8452 | 40362.7673 | 0.72002 | 0.99708 | 7.999 | 7.9898 | 0.49046 | 0.039041 | 7.9164 | 7.8295 | 8.0826 | 8.1695 | 0.1662 | 0.33997 | 0.062076 | 0.062073 | 70.0072 | 119.5318 | 0.0045708 | 0.00065976 | FALSE | 1.0012 | TRUE | [M-H]- | 159.631 | 7.9318 | 8.064 | 0.13221 | 7.9389 | 8.0574 | 0.11847 | -11.091 | 0.97579 | 0.95336 | 1.7685 | 0 | 13 | 6 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1592.6813 | No Control Sample | 191475.4056 | off | off | 1592.6813 | 191475.4056 | off | 40 | off | 0.025 | off | 0.00020795 | off | 412.9692 | 6.735 | 2.7813 | 368.9751 | -2.7415 | -1.0116 | 0 | 0 | 0 | 0 | 0 | 8.7288 | FALSE | 0.00625 | 0.0096785
1139 | 1140 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFNA | 0.0270848726417647 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 0.025 | 4 | 0.00625 | 0.58993 | 8.46 | 8.47 | Baseline | Baseline | 938.841 | 149211.3745 | 0.006292 | 212.4985 | 31629.0369 | 0.0067185 | 4.4181 | 4.7175 | 213.0669 | 31652.1872 | 0.84849 | 0.99532 | 8.4991 | 8.4986 | 0.46217 | 0.0391 | 8.3794 | 8.3698 | 8.6388 | 8.6986 | 0.25942 | 0.32877 | 0.068478 | 0.074219 | 45.6006 | 118.9132 | 0.0054104 | 0.00013774 | FALSE | 1.0001 | TRUE | [M-H]- | 108.3395 | 8.4214 | 8.57 | 0.14854 | 8.4279 | 8.5591 | 0.13117 | 4.4319 | 0.90509 | 0.73633 | 0.73244 | 0 | 15 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 938.841 | No Control Sample | 159143.3885 | off | off | 938.841 | 159143.3885 | off | 40 | off | 0.025 | off | 0.00014748 | off | 462.9635 | 0.57323 | 0.26539 | 418.9738 | 1.526 | 0.63936 | 0 | 0 | 0 | 0 | 0 | 9.8103 | FALSE | 0.00625 | 0.006292
1141 | 1142 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFDA | 0.0296391104001146 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.87 | 8.86 | Baseline | Baseline | 983.8917 | 133721.8498 | 0.0073577 | 167.4824 | 28289.3085 | 0.0059203 | 5.8746 | 4.7269 | 167.4824 | 28296.1424 | 0.95967 | 0.99782 | 8.9096 | 8.908 | 0.44603 | 0.039563 | 8.6562 | 8.7386 | 9.1374 | 9.0865 | 0.48122 | 0.34798 | 0.083839 | 0.073222 | 50.0624 | 141.8317 | 1.0908e-12 | 0.00049645 | FALSE | 1.0002 | TRUE | [M-H]- | 118.5564 | 8.7928 | 9.0149 | 0.22204 | 8.8065 | 8.9994 | 0.19299 | 3.7965e-10 | 0.89773 | 0.75345 | 1.1539 | 0 | 25 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 983.8917 | No Control Sample | 0 | off | off | 983.8917 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9631 | 6.0226 | 3.0893 | 468.9738 | 8.0098 | 3.7563 | 0 | 0 | 0 | 0 | 0 | 3259.9535 | FALSE | 0.00625 | 0.0073577
1143 | 1144 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFUdA | 0.0303743734214202 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.21 | 9.2 | Baseline | Baseline | 665.2986 | 98767.2731 | 0.006736 | 157.7056 | 21835.1639 | 0.0072226 | 4.2186 | 4.5233 | 165.7871 | 21845.2402 | 0.80827 | 0.99573 | 9.2425 | 9.2425 | 0.35246 | 0.032462 | 9.1582 | 9.0874 | 9.3873 | 9.3765 | 0.22903 | 0.28905 | 0.065142 | 0.070851 | 44.2312 | 148.2974 | 0.077475 | 0.00082616 | FALSE | 0.99999 | TRUE | [M-H]- | 121.4975 | 9.1707 | 9.3088 | 0.13811 | 9.1814 | 9.2971 | 0.11564 | -53.3469 | 1.0342 | 1.0626 | 1.3124 | 0 | 13 | 5 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Fail | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 665.2986 | No Control Sample | 0 | off | off | 665.2986 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9608 | 6.946 | 3.9103 | 518.967 | -0.5883 | -0.30531 | 0 | 0 | 0 | 0 | 0 | 11.9735 | FALSE | 0.00625 | 0.006736
1145 | 1146 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFDoA | 0.0266458542832908 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.47 | 9.47 | Baseline | Baseline | 578.9124 | 83956.7749 | 0.0068954 | 160.9803 | 22630.5835 | 0.0071134 | 3.5962 | 3.7099 | 164.7555 | 22631.4045 | 0.85944 | 0.99834 | 9.5067 | 9.5077 | 0.38804 | 0.036748 | 9.4218 | 9.3428 | 9.6048 | 9.6533 | 0.18297 | 0.31048 | 0.053489 | 0.056905 | 59.3666 | 408.1829 | 0.025206 | 6.7672e-05 | FALSE | 0.9999 | TRUE | [M-H]- | 106.5834 | 9.4427 | 9.5805 | 0.13785 | 9.4509 | 9.5633 | 0.11246 | 22.1768 | 1.1242 | 1.1173 | 1.9089 | 0 | 13 | 5 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 578.9124 | No Control Sample | 0 | off | off | 578.9124 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9595 | 9.5341 | 5.844 | 568.9632 | -0.29988 | -0.17062 | 0 | 0 | 0 | 0 | 0 | 6886.4834 | FALSE | 0.00625 | 0.0068954
1147 | 1148 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFTrDA | 0.0212782376591795 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.68 | 9.47 | Manual | Baseline | 336.98 | 83956.7749 | 0.0040137 | 88.6279 | 22630.5835 | 0.0039163 | 3.8022 | 3.7099 | 90.1135 | 22631.4045 | 0.83406 | 0.99834 | 9.7228 | 9.5077 | 0.44185 | 0.042771 | 9.6546 | 9.3428 | 9.7809 | 9.6533 | 0.1263 | 0.31048 | 0.063 | 0.056905 | 56.0631 | 408.1829 | 0.0047291 | 6.7672e-05 | TRUE | 1.0226 | TRUE | [M-H]- | 85.113 | 9.6643 | 9.7761 | 0.11187 | 9.6715 | 9.7732 | 0.1017 | -3.3186 | 1.1675 | 1.4983 | 0.8662 | 0 | 13 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 336.98 | No Control Sample | 0 | off | off | 336.98 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9521 | 2.5143 | 1.6669 | 618.9565 | -8.464 | -5.2389 | 0 | 0 | 0 | 0 | 0 | 14.1366 | FALSE | 0.00625 | 0.0040137
1149 | 1150 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFTeDA | 0.0345148653497078 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.87 | 9.86 | Manual | Baseline | 228.0143 | 18632.9803 | 0.012237 | 60.8509 | 5349.3491 | 0.011375 | 3.7471 | 3.4832 | 61.7823 | 5353.3914 | 0.47954 | 0.99505 | 9.8966 | 9.8967 | 0.26967 | 0.026617 | 9.8405 | 9.7902 | 9.9555 | 9.9905 | 0.11503 | 0.20023 | 0.068327 | 0.054349 | 30.9263 | 900.7167 | 0 | 0.0014345 | TRUE | 0.99999 | TRUE | [M-H]- | 138.0595 | 9.8439 | 9.9526 | 0.10873 | 9.8491 | 9.9485 | 0.099405 | 0 | 1.2147 | 1.5137 | 1.1366 | 0 | 13 | 9 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 228.0143 | No Control Sample | 0 | off | off | 228.0143 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9525 | 7.2914 | 5.1984 | 668.9662 | 12.4523 | 8.3301 | 0 | 0 | 0 | 0 | 0 | 5243.8022 | FALSE | 0.00625 | 0.012237
1156 | 1157 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFBS | 0.0513473326505648 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 3.5 | 3.5 | Manual | Manual | 568.3661 | 41819.2594 | 0.013591 | 75.3388 | 3700.9439 | 0.020357 | 7.5441 | 11.2996 | 75.5977 | 3711.9587 | 0.87265 | 0.98601 | 3.8828 | 3.8547 | 10.9367 | 0.38278 | 3.6642 | 3.6061 | 4.0516 | 4.1273 | 0.38747 | 0.52124 | 0.15575 | 0.16067 | 33.9029 | 742.8744 | 0.001419 | 0.00086112 | TRUE | 1.0073 | TRUE | [M-H]- | 205.3893 | 3.6701 | 4.0405 | 0.3704 | 3.6775 | 4.0329 | 0.35537 | -0.2759 | 0.84983 | 0.68828 | 5.6689 | 0 | 61 | 26 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 568.3661 | No Control Sample | 0 | off | off | 568.3661 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9456 | 8.7402 | 2.6128 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 287.8561 | FALSE | 0.00625 | 0.013591
1158 | 1159 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFPeS | 0.0347625895298739 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 6.38 | 3.5 | Manual | Manual | 387.4383 | 41819.2594 | 0.0092646 | 69.9663 | 3700.9439 | 0.018905 | 5.5375 | 11.2996 | 71.7355 | 3711.9587 | 0.83671 | 0.98601 | 6.4074 | 3.8547 | 0.4297 | 0.027415 | 6.3153 | 3.6061 | 6.4999 | 4.1273 | 0.18458 | 0.52124 | 0.11176 | 0.16067 | 31.6856 | 742.8744 | 0.015457 | 0.00086112 | TRUE | 1.6623 | TRUE | [M-H]- | 139.0504 | 6.3201 | 6.495 | 0.17485 | 6.324 | 6.4929 | 0.16894 | -5.8589 | 1.4013 | 1.8862 | 10.2183 | 0 | 22 | 14 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 387.4383 | No Control Sample | 0 | off | off | 387.4383 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9418 | 5.6097 | 1.9574 | 79.9558 | -16.9127 | -1.3523 | 0 | 0 | 0 | 0 | 0 | 7.0423 | FALSE | 0.00625 | 0.0092646
1160 | 1161 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | L-PFHxS | 0.0337130713181724 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.020275 | 4 | 0.0050688 | 0 | 7.34 | 7.35 | Manual | Baseline | 387.7739 | 61576.7301 | 0.0062974 | 89.5474 | 15061.7201 | 0.0059454 | 4.3304 | 4.0883 | 91.2872 | 15075.0249 | 0.63614 | 0.99145 | 7.3786 | 7.3809 | 0.52643 | 0.03864 | 7.3079 | 7.2669 | 7.4494 | 7.5403 | 0.14148 | 0.27342 | 0.076876 | 0.063022 | 24.3544 | 210.7204 | 0.010246 | 0.0012116 | TRUE | 0.99969 | TRUE | [M-H]- | 166.279 | 7.3101 | 7.444 | 0.13391 | 7.3173 | 7.4394 | 0.12216 | -6.4847 | 0.80528 | 0.6071 | 14.0129 | 0 | 19 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 387.7739 | No Control Sample | 0 | off | off | 387.7739 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9403 | 9.3226 | 3.7191 | 79.9567 | -5.6351 | -0.45056 | 0 | 0 | 0 | 0 | 0 | 743.3392 | FALSE | 0.0050688 | 0.0062974
1162 | 1163 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | Br-PFHxS | 0 | Qualifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0.004725 | 4 | 0.0011813 | 0 | 7.16 | 7.35 | 0 | Baseline | 0 | 61576.7301 | 0 | 0 | 15061.7201 | 0 | 0 | 4.0883 | 91.2872 | 15075.0249 | 0 | 0.99145 | 0 | 7.3809 | 0 | 0 | 0 | 7.2669 | 0 | 7.5403 | 0 | 0.27342 | 0 | 0.063022 | 0 | 210.7204 | 0 | 0.0012116 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 387.7739 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0011813 | 0
1164 | 1165 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFHpS | 0.0342327070923687 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7.98 | 7.35 | Manual | Baseline | 320.8153 | 61576.7301 | 0.00521 | 75.1547 | 15061.7201 | 0.0049898 | 4.2687 | 4.0883 | 76.1097 | 15075.0249 | 0.87309 | 0.99145 | 8.0223 | 7.3809 | 0.53 | 0.042294 | 7.9232 | 7.2669 | 8.1172 | 7.5403 | 0.19403 | 0.27342 | 0.06174 | 0.063022 | 40.0396 | 210.7204 | 0.005089 | 0.0012116 | TRUE | 1.0869 | TRUE | [M-H]- | 136.9308 | 7.9304 | 8.1063 | 0.17586 | 7.9678 | 8.0973 | 0.1295 | -1.9712 | 0.97367 | 1.4455 | 15.6076 | 0 | 15 | 6 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 320.8153 | No Control Sample | 0 | off | off | 320.8153 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9366 | 7.1058 | 3.19 | 79.9577 | 6.7355 | 0.53855 | 0 | 0 | 0 | 0 | 0 | 4.646 | FALSE | 0.00625 | 0.00521
1166 | 1167 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | L-PFOS | 0.0184457791561106 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 0.0197 | 4 | 0.004925 | 0 | 8.46 | 8.45 | Manual | Baseline | 182.7001 | 38105.3886 | 0.0047946 | 50.095 | 8406.5212 | 0.0059591 | 3.6471 | 4.5328 | 51.2785 | 8412.9081 | 0.38263 | 0.9935 | 8.4877 | 8.4972 | 0.3272 | 0.027681 | 8.4342 | 8.3355 | 8.5489 | 8.6325 | 0.11474 | 0.29698 | 0.060888 | 0.070351 | 18.3798 | 156.4947 | 0.015432 | 0.0013716 | TRUE | 0.99888 | TRUE | [M-H]- | 93.6334 | 8.4239 | 8.54 | 0.11605 | 8.4239 | 8.5363 | 0.11241 | -6.7374 | 0.90524 | 0.75372 | 21.8417 | 0 | 8 | 5 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 182.7001 | No Control Sample | 0 | off | off | 182.7001 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9316 | 2.8183 | 1.4061 | 79.9566 | -8.7135 | -0.69671 | 0 | 0 | 0 | 0 | 0 | 9.5564 | FALSE | 0.004925 | 0.0047946
1168 | 1169 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | Br-PFOS | 0.00880731128419409 | Qualifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.005275 | 4 | 0.0013188 | 0 | 8.3 | 8.45 | Manual | Baseline | 186.7528 | 38105.3886 | 0.004901 | 29.77 | 8406.5212 | 0.0035413 | 6.2732 | 4.5328 | 51.2785 | 8412.9081 | 0.42142 | 0.9935 | 8.3647 | 8.4972 | 0.77951 | 0.064699 | 8.2616 | 8.3355 | 8.4291 | 8.6325 | 0.16743 | 0.29698 | 0.12592 | 0.070351 | 10.9986 | 156.4947 | 0.017311 | 0.0013716 | TRUE | 0.98441 | TRUE | [M-H]- | 166.9632 | 8.2658 | 8.4453 | 0.17942 | 8.2708 | 8.4453 | 0.17452 | 3.078 | 0.76056 | 0.54372 | 1.0222 | 0 | 12 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 186.7528 | No Control Sample | 0 | off | off | 182.7001 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9316 | 2.8183 | 1.4061 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 9.5564 | FALSE | 0.0013188 | 0.004901
1170 | 1171 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFNS | 0.0428871275194416 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.85 | 8.45 | Manual | Baseline | 362.2209 | 38105.3886 | 0.0095058 | 74.0513 | 8406.5212 | 0.0088088 | 4.8915 | 4.5328 | 74.9386 | 8412.9081 | 0.853 | 0.9935 | 8.9033 | 8.4972 | 0.6028 | 0.053348 | 8.8304 | 8.3355 | 9.0258 | 8.6325 | 0.19536 | 0.29698 | 0.072183 | 0.070351 | 38.9191 | 156.4947 | 0.0050853 | 0.0013716 | TRUE | 1.0478 | TRUE | [M-H]- | 171.5485 | 8.8383 | 9.0263 | 0.18794 | 8.843 | 9.0128 | 0.16981 | 1.9276 | 1.5233 | 1.9783 | 11.7697 | 0 | 11 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Fail | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 362.2209 | No Control Sample | 0 | off | off | 362.2209 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9315 | 8.0929 | 4.4424 | 79.9564 | -8.1841 | -0.65438 | 0 | 0 | 0 | 0 | 0 | 5.6007 | FALSE | 0.00625 | 0.0095058
1172 | 1173 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFDS | 0.0186890404537566 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.19 | 8.45 | Manual | Baseline | 162.1801 | 38105.3886 | 0.0042561 | 42.7277 | 8406.5212 | 0.0050827 | 3.7957 | 4.5328 | 43.2942 | 8412.9081 | 0.85162 | 0.9935 | 9.2219 | 8.4972 | 0.34716 | 0.031904 | 9.1537 | 8.3355 | 9.3001 | 8.6325 | 0.14635 | 0.29698 | 0.061992 | 0.070351 | 49.014 | 156.4947 | 0.015275 | 0.0013716 | TRUE | 1.0853 | TRUE | [M-H]- | 74.7562 | 9.1631 | 9.2813 | 0.1182 | 9.1669 | 9.2773 | 0.11037 | -4.4597 | 1.0061 | 1.0098 | 25.5605 | 0 | 9 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Fail | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 162.1801 | No Control Sample | 0 | off | off | 162.1801 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9271 | 5.5116 | 3.301 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4.4334 | FALSE | 0.00625 | 0.0042561
1176 | 1177 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FBSA | 0.016537151276887 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 4.6 | 8.86 | Manual | Baseline | 312.0734 | 71568.6856 | 0.0043605 | 82.249 | 15716.4578 | 0.0052333 | 3.7943 | 4.5537 | 83.2888 | 15724.9057 | 0.67382 | 0.99471 | 4.9633 | 9.0317 | 7.8983 | 0.36332 | 4.8772 | 8.8873 | 5.0485 | 9.2475 | 0.1713 | 0.36013 | 0.08083 | 0.06924 | 56.2078 | 174.7741 | 0.0050887 | 0.00053621 | TRUE | 0.54955 | TRUE | [M-H]- | 66.1486 | 4.8806 | 5.0391 | 0.15859 | 4.885 | 5.0374 | 0.15236 | -2.4433 | 0.93297 | 0.89221 | 4.1115 | 0 | 32 | 16 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 312.0734 | No Control Sample | 0 | off | off | 312.0734 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9602 | 4.214 | 1.2556 | 77.9655 | 1.9808 | 0.15443 | 0 | 0 | 0 | 0 | 0 | 5.5689 | FALSE | 0.00625 | 0.0043605
1178 | 1179 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FPeSA | 0.028629510178134 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Asymmetry, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7 | 8.86 | Baseline | Baseline | 838.9891 | 71568.6856 | 0.011723 | 179.5086 | 15716.4578 | 0.011422 | 4.6738 | 4.5537 | 180.4865 | 15724.9057 | 0.94638 | 0.99471 | 7.1221 | 9.0317 | 1.7447 | 0.12213 | 7.0293 | 8.8873 | 7.2282 | 9.2475 | 0.19897 | 0.36013 | 0.083972 | 0.06924 | 85.5516 | 174.7741 | 0.011538 | 0.00053621 | FALSE | 0.78857 | TRUE | [M-H]- | 114.518 | 7.0576 | 7.1925 | 0.1349 | 7.0619 | 7.1797 | 0.11784 | -10.4092 | 0.79162 | 0.45593 | 3.0009 | 0 | 22 | 11 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 838.9891 | No Control Sample | 0 | off | off | 838.9891 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9566 | 2.3416 | 0.81476 | 77.9658 | 6.5753 | 0.51264 | 0 | 0 | 0 | 0 | 0 | 2.419 | FALSE | 0.00625 | 0.011723
1180 | 1181 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FHxSA | 0.0232262004036954 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7.86 | 8.86 | Manual | Baseline | 339.6565 | 71568.6856 | 0.0047459 | 78.1053 | 15716.4578 | 0.0049697 | 4.3487 | 4.5537 | 79.144 | 15724.9057 | 0.96823 | 0.99471 | 8.0026 | 9.0317 | 1.8144 | 0.14261 | 7.937 | 8.8873 | 8.1026 | 9.2475 | 0.16559 | 0.36013 | 0.059723 | 0.06924 | 285.8958 | 174.7741 | 0.005092 | 0.00053621 | TRUE | 0.88606 | TRUE | [M-H]- | 92.9048 | 7.9426 | 8.0918 | 0.14921 | 7.9466 | 8.0836 | 0.13701 | -2.4017 | 1.2703 | 1.5018 | 11.3538 | 0 | 14 | 6 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 339.6565 | No Control Sample | 0 | off | off | 339.6565 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9558 | 8.0528 | 3.2046 | 77.9658 | 4.9988 | 0.38973 | 0 | 0 | 0 | 0 | 0 | 19.0734 | FALSE | 0.00625 | 0.0047459
1182 | 1183 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FHpSA | 0.0249534380807245 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.58 | 8.86 | Baseline | Baseline | 755.8398 | 71568.6856 | 0.010561 | 175.4532 | 15716.4578 | 0.011164 | 4.3079 | 4.5537 | 175.4532 | 15724.9057 | 1 | 0.99471 | 8.5896 | 9.0317 | 0.11245 | 0.0096485 | 8.382 | 8.8873 | 8.8346 | 9.2475 | 0.45253 | 0.36013 | 0.065361 | 0.06924 | 69.4023 | 174.7741 | 0 | 0.00053621 | FALSE | 0.95106 | TRUE | [M-H]- | 99.8138 | 8.522 | 8.6782 | 0.15617 | 8.5284 | 8.6557 | 0.12729 | 0 | 1.1372 | 1.0443 | 4.5061 | 0 | 25 | 5 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 755.8398 | No Control Sample | 0 | off | off | 755.8398 | 0 | off | 0 | off | 0 | off | 0 | off | 447.952 | 5.8914 | 2.639 | 77.9648 | -7.9897 | -0.62292 | 0 | 0 | 0 | 0 | 0 | 10.1785 | FALSE | 0.00625 | 0.010561
1184 | 1185 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FOSA | 0.0384870298787943 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.88 | 8.86 | Baseline | Baseline | 626.7335 | 71568.6856 | 0.0087571 | 157.9665 | 15716.4578 | 0.010051 | 3.9675 | 4.5537 | 157.9665 | 15724.9057 | 0.94965 | 0.99471 | 9.0289 | 9.0317 | 1.6766 | 0.14888 | 8.8145 | 8.8873 | 9.2401 | 9.2475 | 0.42557 | 0.36013 | 0.059188 | 0.06924 | 43.2562 | 174.7741 | 0 | 0.00053621 | FALSE | 0.99969 | TRUE | [M-H]- | 153.9481 | 8.9525 | 9.0997 | 0.14718 | 8.9652 | 9.0912 | 0.1261 | 0 | 1.0217 | 1.124 | 6.7978 | 0 | 22 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 626.7335 | No Control Sample | 0 | off | off | 626.7335 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9481 | 3.8681 | 1.9261 | 77.9652 | -2.254 | -0.17573 | 0 | 0 | 0 | 0 | 0 | 5.3241 | FALSE | 0.00625 | 0.0087571
1203 | 1204 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 4:2 FTS | 0.0217048334481503 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Asymmetry, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 5.94 | 5.92 | Manual | Baseline | 186.0013 | 16877.5802 | 0.011021 | 35.8902 | 3061.5205 | 0.011723 | 5.1825 | 5.5128 | 36.4838 | 3082.3584 | 0.95684 | 0.97336 | 6.0068 | 5.9672 | 1.1241 | 0.066769 | 5.8601 | 5.7818 | 6.0615 | 6.1006 | 0.20144 | 0.3188 | 0.087395 | 0.085678 | 67.0407 | 198.5401 | 0.0051082 | 0.011868 | TRUE | 1.0066 | TRUE | [M-H]- | 86.8193 | 5.8624 | 6.0574 | 0.19503 | 5.866 | 6.0532 | 0.18729 | -0.91014 | 0.64488 | 0.26839 | 4.1813 | 0 | 25 | 12 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Fail | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 186.0013 | No Control Sample | 0 | off | off | 186.0013 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9748 | 1.4241 | 0.46566 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 214.1475 | FALSE | 0.00625 | 0.011021
1205 | 1206 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 6:2 FTS | 0.0256219962083346 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7.92 | 7.92 | Manual | Baseline | 256.5963 | 16506.1195 | 0.015546 | 58.3486 | 3865.6174 | 0.015094 | 4.3976 | 4.27 | 59.1208 | 3868.3763 | 0.96743 | 0.99492 | 7.9501 | 7.951 | 0.3795 | 0.030056 | 7.8887 | 7.8321 | 8.0571 | 8.0938 | 0.16835 | 0.26173 | 0.065484 | 0.065396 | 81.5049 | 106.7824 | 0.0050916 | 0.0015056 | TRUE | 0.99988 | TRUE | [M-H]- | 102.488 | 7.8957 | 8.0471 | 0.1514 | 7.8986 | 8.0376 | 0.13895 | -1.7647 | 1.2523 | 1.4154 | 5.3867 | 0 | 15 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 256.5963 | No Control Sample | 0 | off | off | 256.5963 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9729 | 11.6151 | 4.9593 | 406.9624 | 2.3015 | 0.93663 | 0 | 0 | 0 | 0 | 0 | 10380.9208 | FALSE | 0.00625 | 0.015546
1207 | 1208 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 8:2 FTS | 0.0249167100263445 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.88 | 8.89 | Manual | Baseline | 245.6273 | 20091.4969 | 0.012225 | 46.5721 | 4036.2706 | 0.011538 | 5.2741 | 4.9777 | 47.1538 | 4036.8688 | 0.97165 | 0.99827 | 8.9031 | 8.9065 | 0.2601 | 0.023096 | 8.8199 | 8.7433 | 8.9998 | 9.0737 | 0.17986 | 0.3304 | 0.08738 | 0.077927 | 671.2512 | 139.613 | 0.0050879 | 0.00029083 | TRUE | 0.99961 | TRUE | [M-H]- | 99.6668 | 8.8309 | 8.9868 | 0.15584 | 8.8398 | 8.9813 | 0.14144 | 1.3174 | 1.0997 | 1.2842 | 4.0042 | 0 | 10 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Fail | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 245.6273 | No Control Sample | 0 | off | off | 245.6273 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9672 | 10.82 | 5.7017 | 506.9614 | 11.3326 | 5.7451 | 0 | 0 | 0 | 0 | 0 | 10389.0346 | FALSE | 0.00625 | 0.012225
1209 | 1210 | PFAS CS2 0.025ng/ml | 9 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 10) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 10:2 FTS | 0 | Quantifiers | 10/26/2023 21:20:30 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.5 | 8.89 | 0 | Baseline | 0 | 20091.4969 | 0 | 0 | 4036.2706 | 0 | 0 | 4.9777 | 44.1509 | 4036.8688 | 0 | 0.99827 | 0 | 8.9065 | 0 | 0 | 0 | 8.7433 | 0 | 9.0737 | 0 | 0.3304 | 0 | 0.077927 | 0 | 139.613 | 0 | 0.00029083 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.00625 | 0
1270 | 1271 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFBA | 0.0351121758026359 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 1.49 | 1.48 | Manual | Baseline | 682.8294 | 108415.8957 | 0.0062982 | 124.3645 | 16254.2058 | 0.0076512 | 5.4906 | 6.67 | 139.7748 | 16265.8082 | 0.38226 | 0.994 | 1.4569 | 1.4535 | 2.2222 | 0.03311 | 1.3483 | 1.3047 | 1.5856 | 1.6948 | 0.23725 | 0.39008 | 0.084427 | 0.098558 | 17.9348 | 253.6428 | 0.011297 | 0.0020179 | TRUE | 1.0024 | TRUE | [M-H]- | 140.4487 | 1.3503 | 1.5754 | 0.22509 | 1.3531 | 1.571 | 0.21792 | -5.9219 | 0.93516 | 0.85257 | 1.5378 | 0 | 49 | 19 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Fail | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 682.8294 | No Control Sample | 0 | off | off | 682.8294 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9812 | 9.5674 | 2.0376 | 168.9891 | 0.88232 | 0.1491 | 0 | 0 | 0 | 0 | 0 | 217.6143 | FALSE | 0.00625 | 0.0062982
1272 | 1273 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFPeA | 0.0354497276067048 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy, Integration Quality | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 3.18 | 3.21 | Manual | Manual | 628.632 | 79293.1328 | 0.007928 | 98.2881 | 8775.6439 | 0.0112 | 6.3958 | 9.0356 | 102.6938 | 8796.5731 | 0.49728 | 0.9884 | 3.1808 | 3.18 | 0.024151 | 0.00076799 | 3.0356 | 2.9719 | 3.3198 | 3.4652 | 0.2842 | 0.49325 | 0.13684 | 0.13759 | 19.3701 | 653.645 | 0.036753 | 0.00098717 | TRUE | 1.0002 | TRUE | [M-H]- | 141.7989 | 3.0366 | 3.3128 | 0.27619 | 3.0422 | 3.3095 | 0.26732 | -12.7105 | 0.98947 | 0.99489 | 0.80931 | 0 | 52 | 27 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 628.632 | No Control Sample | 0 | off | off | 628.632 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9785 | 9.5769 | 2.5185 | 218.9832 | -12.0824 | -2.6459 | 0 | 0 | 0 | 0 | 0 | 5.484 | FALSE | 0.00625 | 0.007928
1274 | 1275 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFHxA | 0.0360672251308496 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0.63531 | 6.13 | 6.12 | Baseline | Baseline | 1295.0453 | 154285.1774 | 0.0083938 | 215.9984 | 26700.7149 | 0.0080896 | 5.9956 | 5.7783 | 223.0166 | 26712.2993 | 0.7874 | 0.99588 | 6.1248 | 6.1182 | 0.085078 | 0.0052153 | 6.0205 | 5.9367 | 6.2367 | 6.3566 | 0.21613 | 0.41993 | 0.093901 | 0.0895 | 46.2415 | 204.5524 | 0.036114 | 0.00097393 | FALSE | 1.0011 | TRUE | [M-H]- | 144.2689 | 6.0338 | 6.2247 | 0.19093 | 6.0387 | 6.2137 | 0.17493 | 36.0916 | 0.87613 | 0.68216 | 1.037 | 0 | 24 | 12 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1295.0453 | No Control Sample | 203843.5852 | off | off | 1295.0453 | 203843.5852 | off | 40 | off | 0.025 | off | 0.00015883 | off | 312.9744 | 5.0061 | 1.5668 | 268.981 | -5.2378 | -1.4089 | 0 | 0 | 0 | 0 | 0 | 4661.321 | FALSE | 0.00625 | 0.0083938
1276 | 1277 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFHpA | 0.0410287366775178 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0.772 | 7.26 | 7.27 | Baseline | Baseline | 1762.702 | 171696.8046 | 0.010266 | 344.4507 | 40516.0161 | 0.0085016 | 5.1174 | 4.2378 | 345.9993 | 40536.0509 | 0.87002 | 0.99565 | 7.2747 | 7.2745 | 0.20181 | 0.014651 | 7.1253 | 7.1279 | 7.3681 | 7.4219 | 0.24285 | 0.29408 | 0.079307 | 0.06491 | 65.2366 | 297.3125 | 0.0065244 | 0.00096589 | FALSE | 1 | TRUE | [M-H]- | 164.1149 | 7.191 | 7.352 | 0.16107 | 7.2056 | 7.3424 | 0.13674 | 9.254 | 0.87496 | 0.76729 | 1.4214 | 0 | 28 | 11 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1762.702 | No Control Sample | 228328.7703 | off | off | 1762.702 | 228328.7703 | off | 40 | off | 0.025 | off | 0.000193 | off | 362.9724 | 7.6541 | 2.7782 | 318.9804 | 4.0006 | 1.2761 | 0 | 0 | 0 | 0 | 0 | 16.7086 | FALSE | 0.00625 | 0.010266
1278 | 1279 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFOA | 0.038311167072027 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Integration Quality | 0 | 0 | 0.025 | 4 | 0.00625 | 0.80788 | 7.96 | 7.94 | Baseline | Baseline | 1546.8986 | 166489.8396 | 0.0092912 | 343.568 | 38722.9041 | 0.0088725 | 4.5025 | 4.2995 | 345.3026 | 38730.8021 | 0.74735 | 0.99774 | 7.9737 | 7.9636 | 0.17229 | 0.013714 | 7.8742 | 7.7717 | 8.1269 | 8.182 | 0.25278 | 0.41027 | 0.066266 | 0.065032 | 32.6815 | 118.3256 | 0.0028765 | 0.00042405 | FALSE | 1.0013 | TRUE | [M-H]- | 153.2447 | 7.9046 | 8.0889 | 0.18433 | 7.9133 | 8.0474 | 0.13413 | 3.9095 | 1.3638 | 1.2788 | 1.8517 | 0 | 19 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1546.8986 | No Control Sample | 191475.4056 | off | off | 1546.8986 | 191475.4056 | off | 40 | off | 0.025 | off | 0.00020197 | off | 412.9688 | 5.789 | 2.3907 | 368.9788 | 7.3573 | 2.7147 | 0 | 0 | 0 | 0 | 0 | 7.1433 | FALSE | 0.00625 | 0.0092912
1280 | 1281 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFNA | 0.0267282626421207 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 0.025 | 4 | 0.00625 | 0.55398 | 8.46 | 8.47 | Baseline | Baseline | 881.6244 | 141987.3193 | 0.0062092 | 175.0986 | 27957.883 | 0.0062629 | 5.035 | 5.0786 | 175.0986 | 27973.1538 | 0.91777 | 0.99529 | 8.4734 | 8.4769 | 0.1578 | 0.01335 | 8.2868 | 8.311 | 8.6473 | 8.6379 | 0.36047 | 0.32687 | 0.075162 | 0.079836 | 43.5605 | 161.2104 | 3.3963e-07 | 0.00094166 | FALSE | 0.99958 | TRUE | [M-H]- | 106.9131 | 8.3861 | 8.5753 | 0.18925 | 8.4018 | 8.5583 | 0.15654 | 0.00016498 | 1.0625 | 1.1346 | 0.83373 | 0 | 21 | 5 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 881.6244 | No Control Sample | 159143.3885 | off | off | 881.6244 | 159143.3885 | off | 40 | off | 0.025 | off | 0.0001385 | off | 462.9667 | 7.5967 | 3.517 | 418.9732 | 0.094619 | 0.039643 | 0 | 0 | 0 | 0 | 0 | 10.2189 | FALSE | 0.00625 | 0.0062092
1282 | 1283 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFDA | 0.025115833510128 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.87 | 8.86 | Baseline | Baseline | 815.9307 | 130865.7387 | 0.0062349 | 193.0966 | 26734.8437 | 0.0072227 | 4.2255 | 4.895 | 193.2035 | 26746.337 | 0.84388 | 0.99637 | 8.89 | 8.8938 | 0.22573 | 0.020022 | 8.7802 | 8.7107 | 9.0148 | 9.0416 | 0.2346 | 0.33097 | 0.066281 | 0.076586 | 36.9351 | 150.5904 | 0.00097037 | 0.00076222 | FALSE | 0.99958 | TRUE | [M-H]- | 100.4633 | 8.8189 | 8.9601 | 0.14126 | 8.8268 | 8.9475 | 0.12063 | -0.7987 | 1.0715 | 1.0819 | 1.5483 | 0 | 12 | 5 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Fail | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 815.9307 | No Control Sample | 0 | off | off | 815.9307 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9634 | 6.6225 | 3.3971 | 468.9754 | 11.4662 | 5.3773 | 0 | 0 | 0 | 0 | 0 | 3384.7863 | FALSE | 0.00625 | 0.0062349
1284 | 1285 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFUdA | 0.025051623260021 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.21 | 9.2 | Baseline | Baseline | 576.019 | 103682.319 | 0.0055556 | 129.4271 | 22747.5952 | 0.0056897 | 4.4505 | 4.5579 | 129.5484 | 22750.1557 | 0.88285 | 0.99802 | 9.2265 | 9.2312 | 0.17945 | 0.016528 | 9.1302 | 9.0802 | 9.3956 | 9.4506 | 0.26538 | 0.37033 | 0.06222 | 0.071806 | 36.515 | 132.4805 | 0.0014547 | 0.00025599 | FALSE | 0.99949 | TRUE | [M-H]- | 100.2065 | 9.1607 | 9.3429 | 0.18226 | 9.1713 | 9.3303 | 0.15905 | -0.70944 | 1.3446 | 1.7826 | 1.6476 | 0 | 15 | 5 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 576.019 | No Control Sample | 0 | off | off | 576.019 | 0 | off | 0 | off | 0 | off | 0 | off | 562.961 | 7.3538 | 4.1399 | 518.9707 | 6.4811 | 3.3635 | 0 | 0 | 0 | 0 | 0 | 16.1859 | FALSE | 0.00625 | 0.0055556
1286 | 1287 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFDoA | 0.0230575480883749 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.47 | 9.47 | Baseline | Baseline | 540.1097 | 90519.3394 | 0.0059668 | 114.6637 | 23713.9444 | 0.0048353 | 4.7104 | 3.8171 | 114.6637 | 23715.0455 | 0.92251 | 0.99888 | 9.5035 | 9.4989 | 0.35407 | 0.03353 | 9.3093 | 9.3512 | 9.668 | 9.6719 | 0.3588 | 0.32068 | 0.071131 | 0.059018 | 40.8234 | 415.4149 | 2.1195e-09 | 9.9666e-05 | FALSE | 1.0005 | TRUE | [M-H]- | 92.2302 | 9.3996 | 9.5692 | 0.1696 | 9.4299 | 9.5623 | 0.13243 | -6.7736e-07 | 0.77057 | 0.65936 | 0.78162 | 0 | 24 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 540.1097 | No Control Sample | 0 | off | off | 540.1097 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9571 | 5.5586 | 3.4072 | 568.9613 | -3.7281 | -2.1212 | 0 | 0 | 0 | 0 | 0 | 12511.9984 | FALSE | 0.00625 | 0.0059668
1288 | 1289 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFTrDA | 0.0265897819930765 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.68 | 9.47 | Baseline | Baseline | 454.0136 | 90519.3394 | 0.0050157 | 165.1193 | 23713.9444 | 0.006963 | 2.7496 | 3.8171 | 165.1193 | 23715.0455 | 0.98439 | 0.99888 | 9.7151 | 9.4989 | 0.36288 | 0.035126 | 9.5677 | 9.3512 | 9.8433 | 9.6719 | 0.2756 | 0.32068 | 0.039068 | 0.059018 | 460.2646 | 415.4149 | 0 | 9.9666e-05 | FALSE | 1.0228 | TRUE | [M-H]- | 106.3591 | 9.661 | 9.7749 | 0.11388 | 9.6693 | 9.7639 | 0.094557 | 0 | 1.003 | 0.95288 | 0.33415 | 0 | 24 | 6 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 454.0136 | No Control Sample | 0 | off | off | 454.0136 | 0 | off | 0 | off | 0 | off | 0 | off | 663.3513 | 604.6292 | 400.8392 | 618.9585 | -5.2431 | -3.2453 | 0 | 0 | 0 | 0 | 0 | inf | FALSE | 0.00625 | 0.0050157
1290 | 1291 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFTeDA | 0.0311591797378633 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.87 | 9.86 | Manual | Baseline | 234.9069 | 21263.5715 | 0.011047 | 70.4608 | 6521.6296 | 0.010804 | 3.3339 | 3.2605 | 71.2767 | 6521.6725 | 0.44765 | 0.99833 | 9.8893 | 9.8968 | 0.19514 | 0.019261 | 9.8075 | 9.7759 | 9.9689 | 10.0165 | 0.1614 | 0.24067 | 0.041803 | 0.050205 | 39.8902 | 944.3924 | 0.004705 | 1.3104e-05 | TRUE | 0.99924 | TRUE | [M-H]- | 124.6367 | 9.8126 | 9.9635 | 0.15089 | 9.8231 | 9.957 | 0.13382 | -2.054 | 0.92139 | 0.87549 | 0.80272 | 0 | 18 | 6 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 234.9069 | No Control Sample | 0 | off | off | 234.9069 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9534 | 8.5688 | 6.1091 | 668.9534 | -6.7293 | -4.5016 | 0 | 0 | 0 | 0 | 0 | 12305.6812 | FALSE | 0.00625 | 0.011047
1297 | 1298 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFBS | 0.0267352907009613 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 3.5 | 3.5 | Manual | Manual | 310.5554 | 43885.4107 | 0.0070765 | 61.5686 | 3785.9646 | 0.016262 | 5.0441 | 11.5916 | 61.8574 | 3795.2362 | 0.71676 | 0.98801 | 3.8217 | 3.8167 | 9.1909 | 0.32168 | 3.6789 | 3.5579 | 3.9274 | 4.0982 | 0.24845 | 0.54032 | 0.084447 | 0.1745 | 24.799 | 865.6797 | 0.0010656 | 0.0016707 | TRUE | 1.0013 | TRUE | [M-H]- | 106.9412 | 3.6861 | 3.9222 | 0.23605 | 3.6901 | 3.9171 | 0.22699 | -0.26405 | 0.83443 | 0.651 | 10.289 | 0 | 40 | 15 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 310.5554 | No Control Sample | 0 | off | off | 310.5554 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9453 | 7.6175 | 2.2772 | 79.9573 | 0.72417 | 0.057903 | 0 | 0 | 0 | 0 | 0 | 5.2009 | FALSE | 0.00625 | 0.0070765
1299 | 1300 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFPeS | 0.0277334202457027 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 6.38 | 3.5 | Manual | Manual | 324.3678 | 43885.4107 | 0.0073912 | 91.1956 | 3785.9646 | 0.024088 | 3.5568 | 11.5916 | 92.3334 | 3795.2362 | 0.92608 | 0.98801 | 6.3686 | 3.8167 | 0.17902 | 0.011422 | 6.2594 | 3.5579 | 6.4681 | 4.0982 | 0.20869 | 0.54032 | 0.036958 | 0.1745 | 79.3117 | 865.6797 | 0.0050878 | 0.0016707 | TRUE | 1.6686 | TRUE | [M-H]- | 110.9337 | 6.268 | 6.4559 | 0.1879 | 6.2749 | 6.4334 | 0.15857 | -2.2234 | 0.88894 | 0.60476 | 14.0841 | 0 | 24 | 6 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 324.3678 | No Control Sample | 0 | off | off | 324.3678 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9426 | 8.0091 | 2.7947 | 79.9571 | -0.65885 | -0.05268 | 0 | 0 | 0 | 0 | 0 | 11.2543 | FALSE | 0.00625 | 0.0073912
1301 | 1302 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | L-PFHxS | 0.0291502125610678 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.02028 | 4 | 0.00507 | 0 | 7.34 | 7.35 | Manual | Baseline | 333.4419 | 61237.1287 | 0.0054451 | 81.6981 | 14669.9534 | 0.0055691 | 4.0814 | 4.1743 | 82.9486 | 14676.8539 | 0.76397 | 0.99081 | 7.3483 | 7.3488 | 0.1127 | 0.0082725 | 7.2959 | 7.2236 | 7.4394 | 7.4916 | 0.14353 | 0.26798 | 0.067816 | 0.064533 | 43.3823 | 166.7715 | 0 | 0.00082437 | TRUE | 0.99993 | TRUE | [M-H]- | 143.7387 | 7.3002 | 7.4284 | 0.12821 | 7.304 | 7.4092 | 0.10523 | 0 | 1.4408 | 1.5853 | 15.5173 | 0 | 19 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 333.4419 | No Control Sample | 0 | off | off | 333.4419 | 0 | off | 0 | off | 0 | off | 0 | off | 398.939 | 6.004 | 2.3952 | 79.9564 | -10.0362 | -0.80247 | 0 | 0 | 0 | 0 | 0 | 566.8709 | FALSE | 0.00507 | 0.0054451
1303 | 1304 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | Br-PFHxS | 0 | Qualifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0.00473 | 4 | 0.0011825 | 0 | 7.16 | 7.35 | 0 | Baseline | 0 | 61237.1287 | 0 | 0 | 14669.9534 | 0 | 0 | 4.1743 | 82.9486 | 14676.8539 | 0 | 0.99081 | 0 | 7.3488 | 0 | 0 | 0 | 7.2236 | 0 | 7.4916 | 0 | 0.26798 | 0 | 0.064533 | 0 | 166.7715 | 0 | 0.00082437 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 333.4419 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0011825 | 0
1305 | 1306 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFHpS | 0.0319291810356887 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7.98 | 7.35 | Manual | Baseline | 297.5773 | 61237.1287 | 0.0048594 | 78.4766 | 14669.9534 | 0.0053495 | 3.7919 | 4.1743 | 79.0554 | 14676.8539 | 0.89274 | 0.99081 | 7.982 | 7.3488 | 0.025304 | 0.0020192 | 7.904 | 7.2236 | 8.0589 | 7.4916 | 0.15486 | 0.26798 | 0.056789 | 0.064533 | 35.544 | 166.7715 | 0.0050622 | 0.00082437 | TRUE | 1.0862 | TRUE | [M-H]- | 127.7167 | 7.9122 | 8.0508 | 0.13858 | 7.9199 | 8.0465 | 0.12663 | 2.5652 | 1.1067 | 1.3059 | 18.3471 | 0 | 13 | 6 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 297.5773 | No Control Sample | 0 | off | off | 297.5773 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9367 | 7.222 | 3.2422 | 79.9573 | 1.0868 | 0.086895 | 0 | 0 | 0 | 0 | 0 | 5.3867 | FALSE | 0.00625 | 0.0048594
1307 | 1308 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | L-PFOS | 0.0250102648449061 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.0197 | 4 | 0.004925 | 0 | 8.46 | 8.45 | Manual | Baseline | 251.8744 | 38744.5288 | 0.0065009 | 58.0332 | 7946.5989 | 0.0073029 | 4.3402 | 4.8756 | 58.8923 | 7948.791 | 0.59965 | 0.99507 | 8.4799 | 8.4767 | 0.23517 | 0.019895 | 8.3991 | 8.33 | 8.5456 | 8.7128 | 0.14652 | 0.38273 | 0.070993 | 0.076107 | 26.2156 | 127.9864 | 0.010199 | 0.00075141 | TRUE | 1.0004 | TRUE | [M-H]- | 126.9557 | 8.4035 | 8.537 | 0.13345 | 8.4113 | 8.5316 | 0.12025 | -4.0396 | 0.8784 | 0.76407 | 12.947 | 0 | 9 | 5 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 251.8744 | No Control Sample | 0 | off | off | 251.8744 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9324 | 4.2827 | 2.1368 | 79.957 | -3.7485 | -0.29972 | 0 | 0 | 0 | 0 | 0 | 6.8591 | FALSE | 0.004925 | 0.0065009
1309 | 1310 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | Br-PFOS | 0.00178244530617112 | Qualifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Asymmetry, Total Width | 0 | 0 | 0.00528 | 4 | 0.00132 | 0 | 8.3 | 8.45 | Manual | Baseline | 38.4294 | 38744.5288 | 0.00099187 | 7.9204 | 7946.5989 | 0.0009967 | 4.852 | 4.8756 | 58.8923 | 7948.791 | 0.10467 | 0.99507 | 8.2987 | 8.4767 | 0.016163 | 0.0013415 | 8.2289 | 8.33 | 8.3737 | 8.7128 | 0.14483 | 0.38273 | 0.089284 | 0.076107 | 3.5779 | 127.9864 | 0.0078859 | 0.00075141 | TRUE | 0.979 | TRUE | [M-H]- | 33.7584 | 8.234 | 8.3709 | 0.13695 | 8.2404 | 8.3627 | 0.12233 | -0.43128 | 0.77527 | 0.49282 | 0.15257 | 0 | 10 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 38.4294 | No Control Sample | 0 | off | off | 251.8744 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9324 | 4.2827 | 2.1368 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.8591 | FALSE | 0.00132 | 0.00099187
1311 | 1312 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFNS | 0.0336748395555953 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.85 | 8.45 | Manual | Baseline | 289.1852 | 38744.5288 | 0.0074639 | 72.0472 | 7946.5989 | 0.0090664 | 4.0138 | 4.8756 | 73.15 | 7948.791 | 0.78714 | 0.99507 | 8.8898 | 8.4767 | 0.44932 | 0.039765 | 8.8078 | 8.33 | 8.9626 | 8.7128 | 0.1548 | 0.38273 | 0.064467 | 0.076107 | 25.1239 | 127.9864 | 0 | 0.00075141 | TRUE | 1.0487 | TRUE | [M-H]- | 134.6994 | 8.8169 | 8.9516 | 0.1347 | 8.8274 | 8.9471 | 0.11975 | 0 | 0.933 | 0.93839 | 14.1658 | 0 | 9 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Fail | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 289.1852 | No Control Sample | 0 | off | off | 289.1852 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9311 | 7.4165 | 4.0711 | 79.9565 | -7.2935 | -0.58317 | 0 | 0 | 0 | 0 | 0 | 3.8416 | FALSE | 0.00625 | 0.0074639
1313 | 1314 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFDS | 0.0311836234106869 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.19 | 8.45 | Manual | Baseline | 275.1447 | 38744.5288 | 0.0071015 | 61.3531 | 7946.5989 | 0.0077207 | 4.4846 | 4.8756 | 62.1203 | 7948.791 | 0.75801 | 0.99507 | 9.2139 | 8.4767 | 0.26018 | 0.023911 | 9.1142 | 8.33 | 9.296 | 8.7128 | 0.18184 | 0.38273 | 0.069484 | 0.076107 | 19.6689 | 127.9864 | 0.005088 | 0.00075141 | TRUE | 1.087 | TRUE | [M-H]- | 124.7345 | 9.1282 | 9.2858 | 0.15755 | 9.1401 | 9.2771 | 0.13702 | -1.7167 | 0.93157 | 0.88552 | 12.9111 | 0 | 12 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Fail | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 275.1447 | No Control Sample | 0 | off | off | 275.1447 | 0 | off | 0 | off | 0 | off | 0 | off | 598.929 | 8.6266 | 5.1667 | 79.9567 | -3.0936 | -0.24735 | 0 | 0 | 0 | 0 | 0 | 5.7158 | FALSE | 0.00625 | 0.0071015
1317 | 1318 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FBSA | 0.0173428246611319 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy, Integration Quality | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 4.6 | 8.86 | Manual | Baseline | 324.6987 | 71004.7902 | 0.0045729 | 58.8909 | 14743.0695 | 0.0039945 | 5.5136 | 4.8161 | 60.3317 | 14744.7532 | 0.79185 | 0.9976 | 4.8517 | 9.0181 | 5.4719 | 0.25171 | 4.8018 | 8.8137 | 5.0528 | 9.1793 | 0.25099 | 0.36558 | 0.16616 | 0.075107 | 42.3437 | 121.9821 | 0.015444 | 6.6063e-05 | TRUE | 0.538 | TRUE | [M-H]- | 69.3713 | 4.8021 | 5.0453 | 0.24318 | 4.8045 | 5.0408 | 0.23628 | -3.6237 | 5.9597 | 12.1043 | 4.453 | 0 | 46 | 31 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 324.6987 | No Control Sample | 0 | off | off | 324.6987 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9608 | 6.2655 | 1.8669 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5.5689 | FALSE | 0.00625 | 0.0045729
1319 | 1320 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FPeSA | 0.023413059634361 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7 | 8.86 | Baseline | Baseline | 680.7147 | 71004.7902 | 0.0095869 | 129.5491 | 14743.0695 | 0.0087871 | 5.2545 | 4.8161 | 130.9354 | 14744.7532 | 0.96337 | 0.9976 | 7.09 | 9.0181 | 1.2852 | 0.089966 | 7.0123 | 8.8137 | 7.221 | 9.1793 | 0.20867 | 0.36558 | 0.082815 | 0.075107 | 81.638 | 121.9821 | 0.016734 | 6.6063e-05 | FALSE | 0.78619 | TRUE | [M-H]- | 93.6522 | 7.0246 | 7.1949 | 0.17034 | 7.0315 | 7.1868 | 0.15528 | -10.3889 | 1.3406 | 1.7423 | 3.9336 | 0 | 23 | 11 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 680.7147 | No Control Sample | 0 | off | off | 680.7147 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9568 | 2.7911 | 0.97117 | 77.9647 | -7.1419 | -0.55682 | 0 | 0 | 0 | 0 | 0 | 8.0467 | FALSE | 0.00625 | 0.0095869
1321 | 1322 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FHxSA | 0.0324601918622174 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7.86 | 8.86 | Baseline | Baseline | 470.9529 | 71004.7902 | 0.0066327 | 110.7357 | 14743.0695 | 0.007511 | 4.2529 | 4.8161 | 110.7384 | 14744.7532 | 0.96738 | 0.9976 | 7.9906 | 9.0181 | 1.6617 | 0.13061 | 7.8886 | 8.8137 | 8.0969 | 9.1793 | 0.20827 | 0.36558 | 0.068795 | 0.075107 | 50.2976 | 121.9821 | 4.8335e-05 | 6.6063e-05 | FALSE | 0.88607 | TRUE | [M-H]- | 129.8408 | 7.9194 | 8.0498 | 0.13041 | 7.9262 | 8.0451 | 0.11889 | -0.0257 | 0.83141 | 0.65961 | 7.0195 | 0 | 16 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 470.9529 | No Control Sample | 0 | off | off | 470.9529 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9544 | 4.605 | 1.8326 | 77.9638 | -20.1363 | -1.5699 | 0 | 0 | 0 | 0 | 0 | 15.8155 | FALSE | 0.00625 | 0.0066327
1323 | 1324 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FHpSA | 0.0223651258763608 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.58 | 8.86 | Baseline | Baseline | 672.1022 | 71004.7902 | 0.0094656 | 157.8728 | 14743.0695 | 0.010708 | 4.2572 | 4.8161 | 163.1779 | 14744.7532 | 0.87777 | 0.9976 | 8.5672 | 9.0181 | 0.14974 | 0.012848 | 8.4777 | 8.8137 | 8.649 | 9.1793 | 0.1713 | 0.36558 | 0.071035 | 0.075107 | 70.701 | 121.9821 | 0.0011059 | 6.6063e-05 | FALSE | 0.95 | TRUE | [M-H]- | 89.4605 | 8.502 | 8.6299 | 0.12785 | 8.5079 | 8.6255 | 0.11757 | 1.0192 | 1.0085 | 1.0446 | 5.2219 | 0 | 11 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 672.1022 | No Control Sample | 0 | off | off | 672.1022 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9528 | 7.544 | 3.3793 | 77.9646 | -10.275 | -0.80109 | 0 | 0 | 0 | 0 | 0 | 0.64758 | FALSE | 0.00625 | 0.0094656
1325 | 1326 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FOSA | 0.0368217670171698 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.88 | 8.86 | Baseline | Baseline | 594.8915 | 71004.7902 | 0.0083782 | 118.0659 | 14743.0695 | 0.0080082 | 5.0386 | 4.8161 | 118.0659 | 14744.7532 | 1 | 0.9976 | 9.0272 | 9.0181 | 1.658 | 0.14723 | 8.8038 | 8.8137 | 9.2711 | 9.1793 | 0.4673 | 0.36558 | 0.077367 | 0.075107 | 109.1968 | 121.9821 | 0 | 6.6063e-05 | FALSE | 1.001 | TRUE | [M-H]- | 147.2871 | 8.9408 | 9.1261 | 0.18528 | 8.9529 | 9.0921 | 0.13921 | 0 | 0.97329 | 0.67395 | 6.0194 | 0 | 24 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 594.8915 | No Control Sample | 0 | off | off | 594.8915 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9486 | 4.8778 | 2.4289 | 77.9653 | -0.71833 | -0.056005 | 0 | 0 | 0 | 0 | 0 | 8.9501 | FALSE | 0.00625 | 0.0083782
1344 | 1345 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 4:2 FTS | 0.0289818954188649 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 5.94 | 5.92 | Manual | Baseline | 254.4453 | 17290.931 | 0.014716 | 68.7202 | 3119.4925 | 0.022029 | 3.7026 | 5.5429 | 69.767 | 3131.7919 | 0.92644 | 0.98276 | 5.9254 | 5.9237 | 0.24592 | 0.014608 | 5.8499 | 5.7817 | 5.986 | 6.0822 | 0.13614 | 0.30052 | 0.048448 | 0.083658 | 1416.3766 | 198.3607 | 0.010203 | 0.0053179 | TRUE | 1.0003 | TRUE | [M-H]- | 115.9276 | 5.8535 | 5.9799 | 0.12633 | 5.8569 | 5.9735 | 0.11665 | -5.1505 | 0.96624 | 0.88008 | 3.3987 | 0 | 18 | 7 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Fail | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 254.4453 | No Control Sample | 0 | off | off | 254.4453 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9761 | 5.6046 | 1.8325 | 306.9715 | 12.6203 | 3.874 | 0 | 0 | 0 | 0 | 0 | 239.2764 | FALSE | 0.00625 | 0.014716
1346 | 1347 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 6:2 FTS | 0.0353189118355587 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7.92 | 7.92 | Manual | Baseline | 335.3581 | 15649.8085 | 0.021429 | 64.6483 | 3616.0091 | 0.017878 | 5.1874 | 4.3279 | 65.4565 | 3616.6348 | 0.97008 | 0.99267 | 7.9111 | 7.9262 | 0.11224 | 0.0088896 | 7.8214 | 7.7889 | 8.0058 | 8.0745 | 0.18445 | 0.28567 | 0.096618 | 0.066486 | 43.8922 | 84.4905 | 0.0050879 | 0.00037073 | TRUE | 0.99809 | TRUE | [M-H]- | 141.2756 | 7.8342 | 7.9998 | 0.16556 | 7.8432 | 7.9951 | 0.15194 | 1.7833 | 1.1585 | 1.4329 | 3.7127 | 0 | 17 | 10 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 335.3581 | No Control Sample | 0 | off | off | 335.3581 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9726 | 11.096 | 4.7377 | 406.963 | 3.7011 | 1.5062 | 0 | 0 | 0 | 0 | 0 | 10380.7219 | FALSE | 0.00625 | 0.021429
1348 | 1349 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 8:2 FTS | 0.0262798198806564 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.88 | 8.89 | Manual | Baseline | 244.4637 | 18959.1294 | 0.012894 | 60.3527 | 3821.1543 | 0.015794 | 4.0506 | 4.9616 | 60.9519 | 3821.1543 | 0.97357 | 0.99752 | 8.8816 | 8.8849 | 0.017992 | 0.0015977 | 8.8115 | 8.6781 | 8.9636 | 9.1045 | 0.15208 | 0.42647 | 0.062988 | 0.077909 | 1171.7814 | 182.5856 | 0.01015 | 7.845e-09 | TRUE | 0.99963 | TRUE | [M-H]- | 105.1193 | 8.8197 | 8.9487 | 0.12902 | 8.829 | 8.9406 | 0.11165 | -4.0281 | 0.90636 | 0.80288 | 6.2348 | 0 | 9 | 5 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Fail | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 244.4637 | No Control Sample | 0 | off | off | 244.4637 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9648 | 6.3081 | 3.3241 | 506.957 | 2.5587 | 1.2972 | 0 | 0 | 0 | 0 | 0 | 12184.8011 | FALSE | 0.00625 | 0.012894
1350 | 1351 | PFAS CS2 0.025ng/ml | 10 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 11) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 10:2 FTS | 0 | Quantifiers | 10/26/2023 21:37:05 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.5 | 8.89 | 0 | Baseline | 0 | 18959.1294 | 0 | 0 | 3821.1543 | 0 | 0 | 4.9616 | 53.5612 | 3821.1543 | 0 | 0.99752 | 0 | 8.8849 | 0 | 0 | 0 | 8.6781 | 0 | 9.1045 | 0 | 0.42647 | 0 | 0.077909 | 0 | 182.5856 | 0 | 7.845e-09 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.00625 | 0
1411 | 1412 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFBA | 0.0317445928113097 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 1.49 | 1.48 | Manual | Manual | 600.2661 | 105417.4537 | 0.0056942 | 119.3152 | 15909.2261 | 0.0074997 | 5.0309 | 6.6262 | 140.6051 | 15943.5301 | 0.35779 | 0.99015 | 1.4858 | 1.4837 | 0.27901 | 0.0041573 | 1.4166 | 1.3565 | 1.58 | 1.702 | 0.16333 | 0.34552 | 0.11044 | 0.099593 | 16.452 | 227.8771 | 0.023869 | 0.0019233 | TRUE | 1.0015 | TRUE | [M-H]- | 126.9784 | 1.4209 | 1.5734 | 0.15242 | 1.4234 | 1.5697 | 0.1463 | -17.4371 | 1.2679 | 1.5363 | 1.6177 | 0 | 34 | 24 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Fail | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 600.2661 | No Control Sample | 0 | off | off | 600.2661 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9809 | 8.0337 | 1.711 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 151.5737 | FALSE | 0.00625 | 0.0056942
1413 | 1414 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFPeA | 0.0398343804021768 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy, Integration Quality | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 3.18 | 3.21 | Manual | Baseline | 675.126 | 75784.2152 | 0.0089085 | 89.7267 | 8624.3411 | 0.010404 | 7.5242 | 8.7872 | 93.8729 | 8685.6243 | 0.53057 | 0.97142 | 3.2901 | 3.232 | 3.4638 | 0.11015 | 3.1209 | 3.0318 | 3.4179 | 3.469 | 0.29704 | 0.4372 | 0.16333 | 0.13276 | 18.3439 | 738.946 | 0.031546 | 0.0076034 | TRUE | 1.018 | TRUE | [M-H]- | 159.3375 | 3.1275 | 3.4163 | 0.28874 | 3.1311 | 3.4133 | 0.28223 | -9.5291 | 0.95144 | 0.90474 | 1.0219 | 0 | 51 | 29 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 675.126 | No Control Sample | 0 | off | off | 675.126 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9771 | 4.112 | 1.0814 | 218.9915 | 25.8752 | 5.6663 | 0 | 0 | 0 | 0 | 0 | 5.484 | FALSE | 0.00625 | 0.0089085
1415 | 1416 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFHxA | 0.0376882219119833 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0.67888 | 6.13 | 6.12 | Baseline | Baseline | 1383.8489 | 157773.8346 | 0.0087711 | 276.0777 | 29137.9027 | 0.0094749 | 5.0125 | 5.4147 | 282.1658 | 29158.8728 | 0.83321 | 0.99441 | 6.1783 | 6.1748 | 0.78743 | 0.048269 | 6.0856 | 5.9952 | 6.2973 | 6.3753 | 0.21173 | 0.38012 | 0.085196 | 0.084531 | 58.9441 | 276.0332 | 0.018522 | 0.0010928 | FALSE | 1.0006 | TRUE | [M-H]- | 150.7529 | 6.1006 | 6.2757 | 0.17506 | 6.1067 | 6.2554 | 0.14873 | -24.1512 | 1.1065 | 1.0368 | 1.1646 | 0 | 23 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1383.8489 | No Control Sample | 203843.5852 | off | off | 1383.8489 | 203843.5852 | off | 40 | off | 0.025 | off | 0.00016972 | off | 312.9751 | 7.1978 | 2.2527 | 268.9822 | -0.57328 | -0.1542 | 0 | 0 | 0 | 0 | 0 | 5289.1041 | FALSE | 0.00625 | 0.0087711
1417 | 1418 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFHpA | 0.0440855088511102 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0.8158 | 7.26 | 7.27 | Baseline | Baseline | 1862.7061 | 168857.3388 | 0.011031 | 488.8816 | 38640.6121 | 0.012652 | 3.8101 | 4.3699 | 498.1382 | 38652.8077 | 0.85015 | 0.99638 | 7.3182 | 7.3152 | 0.8013 | 0.058175 | 7.2296 | 7.1753 | 7.4113 | 7.4911 | 0.18167 | 0.31575 | 0.053856 | 0.067687 | 87.4767 | 188.4585 | 0.01208 | 0.00067724 | FALSE | 1.0004 | TRUE | [M-H]- | 176.342 | 7.2498 | 7.392 | 0.14218 | 7.2604 | 7.3853 | 0.12489 | -32.5096 | 1.0501 | 1.1879 | 1.6682 | 0 | 22 | 8 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1862.7061 | No Control Sample | 228328.7703 | off | off | 1862.7061 | 228328.7703 | off | 40 | off | 0.025 | off | 0.00020395 | off | 362.9718 | 6.0203 | 2.1852 | 318.9801 | 3.0115 | 0.96059 | 0 | 0 | 0 | 0 | 0 | 12.7256 | FALSE | 0.00625 | 0.011031
1419 | 1420 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFOA | 0.0366802184129661 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0.80502 | 7.96 | 7.94 | Baseline | Baseline | 1541.4181 | 173276.5651 | 0.0088957 | 383.1046 | 42128.0045 | 0.0090938 | 4.0235 | 4.1131 | 390.8224 | 42132.4983 | 0.65495 | 0.99837 | 7.994 | 7.9953 | 0.42711 | 0.033998 | 7.9142 | 7.8544 | 8.1122 | 8.231 | 0.19808 | 0.37663 | 0.062367 | 0.06329 | 37.0323 | 109.3264 | 0.0032197 | 0.00029101 | FALSE | 0.99983 | TRUE | [M-H]- | 146.7209 | 7.9425 | 8.0824 | 0.13988 | 7.9488 | 8.0678 | 0.11901 | 6.2272 | 1.276 | 1.4536 | 1.7987 | 0 | 15 | 6 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1541.4181 | No Control Sample | 191475.4056 | off | off | 1541.4181 | 191475.4056 | off | 40 | off | 0.025 | off | 0.00020126 | off | 412.9682 | 4.3636 | 1.802 | 368.98 | 10.4542 | 3.8573 | 0 | 0 | 0 | 0 | 0 | 14.036 | FALSE | 0.00625 | 0.0088957
1421 | 1422 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFNA | 0.0332709314693913 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0.72173 | 8.46 | 8.47 | Baseline | Baseline | 1148.5842 | 148605.3884 | 0.0077291 | 240.9131 | 32730.1337 | 0.0073606 | 4.7676 | 4.5403 | 241.5991 | 32743.9159 | 0.92137 | 0.99606 | 8.5055 | 8.504 | 0.53829 | 0.04554 | 8.3569 | 8.2965 | 8.6327 | 8.6561 | 0.27577 | 0.35958 | 0.073079 | 0.070168 | 53.5197 | 155.849 | 0.0043446 | 0.00072028 | FALSE | 1.0002 | TRUE | [M-H]- | 133.0837 | 8.4193 | 8.5846 | 0.16538 | 8.4373 | 8.575 | 0.1377 | -3.7955 | 1.0574 | 1.2884 | 0.91242 | 0 | 16 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1148.5842 | No Control Sample | 159143.3885 | off | off | 1148.5842 | 159143.3885 | off | 40 | off | 0.025 | off | 0.00018043 | off | 462.9676 | 9.3501 | 4.3288 | 418.9757 | 6.0197 | 2.5221 | 0 | 0 | 0 | 0 | 0 | 9.8103 | FALSE | 0.00625 | 0.0077291
1423 | 1424 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFDA | 0.0249595987356792 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.87 | 8.86 | Baseline | Baseline | 760.8265 | 122791.5024 | 0.0061961 | 150.389 | 24835.1655 | 0.0060555 | 5.0591 | 4.9443 | 150.389 | 24839.766 | 0.89398 | 0.99842 | 8.9169 | 8.9125 | 0.52841 | 0.04687 | 8.7231 | 8.7318 | 9.112 | 9.1011 | 0.3889 | 0.3693 | 0.080735 | 0.077474 | 32.6361 | 126.1058 | 1.2148e-12 | 0.00036073 | FALSE | 1.0005 | TRUE | [M-H]- | 99.8384 | 8.8387 | 9.0048 | 0.16607 | 8.8485 | 8.9933 | 0.14483 | 4.6977e-10 | 1.198 | 1.4328 | 1.9153 | 0 | 20 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 760.8265 | No Control Sample | 0 | off | off | 760.8265 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9626 | 4.9878 | 2.5586 | 468.973 | 6.498 | 3.0474 | 0 | 0 | 0 | 0 | 0 | 3211.0158 | FALSE | 0.00625 | 0.0061961
1425 | 1426 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFUdA | 0.021617432097743 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.21 | 9.2 | Manual | Baseline | 464.2391 | 96837.021 | 0.004794 | 90.372 | 21023.7182 | 0.0042986 | 5.137 | 4.6061 | 91.6687 | 21023.7183 | 0.90009 | 0.99945 | 9.255 | 9.247 | 0.48907 | 0.045044 | 9.1389 | 9.0641 | 9.3288 | 9.5318 | 0.1899 | 0.4677 | 0.080399 | 0.072725 | 39.8072 | 190.0261 | 0.010194 | 5.0624e-09 | TRUE | 1.0009 | TRUE | [M-H]- | 86.4697 | 9.1512 | 9.3222 | 0.17092 | 9.1606 | 9.3142 | 0.15353 | -4.8514 | 0.85331 | 0.6916 | 2.3801 | 0 | 11 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Fail | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 464.2391 | No Control Sample | 0 | off | off | 464.2391 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9606 | 6.6162 | 3.7246 | 518.967 | -0.65015 | -0.33741 | 0 | 0 | 0 | 0 | 0 | 29.9613 | FALSE | 0.00625 | 0.004794
1427 | 1428 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFDoA | 0.0217871757903819 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.47 | 9.47 | Baseline | Baseline | 497.951 | 88319.8322 | 0.005638 | 137.6608 | 21798.4868 | 0.0063152 | 3.6172 | 4.0516 | 137.6608 | 21798.4881 | 0.88868 | 0.99917 | 9.5138 | 9.5129 | 0.46287 | 0.043834 | 9.3823 | 9.3369 | 9.6293 | 9.7118 | 0.24692 | 0.37487 | 0.0556 | 0.063405 | 49.3489 | 279.6411 | 6.4623e-07 | 1.2549e-07 | FALSE | 1.0001 | TRUE | [M-H]- | 87.1487 | 9.452 | 9.5781 | 0.12608 | 9.4587 | 9.5691 | 0.11047 | 0.00036029 | 1.0291 | 1.0215 | 1.0665 | 0 | 17 | 5 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 497.951 | No Control Sample | 0 | off | off | 497.951 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9551 | 2.3662 | 1.4504 | 568.9689 | 9.6064 | 5.4657 | 0 | 0 | 0 | 0 | 0 | 15877.1562 | FALSE | 0.00625 | 0.005638
1429 | 1430 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFTrDA | 0.0203168691130365 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.68 | 9.47 | Manual | Baseline | 338.4759 | 88319.8322 | 0.0038324 | 95.7105 | 21798.4868 | 0.0043907 | 3.5365 | 4.0516 | 96.6822 | 21798.4881 | 0.91726 | 0.99917 | 9.7282 | 9.5129 | 0.4984 | 0.048245 | 9.641 | 9.3369 | 9.7862 | 9.7118 | 0.14522 | 0.37487 | 0.056162 | 0.063405 | 120.963 | 279.6411 | 0 | 1.2549e-07 | TRUE | 1.0226 | TRUE | [M-H]- | 81.2675 | 9.6484 | 9.7747 | 0.12633 | 9.682 | 9.7717 | 0.089644 | 0 | 0.75629 | 0.79846 | 0.52399 | 0 | 15 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 338.4759 | No Control Sample | 0 | off | off | 338.4759 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9473 | -4.7084 | -3.1214 | 618.9659 | 6.7603 | 4.1844 | 0 | 0 | 0 | 0 | 0 | 14.1366 | FALSE | 0.00625 | 0.0038324
1431 | 1432 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFTeDA | 0.0300549593576395 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.87 | 9.86 | Manual | Baseline | 229.5676 | 21543.7286 | 0.010656 | 95.1673 | 6627.8762 | 0.014359 | 2.4123 | 3.2505 | 96.649 | 6628.9538 | 0.53665 | 0.99708 | 9.9069 | 9.9021 | 0.37426 | 0.03694 | 9.8587 | 9.7986 | 9.9973 | 10.0177 | 0.13862 | 0.21912 | 0.030958 | 0.050217 | 58.4748 | 599.727 | 0.01531 | 0.00035213 | TRUE | 1.0005 | TRUE | [M-H]- | 120.2198 | 9.8655 | 9.9925 | 0.12708 | 9.8701 | 9.9621 | 0.092004 | 10.5112 | 1.5204 | 1.4763 | 0.59684 | 0 | 17 | 5 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 229.5676 | No Control Sample | 0 | off | off | 229.5676 | 0 | off | 0 | off | 0 | off | 0 | off | 712.957 | 13.6272 | 9.7155 | 668.9556 | -3.4291 | -2.2939 | 0 | 0 | 0 | 0 | 0 | 12316.5746 | FALSE | 0.00625 | 0.010656
1438 | 1439 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFBS | 0.0240018135901538 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Integration Quality, Asymmetry | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 3.5 | 3.5 | Manual | Manual | 274.0957 | 43144.377 | 0.006353 | 43.0281 | 4104.7833 | 0.010482 | 6.3702 | 10.5108 | 43.112 | 4114.4458 | 0.7562 | 0.99082 | 3.9057 | 3.8843 | 11.5922 | 0.40573 | 3.7562 | 3.6491 | 4.0127 | 4.2241 | 0.25656 | 0.57495 | 0.12779 | 0.1609 | 23.176 | 826.4906 | 0.0036687 | 0 | TRUE | 1.0055 | TRUE | [M-H]- | 96.0073 | 3.765 | 4.0051 | 0.24015 | 3.7679 | 4.0022 | 0.23429 | -0.61528 | 0.73755 | 0.46544 | 11.4056 | 0 | 41 | 22 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 274.0957 | No Control Sample | 0 | off | off | 274.0957 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9447 | 5.5584 | 1.6617 | 79.9585 | 15.6532 | 1.2516 | 0 | 0 | 0 | 0 | 0 | 11.1489 | FALSE | 0.00625 | 0.006353
1440 | 1441 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFPeS | 0.0274698380606772 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 6.38 | 3.5 | Manual | Manual | 315.8599 | 43144.377 | 0.007321 | 86.1481 | 4104.7833 | 0.020987 | 3.6665 | 10.5108 | 87.2203 | 4114.4458 | 0.73907 | 0.99082 | 6.4393 | 3.8843 | 0.92924 | 0.059285 | 6.3488 | 3.6491 | 6.5128 | 4.2241 | 0.16406 | 0.57495 | 0.049368 | 0.1609 | 36.4144 | 826.4906 | 0.0050877 | 0 | TRUE | 1.6578 | TRUE | [M-H]- | 109.8794 | 6.3535 | 6.504 | 0.15052 | 6.359 | 6.4851 | 0.12606 | -2.6716 | 0.87379 | 0.56395 | 13.6359 | 0 | 19 | 7 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 315.8599 | No Control Sample | 0 | off | off | 315.8599 | 0 | off | 0 | off | 0 | off | 0 | off | 348.942 | 6.3011 | 2.1987 | 79.9574 | 2.296 | 0.18359 | 0 | 0 | 0 | 0 | 0 | 16.5278 | FALSE | 0.00625 | 0.007321
1442 | 1443 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | L-PFHxS | 0.0349253434828556 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.02028 | 4 | 0.00507 | 0 | 7.34 | 7.35 | Baseline | Baseline | 396.2717 | 60741.9578 | 0.0065239 | 124.5583 | 14824.0763 | 0.0084024 | 3.1814 | 4.0975 | 124.6034 | 14850.5665 | 0.72597 | 0.98502 | 7.3988 | 7.388 | 0.80174 | 0.058848 | 7.3106 | 7.2607 | 7.4816 | 7.5202 | 0.17105 | 0.25947 | 0.045506 | 0.06306 | 59.6454 | 164.9346 | 0.0007553 | 0.0033509 | FALSE | 1.0015 | TRUE | [M-H]- | 172.2157 | 7.3297 | 7.4537 | 0.12397 | 7.3337 | 7.4428 | 0.10917 | -0.55001 | 0.8273 | 0.53849 | 14.2349 | 0 | 21 | 7 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 396.2717 | No Control Sample | 0 | off | off | 396.2717 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9386 | 4.8853 | 1.9489 | 79.9571 | -0.97658 | -0.078084 | 0 | 0 | 0 | 0 | 0 | 807.328 | FALSE | 0.00507 | 0.0065239
1444 | 1445 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | Br-PFHxS | 0 | Qualifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0.00473 | 4 | 0.0011825 | 0 | 7.16 | 7.35 | 0 | Baseline | 0 | 60741.9578 | 0 | 0 | 14824.0763 | 0 | 0 | 4.0975 | 124.6034 | 14850.5665 | 0 | 0.98502 | 0 | 7.388 | 0 | 0 | 0 | 7.2607 | 0 | 7.5202 | 0 | 0.25947 | 0 | 0.06306 | 0 | 164.9346 | 0 | 0.0033509 | TRUE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 396.2717 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0011825 | 0
1446 | 1447 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFHpS | 0.0229114495492904 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7.98 | 7.35 | Manual | Baseline | 211.8061 | 60741.9578 | 0.003487 | 48.0379 | 14824.0763 | 0.0032405 | 4.4091 | 4.0975 | 48.7732 | 14850.5665 | 0.80929 | 0.98502 | 8.0124 | 7.388 | 0.4055 | 0.032359 | 7.9381 | 7.2607 | 8.1014 | 7.5202 | 0.16331 | 0.25947 | 0.069092 | 0.06306 | 32.1359 | 164.9346 | 0 | 0.0033509 | TRUE | 1.0845 | TRUE | [M-H]- | 91.6458 | 7.9467 | 8.0935 | 0.14674 | 7.9539 | 8.0857 | 0.13188 | 0 | 1.3366 | 1.7601 | 23.4353 | 0 | 14 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 211.8061 | No Control Sample | 0 | off | off | 211.8061 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9356 | 4.9573 | 2.2255 | 79.9571 | -0.77464 | -0.061938 | 0 | 0 | 0 | 0 | 0 | 7.0996 | FALSE | 0.00625 | 0.003487
1448 | 1449 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | L-PFOS | 0.0266133460826624 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.0197 | 4 | 0.004925 | 0 | 8.46 | 8.45 | Manual | Baseline | 280.8958 | 40606.0261 | 0.0069176 | 60.1281 | 8710.5194 | 0.0069029 | 4.6716 | 4.6617 | 60.4317 | 8713.8245 | 0.74058 | 0.99652 | 8.51 | 8.5026 | 0.59095 | 0.049995 | 8.4373 | 8.3503 | 8.5882 | 8.6956 | 0.15094 | 0.34528 | 0.07528 | 0.072141 | 30.6574 | 145.7829 | 0 | 0.00080182 | TRUE | 1.0009 | TRUE | [M-H]- | 135.0931 | 8.4176 | 8.5813 | 0.16371 | 8.4415 | 8.5727 | 0.13123 | 0 | 0.80979 | 0.70117 | 10.5753 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 280.8958 | No Control Sample | 0 | off | off | 280.8958 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9357 | 11.0162 | 5.4963 | 79.9556 | -20.9437 | -1.6746 | 0 | 0 | 0 | 0 | 0 | 32.2707 | FALSE | 0.004925 | 0.0069176
1450 | 1451 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | Br-PFOS | 0.00204242916527534 | Qualifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.00528 | 4 | 0.00132 | 0 | 8.3 | 8.45 | Manual | Baseline | 46.1503 | 40606.0261 | 0.0011365 | 12.6121 | 8710.5194 | 0.0014479 | 3.6592 | 4.6617 | 60.4317 | 8713.8245 | 0.12926 | 0.99652 | 8.3894 | 8.5026 | 1.0766 | 0.089357 | 8.3279 | 8.3503 | 8.4398 | 8.6956 | 0.11194 | 0.34528 | 0.097384 | 0.072141 | 6.5103 | 145.7829 | 0.0047189 | 0.00080182 | TRUE | 0.98669 | TRUE | [M-H]- | 38.6824 | 8.3369 | 8.4603 | 0.12339 | 8.3438 | 8.4603 | 0.11652 | 0.53168 | 0.9985 | 1.1216 | 0.1643 | 0 | 8 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 46.1503 | No Control Sample | 0 | off | off | 280.8958 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9357 | 11.0162 | 5.4963 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 32.2707 | FALSE | 0.00132 | 0.0011365
1452 | 1453 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFNS | 0.0291287434392235 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.85 | 8.45 | Manual | Baseline | 262.1636 | 40606.0261 | 0.0064563 | 47.6289 | 8710.5194 | 0.005468 | 5.5043 | 4.6617 | 48.596 | 8713.8245 | 0.82524 | 0.99652 | 8.9096 | 8.5026 | 0.67336 | 0.059592 | 8.7738 | 8.3503 | 9.0404 | 8.6956 | 0.26661 | 0.34528 | 0.073161 | 0.072141 | 23.2247 | 145.7829 | 0.010254 | 0.00080182 | TRUE | 1.0479 | TRUE | [M-H]- | 116.515 | 8.7784 | 9.0295 | 0.2511 | 8.7925 | 9.0177 | 0.22522 | -1.8319 | 0.93703 | 0.8783 | 17.6067 | 0 | 14 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 262.1636 | No Control Sample | 0 | off | off | 262.1636 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9306 | 6.4285 | 3.5288 | 79.9567 | -5.098 | -0.40762 | 0 | 0 | 0 | 0 | 0 | 7.3891 | FALSE | 0.00625 | 0.0064563
1454 | 1455 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFDS | 0.0325072045300766 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.19 | 8.45 | Manual | Baseline | 300.6037 | 40606.0261 | 0.0074029 | 83.2299 | 8710.5194 | 0.0095551 | 3.6117 | 4.6617 | 84.2682 | 8713.8245 | 0.84229 | 0.99652 | 9.2307 | 8.5026 | 0.44334 | 0.040743 | 9.1578 | 8.3503 | 9.2999 | 8.6956 | 0.14212 | 0.34528 | 0.058105 | 0.072141 | 33.8643 | 145.7829 | 0.0050878 | 0.00080182 | TRUE | 1.0856 | TRUE | [M-H]- | 130.0288 | 9.1672 | 9.2901 | 0.12291 | 9.1769 | 9.2831 | 0.10621 | -2.9796 | 0.897 | 0.8069 | 13.1896 | 0 | 9 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Fail | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 300.6037 | No Control Sample | 0 | off | off | 300.6037 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9272 | 5.6324 | 3.3734 | 79.9567 | -2.471 | -0.19758 | 0 | 0 | 0 | 0 | 0 | 11.7196 | FALSE | 0.00625 | 0.0074029
1458 | 1459 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FBSA | 0.0208259974628314 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 4.6 | 8.86 | Manual | Baseline | 384.6912 | 70054.0719 | 0.0054913 | 63.8176 | 14724.5891 | 0.0043341 | 6.028 | 4.7576 | 64.4655 | 14736.8517 | 0.88179 | 0.99254 | 4.9969 | 9.0365 | 8.6277 | 0.39688 | 4.8497 | 8.8578 | 5.1208 | 9.2012 | 0.27108 | 0.34338 | 0.067163 | 0.073657 | 48.5076 | 111.0713 | 0 | 0.00153 | TRUE | 0.55296 | TRUE | [M-H]- | 83.304 | 4.8531 | 5.1148 | 0.26177 | 4.8576 | 5.1124 | 0.25481 | 0 | 0.81471 | 0.63201 | 3.4952 | 0 | 49 | 14 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 384.6912 | No Control Sample | 0 | off | off | 384.6912 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9623 | 11.1143 | 3.3116 | 77.9655 | 2.188 | 0.17059 | 0 | 0 | 0 | 0 | 0 | 5.5689 | FALSE | 0.00625 | 0.0054913
1460 | 1461 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FPeSA | 0.0279513602917614 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7 | 8.86 | Baseline | Baseline | 801.7808 | 70054.0719 | 0.011445 | 175.7692 | 14724.5891 | 0.011937 | 4.5616 | 4.7576 | 175.8162 | 14736.8517 | 0.9742 | 0.99254 | 7.1352 | 9.0365 | 1.9316 | 0.13521 | 7.0314 | 8.8578 | 7.24 | 9.2012 | 0.20862 | 0.34338 | 0.07733 | 0.073657 | 101.5119 | 111.0713 | 0.00050394 | 0.00153 | FALSE | 0.7896 | TRUE | [M-H]- | 111.8054 | 7.0608 | 7.2037 | 0.14291 | 7.0667 | 7.1925 | 0.12583 | 0.42459 | 0.95452 | 0.82405 | 3.7304 | 0 | 23 | 11 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 801.7808 | No Control Sample | 0 | off | off | 801.7808 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9582 | 6.9163 | 2.4066 | 77.9655 | 2.7293 | 0.21279 | 0 | 0 | 0 | 0 | 0 | 9.8884 | FALSE | 0.00625 | 0.011445
1462 | 1463 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FHxSA | 0.0179424780359442 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7.86 | 8.86 | Manual | Baseline | 256.8352 | 70054.0719 | 0.0036662 | 69.2471 | 14724.5891 | 0.0047028 | 3.709 | 4.7576 | 71.2677 | 14736.8517 | 0.61389 | 0.99254 | 8.0145 | 9.0365 | 1.9658 | 0.15451 | 7.9507 | 8.8578 | 8.0697 | 9.2012 | 0.11902 | 0.34338 | 0.061753 | 0.073657 | 19.1485 | 111.0713 | 0.015515 | 0.00153 | TRUE | 0.8869 | TRUE | [M-H]- | 71.7699 | 7.9544 | 8.0822 | 0.12781 | 7.9586 | 8.0636 | 0.10504 | 9.0269 | 1.0389 | 0.83375 | 14.143 | 0 | 10 | 6 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 256.8352 | No Control Sample | 0 | off | off | 256.8352 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9548 | 5.6265 | 2.2391 | 77.9644 | -12.5725 | -0.98022 | 0 | 0 | 0 | 0 | 0 | 3.4059 | FALSE | 0.00625 | 0.0036662
1464 | 1465 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FHpSA | 0.0307997747669238 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.58 | 8.86 | Baseline | Baseline | 913.1817 | 70054.0719 | 0.013035 | 191.575 | 14724.5891 | 0.013011 | 4.7667 | 4.7576 | 191.575 | 14736.8517 | 1 | 0.99254 | 8.5952 | 9.0365 | 0.17744 | 0.015224 | 8.397 | 8.8578 | 8.826 | 9.2012 | 0.42907 | 0.34338 | 0.07534 | 0.073657 | 139.944 | 111.0713 | 0 | 0.00153 | FALSE | 0.95116 | TRUE | [M-H]- | 123.1991 | 8.5239 | 8.6752 | 0.15127 | 8.5336 | 8.6625 | 0.12891 | 0 | 0.97207 | 0.89207 | 3.8843 | 0 | 24 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 913.1817 | No Control Sample | 0 | off | off | 913.1817 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9515 | 4.6792 | 2.0961 | 77.9659 | 6.8361 | 0.53298 | 0 | 0 | 0 | 0 | 0 | 0.57435 | FALSE | 0.00625 | 0.013035
1466 | 1467 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FOSA | 0.0341565217543876 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.88 | 8.86 | Baseline | Baseline | 544.4431 | 70054.0719 | 0.0077718 | 128.9028 | 14724.5891 | 0.0087543 | 4.2237 | 4.7576 | 128.9028 | 14736.8517 | 1 | 0.99254 | 9.0348 | 9.0365 | 1.7436 | 0.15483 | 8.8269 | 8.8578 | 9.2139 | 9.2012 | 0.387 | 0.34338 | 0.066361 | 0.073657 | 936.58 | 111.0713 | 0 | 0.00153 | FALSE | 0.99981 | TRUE | [M-H]- | 136.6261 | 8.9585 | 9.0952 | 0.13675 | 8.9676 | 9.0908 | 0.1232 | 0 | 0.84934 | 0.72504 | 6.4722 | 0 | 20 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 544.4431 | No Control Sample | 0 | off | off | 544.4431 | 0 | off | 0 | off | 0 | off | 0 | off | 497.947 | 1.6163 | 0.80481 | 77.9644 | -13.1386 | -1.0244 | 0 | 0 | 0 | 0 | 0 | 8.8965 | FALSE | 0.00625 | 0.0077718
1485 | 1486 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 4:2 FTS | 0.0165653424910706 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 5.94 | 5.92 | Manual | Baseline | 143.2081 | 17026.2027 | 0.008411 | 34.2045 | 3093.9055 | 0.011055 | 4.1868 | 5.5031 | 34.6979 | 3139.0147 | 0.90679 | 0.94822 | 5.9483 | 5.9805 | 0.13923 | 0.0082704 | 5.9089 | 5.8617 | 6.0625 | 6.119 | 0.15361 | 0.25737 | 0.11088 | 0.086946 | 155.4638 | 154.3152 | 0.0050976 | 0.0095472 | TRUE | 0.99461 | TRUE | [M-H]- | 66.2614 | 5.9108 | 6.058 | 0.14718 | 5.9129 | 6.0547 | 0.14175 | -1.1351 | 3.0274 | 5.3882 | 5.5642 | 0 | 20 | 15 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Fail | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 143.2081 | No Control Sample | 0 | off | off | 143.2081 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9737 | -1.9426 | -0.63519 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 276.7397 | FALSE | 0.00625 | 0.008411
1487 | 1488 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 6:2 FTS | 0.029934085740043 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7.92 | 7.92 | Manual | Baseline | 287.7575 | 15844.1194 | 0.018162 | 83.1566 | 4034.4779 | 0.020611 | 3.4604 | 3.9272 | 83.5766 | 4035.5344 | 0.97774 | 0.99493 | 7.9641 | 7.9586 | 0.55625 | 0.044055 | 7.8701 | 7.817 | 8.0561 | 8.1394 | 0.18607 | 0.32238 | 0.043336 | 0.057318 | 55.7186 | 89.4416 | 0 | 0.00059986 | TRUE | 1.0007 | TRUE | [M-H]- | 119.7363 | 7.8786 | 8.0486 | 0.17004 | 7.9204 | 8.0395 | 0.11915 | 0 | 0.97396 | 1.6193 | 5.1806 | 0 | 16 | 5 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 287.7575 | No Control Sample | 0 | off | off | 287.7575 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9697 | 4.1125 | 1.7559 | 406.9614 | -0.30672 | -0.12482 | 0 | 0 | 0 | 0 | 0 | 95.8602 | FALSE | 0.00625 | 0.018162
1489 | 1490 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 8:2 FTS | 0.0211728542643413 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.88 | 8.89 | Manual | Baseline | 197.6867 | 19029.3727 | 0.010389 | 50.9814 | 3687.5784 | 0.013825 | 3.8776 | 5.1604 | 51.4875 | 3687.5806 | 0.97167 | 0.99767 | 8.9083 | 8.9099 | 0.31853 | 0.028285 | 8.8363 | 8.7365 | 8.9855 | 9.1672 | 0.14915 | 0.43073 | 0.061861 | 0.080505 | 1082.3649 | 146.6934 | 0.01015 | 1.4024e-06 | TRUE | 0.99982 | TRUE | [M-H]- | 84.6914 | 8.8479 | 8.9752 | 0.12727 | 8.8541 | 8.9679 | 0.11386 | -3.4694 | 0.96033 | 0.89453 | 6.7145 | 0 | 9 | 5 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Fail | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 197.6867 | No Control Sample | 0 | off | off | 197.6867 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9626 | 1.9863 | 1.0467 | 506.9598 | 8.1389 | 4.1261 | 0 | 0 | 0 | 0 | 0 | 106.9691 | FALSE | 0.00625 | 0.010389
1491 | 1492 | PFAS CS2 0.025ng/ml | 11 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 12) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 10:2 FTS | 0 | Quantifiers | 10/26/2023 21:53:42 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.5 | 8.89 | 0 | Baseline | 0 | 19029.3727 | 0 | 0 | 3687.5784 | 0 | 0 | 5.1604 | 64.3129 | 3687.5806 | 0 | 0.99767 | 0 | 8.9099 | 0 | 0 | 0 | 8.7365 | 0 | 9.1672 | 0 | 0.43073 | 0 | 0.080505 | 0 | 146.6934 | 0 | 1.4024e-06 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.00625 | 0
1552 | 1553 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFBA | 0.042314560755034 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 1.49 | 1.48 | Manual | Baseline | 834.424 | 109934.8604 | 0.0075902 | 171.251 | 16441.917 | 0.010416 | 4.8725 | 6.6863 | 179.012 | 16479.5409 | 0.44485 | 0.98737 | 1.457 | 1.4617 | 2.2131 | 0.032975 | 1.3847 | 1.3131 | 1.5706 | 1.6882 | 0.1859 | 0.37508 | 0.069784 | 0.10165 | 21.027 | 265.5811 | 0 | 0.0056669 | TRUE | 0.99682 | TRUE | [M-H]- | 169.2582 | 1.3878 | 1.568 | 0.18021 | 1.3933 | 1.5654 | 0.17205 | 0 | 1.0437 | 1.1287 | 1.1761 | 0 | 39 | 15 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Fail | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 834.424 | No Control Sample | 0 | off | off | 834.424 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9763 | -13.7337 | -2.925 | 168.988 | -5.6665 | -0.95758 | 0 | 0 | 0 | 0 | 0 | 206.9977 | FALSE | 0.00625 | 0.0075902
1554 | 1555 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFPeA | 0.0472947578252629 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy, Integration Quality | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 3.18 | 3.21 | Manual | Baseline | 768.5816 | 72665.6101 | 0.010577 | 92.5394 | 7895.5417 | 0.01172 | 8.3055 | 9.2034 | 96.6549 | 8020.7158 | 0.61249 | 0.94863 | 3.1978 | 3.1949 | 0.55858 | 0.017763 | 3.0179 | 3.0017 | 3.3999 | 3.4019 | 0.38193 | 0.40013 | 0.12832 | 0.14169 | 19.2927 | 548.614 | 0.0052486 | 0.013035 | TRUE | 1.0009 | TRUE | [M-H]- | 189.179 | 3.0193 | 3.3974 | 0.37809 | 3.0254 | 3.3945 | 0.36901 | 1.2717 | 1.0537 | 1.13 | 0.87365 | 0 | 68 | 25 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 768.5816 | No Control Sample | 0 | off | off | 768.5816 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9813 | 19.9933 | 5.2578 | 218.9847 | -5.1324 | -1.1239 | 0 | 0 | 0 | 0 | 0 | 51.5881 | FALSE | 0.00625 | 0.010577
1556 | 1557 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFHxA | 0.0359977530559435 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0.61533 | 6.13 | 6.12 | Baseline | Baseline | 1254.3166 | 149721.352 | 0.0083777 | 220.9156 | 27556.2803 | 0.0080169 | 5.6778 | 5.4333 | 222.4119 | 27569.9848 | 0.84604 | 0.99607 | 6.1153 | 6.1226 | 0.23961 | 0.014688 | 6.0099 | 5.9609 | 6.2454 | 6.3183 | 0.23542 | 0.35748 | 0.098768 | 0.084618 | 49.6949 | 304.6923 | 0.010578 | 0.00097245 | FALSE | 0.99881 | TRUE | [M-H]- | 143.991 | 6.0232 | 6.2189 | 0.1957 | 6.0313 | 6.2003 | 0.16902 | -9.9268 | 1.2102 | 1.3217 | 1.3142 | 0 | 26 | 12 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1254.3166 | No Control Sample | 203843.5852 | off | off | 1254.3166 | 203843.5852 | off | 40 | off | 0.025 | off | 0.00015383 | off | 312.9734 | 1.9239 | 0.60213 | 268.9832 | 3.1447 | 0.84588 | 0 | 0 | 0 | 0 | 0 | 4070.3145 | FALSE | 0.00625 | 0.0083777
1558 | 1559 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFHpA | 0.0458343912636302 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0.83369 | 7.26 | 7.27 | Baseline | Baseline | 1903.5506 | 165975.6614 | 0.011469 | 429.8743 | 38626.5985 | 0.011129 | 4.4282 | 4.2969 | 435.5179 | 38647.0193 | 0.88935 | 0.99602 | 7.274 | 7.2754 | 0.19342 | 0.014042 | 7.1729 | 7.128 | 7.3942 | 7.4399 | 0.22138 | 0.31193 | 0.069321 | 0.066542 | 108.2887 | 179.292 | 0.013989 | 0.0010939 | FALSE | 0.99981 | TRUE | [M-H]- | 183.3376 | 7.1916 | 7.352 | 0.16045 | 7.2174 | 7.3444 | 0.12699 | -27.1624 | 1.077 | 1.6117 | 1.6685 | 0 | 26 | 10 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1903.5506 | No Control Sample | 228328.7703 | off | off | 1903.5506 | 228328.7703 | off | 40 | off | 0.025 | off | 0.00020842 | off | 362.9738 | 11.5063 | 4.1764 | 318.9815 | 7.4387 | 2.3728 | 0 | 0 | 0 | 0 | 0 | 14.8378 | FALSE | 0.00625 | 0.011469
1560 | 1561 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFOA | 0.0380446911971807 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0.83032 | 7.96 | 7.94 | Baseline | Baseline | 1589.8541 | 172311.5865 | 0.0092266 | 399.1321 | 41394.4209 | 0.0096422 | 3.9833 | 4.1627 | 409.8973 | 41405.0481 | 0.70989 | 0.99689 | 7.9627 | 7.9653 | 0.034447 | 0.002742 | 7.8738 | 7.8257 | 8.111 | 8.1661 | 0.23713 | 0.34045 | 0.055062 | 0.063534 | 45.7893 | 122.8828 | 0.041023 | 0.00059774 | FALSE | 0.99968 | TRUE | [M-H]- | 152.1788 | 7.8915 | 8.0565 | 0.16503 | 7.9037 | 8.0355 | 0.13183 | -69.0481 | 1.2001 | 1.3317 | 1.9112 | 0 | 18 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1589.8541 | No Control Sample | 191475.4056 | off | off | 1589.8541 | 191475.4056 | off | 40 | off | 0.025 | off | 0.00020758 | off | 412.9699 | 8.3844 | 3.4625 | 368.9788 | 7.3584 | 2.7151 | 0 | 0 | 0 | 0 | 0 | 11.6659 | FALSE | 0.00625 | 0.0092266
1562 | 1563 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFNA | 0.0298400483054612 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0.59716 | 8.46 | 8.47 | Baseline | Baseline | 950.339 | 137093.1435 | 0.0069321 | 226.0874 | 27509.7944 | 0.0082184 | 4.2034 | 4.9834 | 229.5707 | 27512.7642 | 0.80643 | 0.99844 | 8.4803 | 8.4752 | 0.24031 | 0.02033 | 8.3557 | 8.2475 | 8.5817 | 8.7098 | 0.226 | 0.46223 | 0.065135 | 0.077774 | 42.9753 | 99.5536 | 0.011046 | 0.00022191 | FALSE | 1.0006 | TRUE | [M-H]- | 119.3602 | 8.4055 | 8.5468 | 0.14128 | 8.4166 | 8.5399 | 0.1233 | 11.0507 | 1.008 | 1.0899 | 1.6915 | 0 | 13 | 5 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Fail | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 950.339 | No Control Sample | 159143.3885 | off | off | 950.339 | 159143.3885 | off | 40 | off | 0.025 | off | 0.00014929 | off | 462.9666 | 7.3269 | 3.3921 | 418.9755 | 5.5726 | 2.3348 | 0 | 0 | 0 | 0 | 0 | 25.6651 | FALSE | 0.00625 | 0.0069321
1564 | 1565 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFDA | 0.0252892637186339 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.87 | 8.86 | Baseline | Baseline | 856.8559 | 136487.1799 | 0.0062779 | 159.0957 | 28392.0141 | 0.0056035 | 5.3858 | 4.8072 | 159.0957 | 28395.9853 | 0.90592 | 0.99853 | 8.896 | 8.8932 | 0.29331 | 0.026016 | 8.7028 | 8.6742 | 9.0502 | 9.1357 | 0.34742 | 0.46152 | 0.087116 | 0.074194 | 31.9366 | 107.0452 | 6.8834e-09 | 0.00028493 | FALSE | 1.0003 | TRUE | [M-H]- | 101.1571 | 8.8178 | 8.9756 | 0.15774 | 8.8259 | 8.9681 | 0.14218 | 3.1522e-06 | 1.1553 | 1.3635 | 1.9553 | 0 | 18 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 856.8559 | No Control Sample | 0 | off | off | 856.8559 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9615 | 2.8872 | 1.481 | 468.9724 | 5.0555 | 2.3709 | 0 | 0 | 0 | 0 | 0 | 2964.8141 | FALSE | 0.00625 | 0.0062779
1566 | 1567 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFUdA | 0.0288692661700145 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.21 | 9.2 | Baseline | Baseline | 656.3311 | 102515.8607 | 0.0064022 | 149.8841 | 22597.1611 | 0.0066329 | 4.3789 | 4.5367 | 149.8841 | 22600.7097 | 0.92003 | 0.99813 | 9.225 | 9.2293 | 0.16265 | 0.01498 | 9.0511 | 9.0783 | 9.3942 | 9.4163 | 0.3431 | 0.338 | 0.06991 | 0.071052 | 55.0971 | 125.2693 | 8.144e-10 | 0.00030949 | FALSE | 0.99953 | TRUE | [M-H]- | 115.4771 | 9.1607 | 9.3026 | 0.1419 | 9.1698 | 9.2942 | 0.12444 | 3.5577e-07 | 1.0743 | 1.184 | 1.4171 | 0 | 19 | 5 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 656.3311 | No Control Sample | 0 | off | off | 656.3311 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9612 | 7.7591 | 4.3681 | 518.9715 | 8.0093 | 4.1566 | 0 | 0 | 0 | 0 | 0 | 11.9735 | FALSE | 0.00625 | 0.0064022
1568 | 1569 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFDoA | 0.02065344817971 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.47 | 9.47 | Manual | Baseline | 426.2096 | 79744.9562 | 0.0053447 | 84.8284 | 19308.8407 | 0.0043932 | 5.0244 | 4.13 | 86.1268 | 19331.6824 | 0.90976 | 0.99303 | 9.5101 | 9.5027 | 0.42338 | 0.040094 | 9.4294 | 9.367 | 9.5938 | 9.6163 | 0.16443 | 0.2493 | 0.082774 | 0.065402 | 62.0916 | 411.5342 | 0 | 0.0020457 | TRUE | 1.0008 | TRUE | [M-H]- | 82.6138 | 9.4354 | 9.5841 | 0.14871 | 9.4432 | 9.5771 | 0.13396 | 0 | 0.83039 | 0.63833 | 1.4078 | 0 | 12 | 7 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 426.2096 | No Control Sample | 0 | off | off | 426.2096 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9517 | -3.1299 | -1.9185 | 568.9639 | 0.87434 | 0.49747 | 0 | 0 | 0 | 0 | 0 | 14857.4699 | FALSE | 0.00625 | 0.0053447
1570 | 1571 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFTrDA | 0.0206117363872999 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.68 | 9.47 | Manual | Baseline | 310.0491 | 79744.9562 | 0.003888 | 93.3676 | 19308.8407 | 0.0048355 | 3.3207 | 4.13 | 94.5611 | 19331.6824 | 0.73681 | 0.99303 | 9.7217 | 9.5027 | 0.43049 | 0.041672 | 9.6501 | 9.367 | 9.7897 | 9.6163 | 0.13967 | 0.2493 | 0.050025 | 0.065402 | 42.4232 | 411.5342 | 0.0050894 | 0.0020457 | TRUE | 1.023 | TRUE | [M-H]- | 82.4469 | 9.6581 | 9.7939 | 0.13586 | 9.6684 | 9.7709 | 0.10257 | -3.4023 | 1.1285 | 1.0552 | 0.66693 | 0 | 14 | 6 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 310.0491 | No Control Sample | 0 | off | off | 310.0491 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9583 | 11.7712 | 7.8037 | 618.9662 | 7.2744 | 4.5026 | 0 | 0 | 0 | 0 | 0 | 14.1366 | FALSE | 0.00625 | 0.003888
1572 | 1573 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFTeDA | 0.0409110988947903 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.87 | 9.86 | Manual | Baseline | 290.5019 | 20027.8467 | 0.014505 | 97.504 | 5790.2633 | 0.016839 | 2.9794 | 3.4589 | 98.7446 | 5791.461 | 0.53636 | 0.99691 | 9.8961 | 9.8962 | 0.2649 | 0.026146 | 9.8411 | 9.7969 | 9.9501 | 10.0024 | 0.109 | 0.20545 | 0.046343 | 0.054079 | 52.6174 | 869.6377 | 0.0050891 | 0.00038921 | TRUE | 0.99999 | TRUE | [M-H]- | 163.6444 | 9.8474 | 9.9534 | 0.10596 | 9.8513 | 9.9408 | 0.089542 | 4.5525 | 1.0989 | 1.0183 | 0.50896 | 0 | 13 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 290.5019 | No Control Sample | 0 | off | off | 290.5019 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9443 | -4.1047 | -2.9264 | 668.9669 | 13.5188 | 9.0435 | 0 | 0 | 0 | 0 | 0 | 11427.8512 | FALSE | 0.00625 | 0.014505
1579 | 1580 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFBS | 0.0364844584449538 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 3.5 | 3.5 | Manual | Manual | 416.5501 | 43134.558 | 0.009657 | 50.832 | 4109.552 | 0.012369 | 8.1946 | 10.4962 | 50.9269 | 4115.9991 | 0.83837 | 0.99166 | 3.8566 | 3.8297 | 10.189 | 0.35662 | 3.6788 | 3.5596 | 3.9781 | 4.1388 | 0.2993 | 0.57919 | 0.19739 | 0.15127 | 22.7336 | 1031.6508 | 0.00070837 | 0.0015175 | TRUE | 1.007 | TRUE | [M-H]- | 145.9378 | 3.686 | 3.9709 | 0.28496 | 3.6905 | 3.9681 | 0.27752 | -0.12031 | 1.3906 | 1.8343 | 7.2507 | 0 | 47 | 32 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 416.5501 | No Control Sample | 0 | off | off | 416.5501 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9453 | 7.8358 | 2.3425 | 79.9561 | -13.4689 | -1.0769 | 0 | 0 | 0 | 0 | 0 | 16.7653 | FALSE | 0.00625 | 0.009657
1581 | 1582 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFPeS | 0.0242957529056503 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Asymmetry, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 6.38 | 3.5 | Manual | Manual | 279.2993 | 43134.558 | 0.0064751 | 61.5956 | 4109.552 | 0.014988 | 4.5344 | 10.4962 | 61.7661 | 4115.9991 | 0.85045 | 0.99166 | 6.3984 | 3.8297 | 0.28898 | 0.018437 | 6.2978 | 3.5596 | 6.4536 | 4.1388 | 0.15574 | 0.57919 | 0.094369 | 0.15127 | 44.8302 | 1031.6508 | 0.010078 | 0.0015175 | TRUE | 1.6707 | TRUE | [M-H]- | 97.183 | 6.3029 | 6.4463 | 0.14339 | 6.3095 | 6.4427 | 0.1332 | -3.9858 | 0.66467 | 0.31428 | 15.1942 | 0 | 18 | 12 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 279.2993 | No Control Sample | 0 | off | off | 279.2993 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9418 | 5.7972 | 2.0229 | 79.9583 | 13.5459 | 1.0831 | 0 | 0 | 0 | 0 | 0 | 5.4356 | FALSE | 0.00625 | 0.0064751
1583 | 1584 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | L-PFHxS | 0.0337868727016643 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.02028 | 4 | 0.00507 | 0 | 7.34 | 7.35 | Baseline | Baseline | 386.8888 | 61301.9867 | 0.0063112 | 119.4253 | 14683.0136 | 0.0081336 | 3.2396 | 4.175 | 119.4253 | 14697.833 | 0.77655 | 0.98823 | 7.3563 | 7.3508 | 0.22195 | 0.016291 | 7.2506 | 7.2333 | 7.4563 | 7.4925 | 0.20572 | 0.25915 | 0.051239 | 0.06453 | 62.8937 | 188.5716 | 3.048e-09 | 0.0022533 | FALSE | 1.0007 | TRUE | [M-H]- | 166.6019 | 7.3005 | 7.4181 | 0.11757 | 7.3053 | 7.4122 | 0.10696 | 1.7694e-06 | 0.95334 | 0.87965 | 15.497 | 0 | 25 | 8 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 386.8888 | No Control Sample | 0 | off | off | 386.8888 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9402 | 9.0596 | 3.6142 | 79.9559 | -15.9415 | -1.2746 | 0 | 0 | 0 | 0 | 0 | 678.1472 | FALSE | 0.00507 | 0.0063112
1585 | 1586 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | Br-PFHxS | 0 | Qualifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy | 0 | 0 | 0.00473 | 4 | 0.0011825 | 0 | 7.16 | 7.35 | 0 | Baseline | 0 | 61301.9867 | 0 | 0 | 14683.0136 | 0 | 0 | 4.175 | 119.4253 | 14697.833 | 0 | 0.98823 | 0 | 7.3508 | 0 | 0 | 0 | 7.2333 | 0 | 7.4925 | 0 | 0.25915 | 0 | 0.06453 | 0 | 188.5716 | 0 | 0.0022533 | TRUE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 386.8888 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.0011825 | 0
1587 | 1588 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFHpS | 0.0322224646272311 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7.98 | 7.35 | Manual | Baseline | 300.6287 | 61301.9867 | 0.0049041 | 87.5831 | 14683.0136 | 0.0059649 | 3.4325 | 4.175 | 88.2631 | 14697.833 | 0.91887 | 0.98823 | 7.9876 | 7.3508 | 0.095113 | 0.00759 | 7.9045 | 7.2333 | 8.0584 | 7.4925 | 0.1539 | 0.25915 | 0.049643 | 0.06453 | 54.947 | 188.5716 | 0.0050641 | 0.0022533 | TRUE | 1.0866 | TRUE | [M-H]- | 128.8899 | 7.9134 | 8.0489 | 0.13553 | 7.9214 | 8.0436 | 0.12219 | 2.8819 | 0.927 | 0.8779 | 17.5563 | 0 | 13 | 5 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 300.6287 | No Control Sample | 0 | off | off | 300.6287 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9364 | 6.6112 | 2.968 | 79.9578 | 7.9939 | 0.63917 | 0 | 0 | 0 | 0 | 0 | 3.674 | FALSE | 0.00625 | 0.0049041
1589 | 1590 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | L-PFOS | 0.0355812905914331 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.0197 | 4 | 0.004925 | 0 | 8.46 | 8.45 | Manual | Baseline | 371.601 | 40179.0646 | 0.0092486 | 83.0483 | 8382.8264 | 0.009907 | 4.4745 | 4.793 | 83.8914 | 8383.8539 | 0.74153 | 0.99769 | 8.4732 | 8.4755 | 0.1564 | 0.013231 | 8.4031 | 8.3126 | 8.5977 | 8.6416 | 0.19457 | 0.32903 | 0.064725 | 0.076179 | 38.7037 | 123.2789 | 0 | 0.00025731 | TRUE | 0.99973 | TRUE | [M-H]- | 180.6157 | 8.4103 | 8.5858 | 0.17547 | 8.4179 | 8.5687 | 0.15082 | 0 | 1.259 | 1.4277 | 9.5986 | 0 | 12 | 5 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 371.601 | No Control Sample | 0 | off | off | 371.601 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9332 | 5.988 | 2.9876 | 79.9578 | 6.348 | 0.50757 | 0 | 0 | 0 | 0 | 0 | 11.6943 | FALSE | 0.004925 | 0.0092486
1591 | 1592 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | Br-PFOS | 0.00255502936685632 | Qualifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.00528 | 4 | 0.00132 | 0 | 8.3 | 8.45 | Manual | Baseline | 57.1259 | 40179.0646 | 0.0014218 | 15.6777 | 8382.8264 | 0.0018702 | 3.6438 | 4.793 | 83.8914 | 8383.8539 | 0.114 | 0.99769 | 8.292 | 8.4755 | 0.095809 | 0.0079522 | 8.2458 | 8.3126 | 8.3947 | 8.6416 | 0.14889 | 0.32903 | 0.045975 | 0.076179 | 7.3064 | 123.2789 | 0 | 0.00025731 | TRUE | 0.97836 | TRUE | [M-H]- | 48.3907 | 8.2506 | 8.3785 | 0.12784 | 8.2569 | 8.3746 | 0.11763 | 0 | 1.5998 | 2.4946 | 0.15373 | 0 | 10 | 5 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 57.1259 | No Control Sample | 0 | off | off | 371.601 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9332 | 5.988 | 2.9876 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 11.6943 | FALSE | 0.00132 | 0.0014218
1593 | 1594 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFNS | 0.03322647502786 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.85 | 8.45 | Manual | Baseline | 295.8995 | 40179.0646 | 0.0073645 | 57.8977 | 8382.8264 | 0.0069067 | 5.1107 | 4.793 | 58.3607 | 8383.8539 | 0.72865 | 0.99769 | 8.8798 | 8.4755 | 0.33639 | 0.02977 | 8.7866 | 8.3126 | 8.966 | 8.6416 | 0.17939 | 0.32903 | 0.082565 | 0.076179 | 16.2134 | 123.2789 | 0.0050653 | 0.00025731 | TRUE | 1.0477 | TRUE | [M-H]- | 132.9059 | 8.793 | 8.9527 | 0.15968 | 8.8031 | 8.9473 | 0.14419 | 1.6348 | 0.81927 | 0.65063 | 12.7239 | 0 | 10 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Fail | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 295.8995 | No Control Sample | 0 | off | off | 295.8995 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9296 | 4.6336 | 2.5435 | 79.9612 | 51.4141 | 4.1109 | 0 | 0 | 0 | 0 | 0 | 14.5878 | FALSE | 0.00625 | 0.0073645
1595 | 1596 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | PFDS | 0.043150123716851 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.19 | 8.45 | Manual | Baseline | 394.8263 | 40179.0646 | 0.0098267 | 67.9858 | 8382.8264 | 0.0081101 | 5.8075 | 4.793 | 68.5046 | 8383.8539 | 0.93081 | 0.99769 | 9.208 | 8.4755 | 0.19631 | 0.018041 | 9.0957 | 8.3126 | 9.3296 | 8.6416 | 0.23392 | 0.32903 | 0.084867 | 0.076179 | 46.3559 | 123.2789 | 0.0050635 | 0.00025731 | TRUE | 1.0864 | TRUE | [M-H]- | 172.6005 | 9.107 | 9.3191 | 0.21217 | 9.1197 | 9.3079 | 0.18825 | -1.4716 | 1.0195 | 1.0605 | 10.0949 | 0 | 14 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Fail | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 394.8263 | No Control Sample | 0 | off | off | 394.8263 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9264 | 4.3165 | 2.5853 | 79.9554 | -18.4452 | -1.4748 | 0 | 0 | 0 | 0 | 0 | 5.7158 | FALSE | 0.00625 | 0.0098267
1599 | 1600 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FBSA | 0.0176108110824991 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 4.6 | 8.86 | Manual | Baseline | 342.6915 | 73799.0823 | 0.0046436 | 51.9573 | 15451.6394 | 0.0033626 | 6.5956 | 4.7761 | 52.3506 | 15458.2047 | 0.80928 | 0.99625 | 4.9304 | 9.0168 | 7.1831 | 0.33042 | 4.8146 | 8.8149 | 5.0565 | 9.2367 | 0.24192 | 0.42182 | 0.12761 | 0.073901 | 31.1727 | 128.9174 | 0.0089051 | 0.00083606 | TRUE | 0.5468 | TRUE | [M-H]- | 70.4432 | 4.817 | 5.0532 | 0.23626 | 4.8211 | 5.0516 | 0.23051 | -1.9126 | 1.4356 | 1.9482 | 4.109 | 0 | 44 | 25 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 342.6915 | No Control Sample | 0 | off | off | 342.6915 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9621 | 10.5905 | 3.1555 | 77.9641 | -15.6907 | -1.2233 | 0 | 0 | 0 | 0 | 0 | 5.5689 | FALSE | 0.00625 | 0.0046436
1601 | 1602 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FPeSA | 0.0305277217580736 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7 | 8.86 | Baseline | Baseline | 922.4964 | 73799.0823 | 0.0125 | 208.5521 | 15451.6394 | 0.013497 | 4.4233 | 4.7761 | 208.5521 | 15458.2047 | 0.98897 | 0.99625 | 7.097 | 9.0168 | 1.3861 | 0.097029 | 6.9511 | 8.8149 | 7.2307 | 9.2367 | 0.27967 | 0.42182 | 0.069745 | 0.073901 | 178.6659 | 128.9174 | 0 | 0.00083606 | FALSE | 0.78709 | TRUE | [M-H]- | 122.1109 | 7.0193 | 7.1657 | 0.14635 | 7.0251 | 7.1587 | 0.13365 | 0 | 0.93293 | 0.83925 | 3.2437 | 0 | 30 | 9 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 922.4964 | No Control Sample | 0 | off | off | 922.4964 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9574 | 4.5598 | 1.5866 | 77.9669 | 19.9298 | 1.5538 | 0 | 0 | 0 | 0 | 0 | 9.7001 | FALSE | 0.00625 | 0.0125
1603 | 1604 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FHxSA | 0.035700011480111 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7.86 | 8.86 | Baseline | Baseline | 538.3417 | 73799.0823 | 0.0072947 | 129.1614 | 15451.6394 | 0.0083591 | 4.168 | 4.7761 | 129.6505 | 15458.2047 | 0.88518 | 0.99625 | 7.9787 | 9.0168 | 1.51 | 0.11869 | 7.8747 | 8.8149 | 8.1286 | 9.2367 | 0.25393 | 0.42182 | 0.061596 | 0.073901 | 31.0794 | 128.9174 | 0.0097325 | 0.00083606 | FALSE | 0.88487 | TRUE | [M-H]- | 142.8 | 7.9149 | 8.0945 | 0.17959 | 7.9232 | 8.0752 | 0.15194 | 4.9504 | 1.547 | 2.0573 | 7.4454 | 0 | 19 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 538.3417 | No Control Sample | 0 | off | off | 538.3417 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9552 | 6.6727 | 2.6554 | 77.9648 | -8.2081 | -0.63995 | 0 | 0 | 0 | 0 | 0 | 0.48462 | FALSE | 0.00625 | 0.0072947
1605 | 1606 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FHpSA | 0.0249242421256404 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.58 | 8.86 | Baseline | Baseline | 778.4832 | 73799.0823 | 0.010549 | 163.527 | 15451.6394 | 0.010583 | 4.7606 | 4.7761 | 163.527 | 15458.2047 | 1 | 0.99625 | 8.5775 | 9.0168 | 0.02871 | 0.0024633 | 8.377 | 8.8149 | 8.7621 | 9.2367 | 0.38505 | 0.42182 | 0.077664 | 0.073901 | 187.396 | 128.9174 | 0 | 0.00083606 | FALSE | 0.95128 | TRUE | [M-H]- | 99.697 | 8.5077 | 8.6478 | 0.14008 | 8.5162 | 8.6423 | 0.12617 | 0 | 0.92534 | 0.87643 | 4.1536 | 0 | 22 | 7 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 778.4832 | No Control Sample | 0 | off | off | 778.4832 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9517 | 5.1026 | 2.2857 | 77.9656 | 2.472 | 0.19273 | 0 | 0 | 0 | 0 | 0 | 5.0102 | FALSE | 0.00625 | 0.010549
1607 | 1608 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | FOSA | 0.0390608064125905 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.88 | 8.86 | Baseline | Baseline | 655.9 | 73799.0823 | 0.0088876 | 151.3146 | 15451.6394 | 0.0097928 | 4.3347 | 4.7761 | 151.3193 | 15458.2047 | 0.96459 | 0.99625 | 9.014 | 9.0168 | 1.5092 | 0.13402 | 8.8259 | 8.8149 | 9.1501 | 9.2367 | 0.32425 | 0.42182 | 0.063047 | 0.073901 | 57.2169 | 128.9174 | 5.3267e-05 | 0.00083606 | FALSE | 0.99969 | TRUE | [M-H]- | 156.2432 | 8.9247 | 9.0999 | 0.17523 | 8.9536 | 9.088 | 0.1344 | 0.024857 | 0.96752 | 1.1786 | 5.0396 | 0 | 17 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 655.9 | No Control Sample | 0 | off | off | 655.9 | 0 | off | 0 | off | 0 | off | 0 | off | 497.948 | 3.6772 | 1.8311 | 77.9648 | -7.4902 | -0.58398 | 0 | 0 | 0 | 0 | 0 | 5.3241 | FALSE | 0.00625 | 0.0088876
1626 | 1627 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 4:2 FTS | 0.0201985528603571 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 5.94 | 5.92 | Manual | Baseline | 197.5 | 19257.3943 | 0.010256 | 52.0015 | 3394.9673 | 0.015317 | 3.798 | 5.6723 | 52.7128 | 3398.5228 | 0.96665 | 0.99318 | 5.94 | 5.9286 | 0.00025396 | 1.5085e-05 | 5.8659 | 5.7671 | 6.0228 | 6.0904 | 0.15688 | 0.32338 | 0.046669 | 0.086424 | 21367.5297 | 183.2945 | 0.015282 | 0.0020877 | TRUE | 1.0019 | TRUE | [M-H]- | 80.7942 | 5.8707 | 6.0166 | 0.14596 | 5.8752 | 6.0109 | 0.13569 | -5.0653 | 1.1408 | 1.2835 | 4.24 | 0 | 21 | 7 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Fail | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 197.5 | No Control Sample | 0 | off | off | 197.5 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9762 | 5.9095 | 1.9323 | 306.9646 | -9.7302 | -2.9868 | 0 | 0 | 0 | 0 | 0 | 150.9147 | FALSE | 0.00625 | 0.010256
1628 | 1629 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 6:2 FTS | 0.0132511881528294 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy, Asymmetry, Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 7.92 | 7.92 | Manual | Baseline | 124.3422 | 15465.763 | 0.0080398 | 38.393 | 3835.4255 | 0.01001 | 3.2387 | 4.0323 | 38.7828 | 3835.4271 | 0.96585 | 0.9984 | 7.9404 | 7.9254 | 0.25698 | 0.020353 | 7.8342 | 7.7465 | 7.9912 | 8.1541 | 0.15704 | 0.40755 | 0.042131 | 0.059426 | 70.506 | 87.7753 | 0 | 9.5347e-07 | TRUE | 1.0019 | TRUE | [M-H]- | 53.0048 | 7.8417 | 7.9856 | 0.14387 | 7.8487 | 7.981 | 0.13238 | 0 | 0.72486 | 0.43463 | 10.6698 | 0 | 15 | 5 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 124.3422 | No Control Sample | 0 | off | off | 124.3422 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9725 | 10.8199 | 4.6197 | 406.9731 | 28.5077 | 11.6015 | 0 | 0 | 0 | 0 | 0 | 9969.0036 | FALSE | 0.00625 | 0.0080398
1630 | 1631 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 8:2 FTS | 0.0225447973749808 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Total Width | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 8.88 | 8.89 | Manual | Baseline | 214.5484 | 19395.697 | 0.011062 | 46.0366 | 3791.4846 | 0.012142 | 4.6604 | 5.1156 | 46.733 | 3794.0172 | 0.96825 | 0.99573 | 8.8833 | 8.8894 | 0.036847 | 0.003272 | 8.8111 | 8.7354 | 8.9854 | 9.0835 | 0.17427 | 0.34817 | 0.074967 | 0.079375 | 606.2476 | 131.1381 | 0.015303 | 0.001532 | TRUE | 0.99931 | TRUE | [M-H]- | 90.1792 | 8.819 | 8.9716 | 0.15265 | 8.8285 | 8.9623 | 0.13374 | -4.0427 | 1.4604 | 2.131 | 6.7716 | 0 | 10 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Fail | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 214.5484 | No Control Sample | 0 | off | off | 214.5484 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9643 | 5.2515 | 2.7674 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 90.5318 | FALSE | 0.00625 | 0.011062
1632 | 1633 | PFAS CS2 0.025ng/ml | 12 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 13) | CS2 | 0 | 0 | Standard | 0 | 0 | 0 | 3 | 1 | 10:2 FTS | 0 | Quantifiers | 10/26/2023 22:10:16 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0.025 | 4 | 0.00625 | 0 | 9.5 | 8.89 | 0 | Baseline | 0 | 19395.697 | 0 | 0 | 3791.4846 | 0 | 0 | 5.1156 | 19.9471 | 3794.0172 | 0 | 0.99573 | 0 | 8.8894 | 0 | 0 | 0 | 8.7354 | 0 | 9.0835 | 0 | 0.34817 | 0 | 0.079375 | 0 | 131.1381 | 0 | 0.001532 | FALSE | 0 | TRUE | [M-H]- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 0 | No Control Sample | 0 | off | off | 0 | 0 | off | 0 | off | 0 | off | 0 | off | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | FALSE | 0.00625 | 0
1693 | 1694 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFBA | 0.116742828682196 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 1.49 | 1.48 | Manual | Baseline | 2352.6359 | 112347.3887 | 0.020941 | 396.6967 | 15850.7757 | 0.025027 | 5.9306 | 7.0878 | 407.6813 | 15874.5134 | 0.69572 | 0.98991 | 1.467 | 1.4666 | 1.5456 | 0.023029 | 1.3848 | 1.2925 | 1.6162 | 1.6976 | 0.23136 | 0.40508 | 0.10385 | 0.10735 | 39.5754 | 216.0398 | 0.010331 | 0.0036161 | TRUE | 1.0002 | TRUE | [M-H]- | 116.7428 | 1.3886 | 1.6069 | 0.21835 | 1.3922 | 1.6029 | 0.21061 | -17.7136 | 1.6866 | 2.4488 | 1.6571 | 0 | 48 | 22 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Fail | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2352.6359 | No Control Sample | 0 | off | off | 2352.6359 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9806 | 6.3983 | 1.3627 | 168.9875 | -8.6313 | -1.4586 | 0 | 0 | 0 | 0 | 0 | 55.5687 | FALSE | 0.025 | 0.020941
1695 | 1696 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFPeA | 0.106844717170572 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Integration Quality | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 3.18 | 3.21 | Manual | Baseline | 1887.8104 | 79005.4918 | 0.023895 | 204.6035 | 9060.1768 | 0.022583 | 9.2267 | 8.7201 | 210.9016 | 9240.9177 | 0.77689 | 0.93279 | 3.2205 | 3.2163 | 1.2738 | 0.040507 | 3.0697 | 3.017 | 3.4221 | 3.4091 | 0.35238 | 0.39212 | 0.14041 | 0.13326 | 46.5248 | 549.3571 | 0.01036 | 0.026149 | TRUE | 1.0013 | TRUE | [M-H]- | 106.8447 | 3.0736 | 3.4175 | 0.34386 | 3.0802 | 3.4095 | 0.32924 | -6.0151 | 0.94266 | 0.87296 | 1.6245 | 0 | 61 | 26 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1887.8104 | No Control Sample | 0 | off | off | 1887.8104 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9791 | 11.8921 | 3.1273 | 218.9867 | 4.0758 | 0.89254 | 0 | 0 | 0 | 0 | 0 | 5.484 | FALSE | 0.025 | 0.023895
1697 | 1698 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFHxA | 0.125313395021628 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 2.3314 | 6.13 | 6.12 | Baseline | Baseline | 4752.3114 | 162951.8732 | 0.029164 | 884.2333 | 29934.5394 | 0.029539 | 5.3745 | 5.4436 | 884.3799 | 29939.5339 | 0.95439 | 0.9974 | 6.1552 | 6.1558 | 0.41057 | 0.025168 | 6.0018 | 5.9706 | 6.2874 | 6.3661 | 0.28563 | 0.39547 | 0.078355 | 0.083216 | 109.089 | 254.5975 | 0.00029783 | 0.00035646 | FALSE | 0.9999 | TRUE | [M-H]- | 125.3134 | 6.0479 | 6.2523 | 0.20444 | 6.0667 | 6.2397 | 0.17304 | 0.922 | 0.9822 | 1.0295 | 1.2525 | 0 | 31 | 10 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4752.3114 | No Control Sample | 203843.5852 | off | off | 4752.3114 | 203843.5852 | off | 10 | off | 0.1 | off | 0.0023314 | off | 312.9749 | 6.8258 | 2.1363 | 268.9841 | 6.254 | 1.6822 | 0 | 0 | 0 | 0 | 0 | 1665.9158 | FALSE | 0.025 | 0.029164
1699 | 1700 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFHpA | 0.111034911513277 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 2.1708 | 7.26 | 7.27 | Baseline | Baseline | 4956.4896 | 178396.4075 | 0.027784 | 1200.3758 | 41897.5085 | 0.02865 | 4.1291 | 4.2579 | 1218.0561 | 41919.1426 | 0.89036 | 0.99589 | 7.3019 | 7.3029 | 0.57718 | 0.041903 | 7.2108 | 7.1469 | 7.3939 | 7.4478 | 0.18315 | 0.3009 | 0.065947 | 0.066615 | 117.5738 | 209.3065 | 0.0074235 | 0.00084659 | FALSE | 0.99987 | TRUE | [M-H]- | 111.0349 | 7.2376 | 7.3679 | 0.13032 | 7.2437 | 7.3581 | 0.1144 | 48.6538 | 0.91397 | 0.75629 | 1.6936 | 0 | 22 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Pass | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4956.4896 | No Control Sample | 228328.7703 | off | off | 4956.4896 | 228328.7703 | off | 10 | off | 0.1 | off | 0.0021708 | off | 362.9721 | 6.819 | 2.4751 | 318.9811 | 6.4164 | 2.0467 | 0 | 0 | 0 | 0 | 0 | 9.1758 | FALSE | 0.025 | 0.027784
1701 | 1702 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFOA | 0.108475225996121 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 2.5477 | 7.96 | 7.94 | Baseline | Baseline | 4878.1638 | 185428.7604 | 0.026307 | 1106.9909 | 41917.7252 | 0.026409 | 4.4067 | 4.4236 | 1130.1797 | 41933.5379 | 0.86118 | 0.99655 | 7.9853 | 7.9854 | 0.31804 | 0.025316 | 7.8609 | 7.8265 | 8.0955 | 8.1663 | 0.23462 | 0.33987 | 0.069062 | 0.068932 | 81.9078 | 109.5591 | 0.0022477 | 0.00072561 | FALSE | 0.99999 | TRUE | [M-H]- | 108.4752 | 7.9206 | 8.0687 | 0.14806 | 7.9297 | 8.0591 | 0.12938 | 10.6053 | 1.1582 | 1.3607 | 1.7625 | 0 | 18 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4878.1638 | No Control Sample | 191475.4056 | off | off | 4878.1638 | 191475.4056 | off | 10 | off | 0.1 | off | 0.0025477 | off | 412.9689 | 6.105 | 2.5212 | 368.9771 | 2.7023 | 0.99708 | 0 | 0 | 0 | 0 | 0 | 3.9408 | FALSE | 0.025 | 0.026307
1703 | 1704 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFNA | 0.107164253332586 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 2.3963 | 8.46 | 8.47 | Baseline | Baseline | 3813.4756 | 153181.9932 | 0.024895 | 800.2227 | 32106.7257 | 0.024924 | 4.7655 | 4.771 | 800.4244 | 32131.5354 | 0.98059 | 0.99435 | 8.4963 | 8.4944 | 0.42953 | 0.036338 | 8.319 | 8.2794 | 8.6147 | 8.6717 | 0.2957 | 0.39235 | 0.073858 | 0.074259 | 108.309 | 168.2092 | 0.00023102 | 0.0013996 | FALSE | 1.0002 | TRUE | [M-H]- | 107.1643 | 8.4125 | 8.5737 | 0.16116 | 8.4213 | 8.5624 | 0.14113 | 0.62518 | 0.97085 | 0.90093 | 1.0829 | 0 | 17 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3813.4756 | No Control Sample | 159143.3885 | off | off | 3813.4756 | 159143.3885 | off | 10 | off | 0.1 | off | 0.0023963 | off | 462.9658 | 5.6341 | 2.6084 | 418.9774 | 9.9156 | 4.1544 | 0 | 0 | 0 | 0 | 0 | 5.4738 | FALSE | 0.025 | 0.024895
1705 | 1706 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFDA | 0.0929294171438869 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.87 | 8.86 | Baseline | Baseline | 3136.0233 | 135939.7079 | 0.023069 | 681.9057 | 27679.0942 | 0.024636 | 4.5989 | 4.9113 | 688.1223 | 27694.7723 | 0.92959 | 0.99536 | 8.9078 | 8.9038 | 0.426 | 0.037786 | 8.8014 | 8.7491 | 9.0152 | 9.0421 | 0.21377 | 0.293 | 0.07241 | 0.076585 | 76.4925 | 178.6552 | 0.016292 | 0.001056 | FALSE | 1.0004 | TRUE | [M-H]- | 92.9294 | 8.8327 | 8.9885 | 0.15581 | 8.8438 | 8.9789 | 0.13512 | 51.9708 | 1.0615 | 1.1689 | 1.8259 | 0 | 11 | 5 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Fail | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3136.0233 | No Control Sample | 0 | off | off | 3136.0233 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9636 | 6.9654 | 3.573 | 468.9712 | 2.5971 | 1.218 | 0 | 0 | 0 | 0 | 0 | 1191.7764 | FALSE | 0.025 | 0.023069
1707 | 1708 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFUdA | 0.0958030952537553 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.21 | 9.2 | Baseline | Baseline | 2197.772 | 103444.377 | 0.021246 | 426.1169 | 22424.0371 | 0.019003 | 5.1577 | 4.6131 | 426.1169 | 22425.6697 | 0.94296 | 0.9989 | 9.2392 | 9.2404 | 0.31754 | 0.029246 | 9.0347 | 9.0807 | 9.4126 | 9.4512 | 0.37788 | 0.37047 | 0.081471 | 0.072786 | 66.5285 | 153.5555 | 5.6871e-10 | 0.0001711 | FALSE | 0.99988 | TRUE | [M-H]- | 95.8031 | 9.1539 | 9.3187 | 0.16478 | 9.169 | 9.3083 | 0.13933 | -6.4131e-07 | 1.0983 | 1.3237 | 1.8015 | 0 | 21 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2197.772 | No Control Sample | 0 | off | off | 2197.772 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9615 | 8.3311 | 4.69 | 518.9698 | 4.8071 | 2.4947 | 0 | 0 | 0 | 0 | 0 | 10.691 | FALSE | 0.025 | 0.021246
1709 | 1710 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFDoA | 0.0786813721198559 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.47 | 9.47 | Baseline | Baseline | 1835.6073 | 90153.0487 | 0.020361 | 472.7082 | 22719.6675 | 0.020806 | 3.8832 | 3.9681 | 472.8364 | 22722.5228 | 0.96446 | 0.99824 | 9.5055 | 9.5048 | 0.37466 | 0.03548 | 9.3798 | 9.3344 | 9.6403 | 9.6843 | 0.2605 | 0.34987 | 0.058848 | 0.061787 | 93.0632 | 364.1585 | 0.00053015 | 0.00025676 | FALSE | 1.0001 | TRUE | [M-H]- | 78.6814 | 9.4394 | 9.5745 | 0.13516 | 9.4491 | 9.5629 | 0.11384 | 0.96202 | 0.95311 | 0.85999 | 1.5852 | 0 | 18 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1835.6073 | No Control Sample | 0 | off | off | 1835.6073 | 0 | off | 0 | off | 0 | off | 0 | off | 612.959 | 8.7182 | 5.3439 | 568.9686 | 9.076 | 5.1639 | 0 | 0 | 0 | 0 | 0 | 4131.948 | FALSE | 0.025 | 0.020361
1711 | 1712 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFTrDA | 0.0911591366252093 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.68 | 9.47 | Baseline | Baseline | 1550.22 | 90153.0487 | 0.017195 | 426.5909 | 22719.6675 | 0.018776 | 3.634 | 3.9681 | 426.5909 | 22722.5228 | 0.98774 | 0.99824 | 9.72 | 9.5048 | 0.41285 | 0.039964 | 9.5542 | 9.3344 | 9.8423 | 9.6843 | 0.28817 | 0.34987 | 0.058092 | 0.061787 | 328.8954 | 364.1585 | 4.2827e-13 | 0.00025676 | FALSE | 1.0226 | TRUE | [M-H]- | 91.1591 | 9.6606 | 9.7772 | 0.11664 | 9.6688 | 9.7701 | 0.10125 | 6.3399e-10 | 1.0282 | 1.0889 | 0.7337 | 0 | 25 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1550.22 | No Control Sample | 0 | off | off | 1550.22 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9578 | 11.1511 | 7.3926 | 618.9665 | 7.8161 | 4.8379 | 0 | 0 | 0 | 0 | 0 | 10.2207 | FALSE | 0.025 | 0.017195
1713 | 1714 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFTeDA | 0.0881864630809908 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Integration Quality, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.87 | 9.86 | Baseline | Baseline | 661.916 | 21170.3152 | 0.031266 | 191.6289 | 6712.761 | 0.028547 | 3.4542 | 3.1537 | 191.6318 | 6719.6976 | 0.74376 | 0.99093 | 9.8949 | 9.8952 | 0.25213 | 0.024885 | 9.8132 | 9.7952 | 9.9858 | 9.9953 | 0.17267 | 0.20008 | 0.053866 | 0.048864 | 82.4173 | 559.6697 | 2.7886e-05 | 0.0020342 | FALSE | 0.99997 | TRUE | [M-H]- | 88.1865 | 9.8466 | 9.9597 | 0.11316 | 9.8518 | 9.9537 | 0.10182 | 0.030949 | 0.99434 | 0.97252 | 1.7391 | 0 | 19 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 661.916 | No Control Sample | 0 | off | off | 661.916 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9515 | 5.9234 | 4.2231 | 668.97 | 18.0448 | 12.0712 | 0 | 0 | 0 | 0 | 0 | 2828.9753 | FALSE | 0.025 | 0.031266
1720 | 1721 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFBS | 0.115803654837705 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 3.5 | 3.5 | Manual | Manual | 1397.3345 | 45587.3318 | 0.030652 | 151.586 | 4075.9428 | 0.03719 | 9.2181 | 11.1845 | 152.1504 | 4083.7449 | 0.93879 | 0.99208 | 3.874 | 3.8649 | 10.6852 | 0.37398 | 3.6691 | 3.6146 | 4.0762 | 4.2138 | 0.40716 | 0.5992 | 0.10813 | 0.16161 | 59.824 | 818.4599 | 0.0019263 | 0.001012 | TRUE | 1.0023 | TRUE | [M-H]- | 115.8037 | 3.664 | 4.0719 | 0.4079 | 3.6691 | 4.0643 | 0.39521 | -0.71716 | 0.86404 | 0.71147 | 9.5511 | 0 | 64 | 18 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1397.3345 | No Control Sample | 0 | off | off | 1397.3345 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9448 | 6.0562 | 1.8105 | 79.957 | -2.265 | -0.1811 | 0 | 0 | 0 | 0 | 0 | 5.2009 | FALSE | 0.025 | 0.030652
1722 | 1723 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFPeS | 0.106904933931005 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 6.38 | 3.5 | Baseline | Manual | 1298.8412 | 45587.3318 | 0.028491 | 232.7104 | 4075.9428 | 0.057094 | 5.5814 | 11.1845 | 236.6821 | 4083.7449 | 0.9582 | 0.99208 | 6.4074 | 3.8649 | 0.43004 | 0.027437 | 6.2986 | 3.6146 | 6.487 | 4.2138 | 0.18835 | 0.5992 | 0.093831 | 0.16161 | 129.0313 | 818.4599 | 0.027768 | 0.001012 | FALSE | 1.6579 | TRUE | [M-H]- | 106.9049 | 6.3203 | 6.4771 | 0.15683 | 6.3257 | 6.4742 | 0.14852 | 34.3085 | 0.86029 | 0.73154 | 13.7702 | 0 | 21 | 12 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1298.8412 | No Control Sample | 0 | off | off | 1298.8412 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9413 | 4.2103 | 1.4692 | 79.9562 | -13.0561 | -1.0439 | 0 | 0 | 0 | 0 | 0 | 6.1535 | FALSE | 0.025 | 0.028491
1724 | 1725 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | L-PFHxS | 0.0892434754306053 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.0811 | 4 | 0.020275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 1119.2339 | 67139.9143 | 0.01667 | 301.2032 | 16986.3403 | 0.017732 | 3.7159 | 3.9526 | 306.4002 | 17002.289 | 0.73217 | 0.99055 | 7.3758 | 7.3743 | 0.48787 | 0.03581 | 7.2934 | 7.2607 | 7.4642 | 7.5345 | 0.17077 | 0.27372 | 0.057771 | 0.060284 | 65.506 | 183.1249 | 0.033937 | 0.00053885 | FALSE | 1.0002 | TRUE | [M-H]- | 110.0413 | 7.3183 | 7.4389 | 0.12058 | 7.3242 | 7.4319 | 0.1077 | -59.8588 | 0.99533 | 0.96795 | 15.0943 | 0 | 21 | 9 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1119.2339 | No Control Sample | 0 | off | off | 1119.2339 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9389 | 5.6459 | 2.2524 | 79.9582 | 12.1724 | 0.97327 | 0 | 0 | 0 | 0 | 0 | 301.9032 | FALSE | 0.020275 | 0.01667
1726 | 1727 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | Br-PFHxS | 0.0150108397569986 | Qualifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.0189 | 4 | 0.004725 | 0 | 7.16 | 7.35 | Manual | Baseline | 255.7721 | 67139.9143 | 0.0038095 | 57.4327 | 16986.3403 | 0.0033811 | 4.4534 | 3.9526 | 306.4002 | 17002.289 | 0.17051 | 0.99055 | 7.1826 | 7.3743 | 0.31601 | 0.022627 | 7.1161 | 7.2607 | 7.2546 | 7.5345 | 0.13846 | 0.27372 | 0.085029 | 0.060284 | 12.4905 | 183.1249 | 0 | 0.00053885 | TRUE | 0.974 | TRUE | [M-H]- | 79.4224 | 7.119 | 7.259 | 0.14004 | 7.1221 | 7.2502 | 0.12811 | 0 | 1.0765 | 1.0682 | 0.22852 | 0 | 17 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 255.7721 | No Control Sample | 0 | off | off | 1119.2339 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9389 | 5.6459 | 2.2524 | 79.9568 | -5.218 | -0.41722 | 0 | 0 | 0 | 0 | 0 | 301.9032 | FALSE | 0.004725 | 0.0038095
1728 | 1729 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFHpS | 0.0901806797166274 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7.98 | 7.35 | Baseline | Baseline | 921.4918 | 67139.9143 | 0.013725 | 198.9253 | 16986.3403 | 0.011711 | 4.6324 | 3.9526 | 198.9253 | 17002.289 | 0.96356 | 0.99055 | 8.0024 | 7.3743 | 0.28117 | 0.022437 | 7.8635 | 7.2607 | 8.1942 | 7.5345 | 0.33072 | 0.27372 | 0.079971 | 0.060284 | 59.7973 | 183.1249 | 9.1841e-13 | 0.00053885 | FALSE | 1.0852 | TRUE | [M-H]- | 90.1807 | 7.943 | 8.0891 | 0.14607 | 7.9519 | 8.0811 | 0.12925 | -5.5242e-10 | 1.2612 | 1.6342 | 22.7571 | 0 | 24 | 8 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 921.4918 | No Control Sample | 0 | off | off | 921.4918 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9356 | 4.832 | 2.1692 | 79.9571 | -1.5417 | -0.12327 | 0 | 0 | 0 | 0 | 0 | 0.25412 | FALSE | 0.025 | 0.013725
1730 | 1731 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | L-PFOS | 0.0911435826075196 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Asymmetry, Total Width | 0 | 0 | 0.0788 | 4 | 0.0197 | 0 | 8.46 | 8.45 | Manual | Baseline | 985.364 | 41592.5252 | 0.023691 | 170.6085 | 8954.6279 | 0.019053 | 5.7756 | 4.6448 | 171.3175 | 8958.2545 | 0.65977 | 0.99471 | 8.4908 | 8.4925 | 0.36408 | 0.030801 | 8.3953 | 8.3489 | 8.5731 | 8.6938 | 0.17786 | 0.34493 | 0.092148 | 0.072542 | 24.9255 | 143.7597 | 0.010092 | 0.00063797 | TRUE | 0.9998 | TRUE | [M-H]- | 115.6644 | 8.3785 | 8.5641 | 0.18557 | 8.3785 | 8.5574 | 0.17884 | -9.6805 | 0.76459 | 0.47372 | 12.7844 | 0 | 12 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 985.364 | No Control Sample | 0 | off | off | 985.364 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9322 | 4.0548 | 2.0231 | 79.9581 | 9.4812 | 0.75809 | 0 | 0 | 0 | 0 | 0 | 1.8411 | FALSE | 0.0197 | 0.023691
1732 | 1733 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | Br-PFOS | 0.0155477533271369 | Qualifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.0211 | 4 | 0.005275 | 0 | 8.3 | 8.45 | Manual | Baseline | 359.849 | 41592.5252 | 0.0086518 | 78.2946 | 8954.6279 | 0.0087435 | 4.5961 | 4.6448 | 171.3175 | 8958.2545 | 0.24227 | 0.99471 | 8.3542 | 8.4925 | 0.65327 | 0.054221 | 8.2267 | 8.3489 | 8.3995 | 8.6938 | 0.17278 | 0.34493 | 0.064437 | 0.072542 | 11.4386 | 143.7597 | 0.0089512 | 0.00063797 | TRUE | 0.98371 | TRUE | [M-H]- | 73.686 | 8.2401 | 8.4161 | 0.17595 | 8.2512 | 8.4161 | 0.16483 | -4.0562 | 0.73465 | 0.51733 | 0.36519 | 0 | 13 | 5 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 359.849 | No Control Sample | 0 | off | off | 985.364 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9322 | 4.0548 | 2.0231 | 79.9583 | 12.4032 | 0.99172 | 0 | 0 | 0 | 0 | 0 | 1.8411 | FALSE | 0.005275 | 0.0086518
1734 | 1735 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFNS | 0.120957754848982 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.85 | 8.45 | Baseline | Baseline | 1115.088 | 41592.5252 | 0.02681 | 224.9342 | 8954.6279 | 0.025119 | 4.9574 | 4.6448 | 224.9369 | 8958.2545 | 0.93567 | 0.99471 | 8.8899 | 8.4925 | 0.45068 | 0.039885 | 8.7256 | 8.3489 | 9.0571 | 8.6938 | 0.33147 | 0.34493 | 0.074129 | 0.072542 | 39.4942 | 143.7597 | 2.3535e-05 | 0.00063797 | FALSE | 1.0468 | TRUE | [M-H]- | 120.9578 | 8.7992 | 8.9824 | 0.18314 | 8.8135 | 8.97 | 0.15657 | -0.015971 | 1.0141 | 1.0579 | 15.9827 | 0 | 17 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1115.088 | No Control Sample | 0 | off | off | 1115.088 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9303 | 5.9698 | 3.277 | 79.9562 | -11.7979 | -0.94332 | 0 | 0 | 0 | 0 | 0 | 3.3186 | FALSE | 0.025 | 0.02681
1736 | 1737 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFDS | 0.105569582607253 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.19 | 8.45 | Baseline | Baseline | 999.9501 | 41592.5252 | 0.024042 | 220.6786 | 8954.6279 | 0.024644 | 4.5313 | 4.6448 | 220.7211 | 8958.2545 | 0.88908 | 0.99471 | 9.221 | 8.4925 | 0.33765 | 0.03103 | 9.0959 | 8.3489 | 9.366 | 8.6938 | 0.27012 | 0.34493 | 0.070696 | 0.072542 | 37.1746 | 143.7597 | 0.00041324 | 0.00063797 | FALSE | 1.0858 | TRUE | [M-H]- | 105.5696 | 9.1537 | 9.3066 | 0.15291 | 9.16 | 9.2922 | 0.1322 | 0.33761 | 1.2827 | 1.4822 | 15.5976 | 0 | 15 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 999.9501 | No Control Sample | 0 | off | off | 999.9501 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9259 | 3.4706 | 2.0786 | 79.9564 | -6.0093 | -0.48049 | 0 | 0 | 0 | 0 | 0 | 1.7321 | FALSE | 0.025 | 0.024042
1740 | 1741 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FBSA | 0.0817685815812704 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 4.6 | 8.86 | Manual | Baseline | 1613.6605 | 74843.2582 | 0.021561 | 184.1524 | 16477.8433 | 0.011176 | 8.7626 | 4.5421 | 185.0824 | 16480.4344 | 0.93983 | 0.99809 | 4.9622 | 9.0292 | 7.8739 | 0.3622 | 4.7862 | 8.8559 | 5.145 | 9.2589 | 0.35883 | 0.403 | 0.15103 | 0.070429 | 65.6859 | 179.7119 | 0 | 0.00035921 | TRUE | 0.54957 | TRUE | [M-H]- | 81.7686 | 4.7951 | 5.1409 | 0.34585 | 4.8073 | 5.1241 | 0.31677 | 0 | 1.173 | 1.3431 | 5.0077 | 0 | 65 | 29 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1613.6605 | No Control Sample | 0 | off | off | 1613.6605 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9609 | 6.3539 | 1.8932 | 77.9663 | 13.3116 | 1.0378 | 0 | 0 | 0 | 0 | 0 | 6.3821 | FALSE | 0.025 | 0.021561
1742 | 1743 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FPeSA | 0.0959202176115174 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7 | 8.86 | Baseline | Baseline | 2939.559 | 74843.2582 | 0.039276 | 716.7828 | 16477.8433 | 0.0435 | 4.101 | 4.5421 | 733.2999 | 16480.4344 | 0.91885 | 0.99809 | 7.1183 | 9.0292 | 1.6896 | 0.11827 | 7.0409 | 8.8559 | 7.2211 | 9.2589 | 0.1802 | 0.403 | 0.065977 | 0.070429 | 170.009 | 179.7119 | 0.025062 | 0.00035921 | FALSE | 0.78836 | TRUE | [M-H]- | 95.9202 | 7.0547 | 7.1884 | 0.13376 | 7.0624 | 7.1813 | 0.11894 | -99.6876 | 1.0785 | 1.1914 | 3.6525 | 0 | 20 | 9 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2939.559 | No Control Sample | 0 | off | off | 2939.559 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9579 | 5.9948 | 2.0859 | 77.9648 | -5.8984 | -0.45987 | 0 | 0 | 0 | 0 | 0 | 1.4563 | FALSE | 0.025 | 0.039276
1744 | 1745 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FHxSA | 0.138376962579098 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7.86 | 8.86 | Baseline | Baseline | 2116.1927 | 74843.2582 | 0.028275 | 473.4784 | 16477.8433 | 0.028734 | 4.4695 | 4.5421 | 473.4785 | 16480.4344 | 0.98227 | 0.99809 | 8.007 | 9.0292 | 1.8699 | 0.14697 | 7.8617 | 8.8559 | 8.1448 | 9.2589 | 0.28308 | 0.403 | 0.069721 | 0.070429 | 67.175 | 179.7119 | 2.5351e-07 | 0.00035921 | FALSE | 0.88679 | TRUE | [M-H]- | 138.377 | 7.9396 | 8.0827 | 0.14312 | 7.9484 | 8.0756 | 0.1272 | 0.00042402 | 0.94752 | 0.90511 | 7.3064 | 0 | 21 | 6 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2116.1927 | No Control Sample | 0 | off | off | 2116.1927 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9548 | 5.6778 | 2.2595 | 77.9645 | -11.283 | -0.87968 | 0 | 0 | 0 | 0 | 0 | 4.1268 | FALSE | 0.025 | 0.028275
1746 | 1747 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FHpSA | 0.10641411637856 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.58 | 8.86 | Baseline | Baseline | 3370.7631 | 74843.2582 | 0.045038 | 704.1981 | 16477.8433 | 0.042736 | 4.7867 | 4.5421 | 705.7823 | 16480.4344 | 0.97712 | 0.99809 | 8.5888 | 9.0292 | 0.10278 | 0.0088187 | 8.457 | 8.8559 | 8.7613 | 9.2589 | 0.30428 | 0.403 | 0.075465 | 0.070429 | 102.67 | 179.7119 | 0.0013994 | 0.00035921 | FALSE | 0.95123 | TRUE | [M-H]- | 106.4141 | 8.5125 | 8.6626 | 0.15003 | 8.5201 | 8.6499 | 0.12983 | 3.2387 | 1.068 | 1.07 | 4.0326 | 0 | 18 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3370.7631 | No Control Sample | 0 | off | off | 3370.7631 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9515 | 4.7975 | 2.149 | 77.965 | -4.9374 | -0.38494 | 0 | 0 | 0 | 0 | 0 | 5.2691 | FALSE | 0.025 | 0.045038
1748 | 1749 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FOSA | 0.119617210506591 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.88 | 8.86 | Baseline | Baseline | 2037.0037 | 74843.2582 | 0.027217 | 406.7731 | 16477.8433 | 0.024686 | 5.0077 | 4.5421 | 406.8157 | 16480.4344 | 0.98796 | 0.99809 | 9.0276 | 9.0292 | 1.6627 | 0.14764 | 8.8668 | 8.8559 | 9.1915 | 9.2589 | 0.3247 | 0.403 | 0.07737 | 0.070429 | 71.6865 | 179.7119 | 0.00022412 | 0.00035921 | FALSE | 0.99983 | TRUE | [M-H]- | 119.6172 | 8.9398 | 9.1079 | 0.16811 | 8.9527 | 9.0944 | 0.14175 | 0.28076 | 1.078 | 1.1763 | 6.0878 | 0 | 17 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2037.0037 | No Control Sample | 0 | off | off | 2037.0037 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9484 | 4.3949 | 2.1884 | 77.9658 | 5.7016 | 0.44453 | 0 | 0 | 0 | 0 | 0 | 3.2817 | FALSE | 0.025 | 0.027217
1767 | 1768 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 4:2 FTS | 0.0854232485294213 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 5.94 | 5.92 | Baseline | Baseline | 819.4484 | 18892.798 | 0.043374 | 161.0534 | 3336.6596 | 0.048268 | 5.0881 | 5.6622 | 163.0063 | 3344.5992 | 0.93871 | 0.98665 | 5.972 | 5.9644 | 0.53908 | 0.032021 | 5.8515 | 5.7819 | 6.0514 | 6.1102 | 0.1998 | 0.32833 | 0.080599 | 0.089138 | 75.4566 | 184.1823 | 0.006788 | 0.004147 | FALSE | 1.0013 | TRUE | [M-H]- | 85.4232 | 5.8707 | 6.0406 | 0.16993 | 5.882 | 6.0361 | 0.1541 | 5.4716 | 0.79027 | 0.6019 | 4.9953 | 0 | 24 | 11 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Fail | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 819.4484 | No Control Sample | 0 | off | off | 819.4484 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9774 | 9.4397 | 3.0865 | 306.9679 | 0.98861 | 0.30347 | 0 | 0 | 0 | 0 | 0 | 3709.0285 | FALSE | 0.025 | 0.043374
1769 | 1770 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 6:2 FTS | 0.0914012531149605 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7.92 | 7.92 | Baseline | Baseline | 886.5895 | 15987.4022 | 0.055456 | 244.072 | 3779.1317 | 0.064584 | 3.6325 | 4.2304 | 244.072 | 3781.5873 | 1 | 0.99514 | 7.9418 | 7.9469 | 0.27588 | 0.021849 | 7.7835 | 7.8037 | 8.1325 | 8.0749 | 0.3489 | 0.27122 | 0.051746 | 0.063909 | 137.7115 | 104.8469 | 0 | 0.0012065 | FALSE | 0.99936 | TRUE | [M-H]- | 91.4013 | 7.8754 | 8.0141 | 0.13871 | 7.8829 | 8.0016 | 0.11868 | 0 | 1.0496 | 1.0271 | 5.5411 | 0 | 26 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 886.5895 | No Control Sample | 0 | off | off | 886.5895 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9711 | 7.5877 | 3.2397 | 406.9638 | 5.5956 | 2.2772 | 0 | 0 | 0 | 0 | 0 | 3686.1362 | FALSE | 0.025 | 0.055456
1771 | 1772 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 8:2 FTS | 0.0848661020106249 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.88 | 8.89 | Baseline | Baseline | 822.5519 | 19754.0179 | 0.04164 | 181.1906 | 3967.2335 | 0.045672 | 4.5397 | 4.9793 | 181.1906 | 3967.2336 | 1 | 0.9988 | 8.9055 | 8.9014 | 0.2876 | 0.025539 | 8.67 | 8.7156 | 9.1159 | 9.125 | 0.44587 | 0.4094 | 0.066464 | 0.077629 | 131.5319 | 206.4778 | 0 | 1.4773e-08 | FALSE | 1.0005 | TRUE | [M-H]- | 84.8661 | 8.8101 | 8.9764 | 0.16625 | 8.8234 | 8.9681 | 0.14471 | 0 | 0.81448 | 0.62936 | 5.8053 | 0 | 23 | 5 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 822.5519 | No Control Sample | 0 | off | off | 822.5519 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9656 | 7.835 | 4.1288 | 506.9531 | -5.0595 | -2.5649 | 0 | 0 | 0 | 0 | 0 | 4406.9617 | FALSE | 0.025 | 0.04164
1773 | 1774 | PFAS CS3 0.1ng/ml | 13 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 14) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 10:2 FTS | 0.103790307952679 | Quantifiers | 10/26/2023 22:26:52 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.5 | 8.89 | Manual | Baseline | 1047.2377 | 19754.0179 | 0.053014 | 285.9098 | 3967.2335 | 0.072068 | 3.6628 | 4.9793 | 287.3538 | 3967.2336 | 0.95884 | 0.9988 | 9.5183 | 8.9014 | 0.19229 | 0.018268 | 9.4443 | 8.7156 | 9.5799 | 9.125 | 0.13559 | 0.4094 | 0.058508 | 0.077629 | 163.9402 | 206.4778 | 0 | 1.4773e-08 | TRUE | 1.0693 | TRUE | [M-H]- | 103.7903 | 9.4557 | 9.5694 | 0.11372 | 9.4634 | 9.5653 | 0.10188 | 0 | 0.9874 | 1.0426 | 2.8721 | 0 | 11 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Fail | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1047.2377 | No Control Sample | 0 | off | off | 1047.2377 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9583 | 5.0386 | 3.159 | 606.9524 | 7.389 | 4.4848 | 0 | 0 | 0 | 0 | 0 | 3.8674 | FALSE | 0.025 | 0.053014
1834 | 1835 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFBA | 0.118412189790028 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy, Integration Quality | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 1.49 | 1.48 | Manual | Baseline | 2384.9805 | 112286.3297 | 0.02124 | 316.2097 | 15567.996 | 0.020312 | 7.5424 | 7.2126 | 329.5333 | 15682.0725 | 0.69677 | 0.96661 | 1.4602 | 1.4576 | 2.0011 | 0.029816 | 1.3494 | 1.3084 | 1.6306 | 1.6425 | 0.2812 | 0.33412 | 0.12913 | 0.11053 | 29.7758 | 229.3461 | 0.036668 | 0.018385 | TRUE | 1.0017 | TRUE | [M-H]- | 118.4122 | 1.357 | 1.6259 | 0.26897 | 1.3617 | 1.6214 | 0.25968 | 41.234 | 1.7532 | 2.6095 | 1.3298 | 0 | 57 | 28 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Fail | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2384.9805 | No Control Sample | 0 | off | off | 2384.9805 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9799 | 3.286 | 0.69984 | 168.9882 | -4.6129 | -0.77954 | 0 | 0 | 0 | 0 | 0 | 49.8489 | FALSE | 0.025 | 0.02124
1836 | 1837 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFPeA | 0.112292466608013 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 3.18 | 3.21 | Manual | Baseline | 2084.4943 | 83004.5799 | 0.025113 | 246.731 | 8504.0694 | 0.029013 | 8.4484 | 9.7606 | 251.3761 | 8557.0615 | 0.82628 | 0.97574 | 3.2122 | 3.1971 | 1.0119 | 0.032179 | 3.0828 | 2.9721 | 3.4182 | 3.4446 | 0.33534 | 0.47248 | 0.14673 | 0.15158 | 56.9825 | 524.901 | 0.0051197 | 0.0046295 | TRUE | 1.0047 | TRUE | [M-H]- | 112.2925 | 3.0909 | 3.3993 | 0.30834 | 3.0992 | 3.3747 | 0.27545 | -3.7669 | 1.5398 | 2.001 | 1.5297 | 0 | 59 | 28 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2084.4943 | No Control Sample | 0 | off | off | 2084.4943 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9769 | 3.258 | 0.85678 | 218.9872 | 6.4877 | 1.4207 | 0 | 0 | 0 | 0 | 0 | 5.484 | FALSE | 0.025 | 0.025113
1838 | 1839 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFHxA | 0.117341797638649 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 2.2193 | 6.13 | 6.12 | Baseline | Baseline | 4523.8446 | 165655.8991 | 0.027309 | 843.2 | 29768.8953 | 0.028325 | 5.3651 | 5.5647 | 844.5894 | 29782.3766 | 0.9365 | 0.99591 | 6.1533 | 6.1521 | 0.37999 | 0.023293 | 6.0114 | 5.9709 | 6.277 | 6.3753 | 0.26558 | 0.40445 | 0.080703 | 0.08655 | 109.3556 | 285.6495 | 0.0030965 | 0.00054918 | FALSE | 1.0002 | TRUE | [M-H]- | 117.3418 | 6.0535 | 6.2518 | 0.19837 | 6.0711 | 6.2407 | 0.1696 | 9.8312 | 1.0044 | 1.0905 | 1.314 | 0 | 29 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4523.8446 | No Control Sample | 203843.5852 | off | off | 4523.8446 | 203843.5852 | off | 10 | off | 0.1 | off | 0.0022193 | off | 312.9744 | 5.1323 | 1.6063 | 268.9838 | 5.1755 | 1.3921 | 0 | 0 | 0 | 0 | 0 | 1489.0259 | FALSE | 0.025 | 0.027309
1840 | 1841 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFHpA | 0.121229577181363 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 2.3032 | 7.26 | 7.27 | Baseline | Baseline | 5258.9627 | 173365.6132 | 0.030335 | 1293.7247 | 42037.8837 | 0.030775 | 4.065 | 4.124 | 1293.9645 | 42124.2855 | 0.9634 | 0.98954 | 7.3001 | 7.2975 | 0.55211 | 0.040083 | 7.1997 | 7.1645 | 7.4174 | 7.4115 | 0.2178 | 0.24708 | 0.065173 | 0.064405 | 191.0379 | 305.4347 | 0.00020243 | 0.0036937 | FALSE | 1.0004 | TRUE | [M-H]- | 121.2296 | 7.2307 | 7.3661 | 0.13539 | 7.2414 | 7.3606 | 0.11918 | -1.2024 | 0.99917 | 1.0901 | 1.6765 | 0 | 26 | 10 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5258.9627 | No Control Sample | 228328.7703 | off | off | 5258.9627 | 228328.7703 | off | 10 | off | 0.1 | off | 0.0023032 | off | 362.9718 | 6.0488 | 2.1955 | 318.9819 | 8.8973 | 2.838 | 0 | 0 | 0 | 0 | 0 | 5.0827 | FALSE | 0.025 | 0.030335
1842 | 1843 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFOA | 0.116291420240759 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 2.6384 | 7.96 | 7.94 | Baseline | Baseline | 5051.958 | 179127.9432 | 0.028203 | 1079.8688 | 41555.5261 | 0.025986 | 4.6783 | 4.3106 | 1081.2783 | 41562.7644 | 0.89184 | 0.99742 | 7.9846 | 7.983 | 0.30911 | 0.024605 | 7.8857 | 7.8114 | 8.1207 | 8.1918 | 0.235 | 0.3803 | 0.073383 | 0.065995 | 67.4235 | 149.04 | 0.0022116 | 0.00036924 | FALSE | 1.0002 | TRUE | [M-H]- | 116.2914 | 7.9149 | 8.0694 | 0.1545 | 7.9243 | 8.0591 | 0.13473 | -10.1625 | 1.1509 | 1.3354 | 1.6717 | 0 | 18 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5051.958 | No Control Sample | 191475.4056 | off | off | 5051.958 | 191475.4056 | off | 10 | off | 0.1 | off | 0.0026384 | off | 412.9684 | 4.7089 | 1.9446 | 368.9777 | 4.3512 | 1.6055 | 0 | 0 | 0 | 0 | 0 | 6.7941 | FALSE | 0.025 | 0.028203
1844 | 1845 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFNA | 0.11147713903355 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 2.501 | 8.46 | 8.47 | Baseline | Baseline | 3980.2262 | 153694.6053 | 0.025897 | 820.4254 | 31807.0757 | 0.025794 | 4.8514 | 4.8321 | 827.0443 | 31822.0827 | 0.95525 | 0.99677 | 8.4965 | 8.4945 | 0.43181 | 0.036531 | 8.367 | 8.3048 | 8.5948 | 8.6856 | 0.2278 | 0.3808 | 0.077246 | 0.075736 | 80.6798 | 120.3307 | 0.014356 | 0.00085155 | FALSE | 1.0002 | TRUE | [M-H]- | 111.4771 | 8.4133 | 8.5724 | 0.15907 | 8.4264 | 8.5628 | 0.13641 | 51.7019 | 1.042 | 1.1555 | 1.1819 | 0 | 13 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Fail | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3980.2262 | No Control Sample | 159143.3885 | off | off | 3980.2262 | 159143.3885 | off | 10 | off | 0.1 | off | 0.002501 | off | 462.9653 | 4.419 | 2.0458 | 418.9736 | 0.93357 | 0.39114 | 0 | 0 | 0 | 0 | 0 | 11.9029 | FALSE | 0.025 | 0.025897
1846 | 1847 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFDA | 0.118604078431724 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.87 | 8.86 | Baseline | Baseline | 3952.6711 | 134249.0791 | 0.029443 | 837.38 | 27680.939 | 0.030251 | 4.7203 | 4.8499 | 837.5922 | 27701.9287 | 0.97687 | 0.99595 | 8.9046 | 8.9054 | 0.3903 | 0.03462 | 8.7761 | 8.7447 | 9.0277 | 9.0359 | 0.25158 | 0.29123 | 0.07465 | 0.0758 | 98.1626 | 145.0018 | 0.00047346 | 0.0013076 | FALSE | 0.99991 | TRUE | [M-H]- | 118.6041 | 8.8306 | 8.9867 | 0.15609 | 8.8411 | 8.9793 | 0.13825 | 1.5759 | 1.0988 | 1.281 | 1.3961 | 0 | 13 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3952.6711 | No Control Sample | 0 | off | off | 3952.6711 | 0 | off | 0 | off | 0 | off | 0 | off | 512.963 | 5.8262 | 2.9886 | 468.971 | 2.2216 | 1.0419 | 0 | 0 | 0 | 0 | 0 | 756.0209 | FALSE | 0.025 | 0.029443
1848 | 1849 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFUdA | 0.108655575038468 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.21 | 9.2 | Baseline | Baseline | 2523.4039 | 104722.1508 | 0.024096 | 513.9373 | 23057.2289 | 0.02229 | 4.9099 | 4.5418 | 515.1528 | 23062.8333 | 0.94729 | 0.99696 | 9.236 | 9.2368 | 0.28258 | 0.026026 | 9.124 | 9.0931 | 9.3566 | 9.4139 | 0.23253 | 0.3207 | 0.077561 | 0.071096 | 78.9625 | 116.4908 | 0.0042008 | 0.00056481 | FALSE | 0.99991 | TRUE | [M-H]- | 108.6556 | 9.1657 | 9.3217 | 0.15599 | 9.1718 | 9.3131 | 0.14129 | 9.2844 | 1.0192 | 1.0058 | 1.7688 | 0 | 13 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Fail | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2523.4039 | No Control Sample | 0 | off | off | 2523.4039 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9601 | 5.7206 | 3.2205 | 518.9682 | 1.7967 | 0.93243 | 0 | 0 | 0 | 0 | 0 | 20.8819 | FALSE | 0.025 | 0.024096
1850 | 1851 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFDoA | 0.0962376335626257 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.47 | 9.47 | Baseline | Baseline | 2337.7403 | 93869.3667 | 0.024904 | 568.2903 | 23449.3275 | 0.024235 | 4.1136 | 4.0031 | 571.7875 | 23461.802 | 0.96007 | 0.99546 | 9.5114 | 9.5041 | 0.43738 | 0.04142 | 9.4091 | 9.3811 | 9.5937 | 9.657 | 0.1846 | 0.27582 | 0.06724 | 0.062199 | 130.2467 | 238.2312 | 0.011367 | 0.0011007 | FALSE | 1.0008 | TRUE | [M-H]- | 96.2376 | 9.4441 | 9.5709 | 0.12683 | 9.4493 | 9.5636 | 0.11431 | 34.9946 | 0.99679 | 0.9576 | 1.0827 | 0 | 13 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2337.7403 | No Control Sample | 0 | off | off | 2337.7403 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9589 | 8.5262 | 5.2262 | 568.9663 | 5.0408 | 2.868 | 0 | 0 | 0 | 0 | 0 | 3496.567 | FALSE | 0.025 | 0.024904
1852 | 1853 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFTrDA | 0.0927088447054931 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.68 | 9.47 | Baseline | Baseline | 1641.5638 | 93869.3667 | 0.017488 | 474.2151 | 23449.3275 | 0.020223 | 3.4616 | 4.0031 | 475.3778 | 23461.802 | 0.97088 | 0.99546 | 9.7135 | 9.5041 | 0.34653 | 0.033544 | 9.6092 | 9.3811 | 9.8208 | 9.657 | 0.21152 | 0.27582 | 0.052083 | 0.062199 | 139.7425 | 238.2312 | 0.0045568 | 0.0011007 | FALSE | 1.022 | TRUE | [M-H]- | 92.7088 | 9.6566 | 9.7771 | 0.1205 | 9.6654 | 9.7694 | 0.104 | -10.2162 | 1.0215 | 1.07 | 0.84527 | 0 | 19 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1641.5638 | No Control Sample | 0 | off | off | 1641.5638 | 0 | off | 0 | off | 0 | off | 0 | off | 662.952 | 2.3732 | 1.5733 | 618.9635 | 2.9734 | 1.8404 | 0 | 0 | 0 | 0 | 0 | 14.1366 | FALSE | 0.025 | 0.017488
1854 | 1855 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFTeDA | 0.0849406612566645 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.87 | 9.86 | Baseline | Baseline | 747.202 | 24811.2587 | 0.030115 | 231.075 | 7217.8095 | 0.032015 | 3.2336 | 3.4375 | 234.1165 | 7223.9415 | 0.65774 | 0.99352 | 9.8998 | 9.8928 | 0.30195 | 0.029802 | 9.8316 | 9.7936 | 9.9592 | 9.9933 | 0.12757 | 0.19977 | 0.053947 | 0.05474 | 73.6459 | 692.6748 | 0.027274 | 0.0016668 | FALSE | 1.0007 | TRUE | [M-H]- | 84.9407 | 9.8445 | 9.9445 | 0.099945 | 9.8516 | 9.9404 | 0.088829 | -49.4041 | 0.87467 | 0.77441 | 1.3127 | 0 | 14 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 747.202 | No Control Sample | 0 | off | off | 747.202 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9522 | 6.8627 | 4.8927 | 668.959 | 1.7107 | 1.1444 | 0 | 0 | 0 | 0 | 0 | 2882.5608 | FALSE | 0.025 | 0.030115
1861 | 1862 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFBS | 0.0968070704871914 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 3.5 | 3.5 | Manual | Manual | 1145.008 | 44685.5923 | 0.025624 | 120.684 | 4236.6376 | 0.028486 | 9.4877 | 10.5474 | 121.326 | 4248.8689 | 0.9151 | 0.98703 | 3.85 | 3.8408 | 10.0007 | 0.35003 | 3.7084 | 3.6085 | 4.0219 | 4.1365 | 0.31351 | 0.52807 | 0.23355 | 0.16334 | 63.9872 | 989.3801 | 0.001777 | 0 | TRUE | 1.0024 | TRUE | [M-H]- | 96.8071 | 3.7195 | 4.0137 | 0.29426 | 3.7243 | 4.0093 | 0.28492 | 0.68405 | 1.0239 | 1.0527 | 12.0737 | 0 | 50 | 38 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1145.008 | No Control Sample | 0 | off | off | 1145.008 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9455 | 8.4453 | 2.5247 | 79.9565 | -8.8408 | -0.70688 | 0 | 0 | 0 | 0 | 0 | 73.8028 | FALSE | 0.025 | 0.025624
1863 | 1864 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFPeS | 0.114735692782112 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 6.38 | 3.5 | Baseline | Manual | 1366.4074 | 44685.5923 | 0.030578 | 252.8865 | 4236.6376 | 0.05969 | 5.4032 | 10.5474 | 252.8865 | 4248.8689 | 0.98377 | 0.98703 | 6.4076 | 3.8408 | 0.43241 | 0.027588 | 6.2196 | 3.6085 | 6.5523 | 4.1365 | 0.33268 | 0.52807 | 0.077074 | 0.16334 | 107.7702 | 989.3801 | 0 | 0 | FALSE | 1.6683 | TRUE | [M-H]- | 114.7357 | 6.2968 | 6.4857 | 0.18896 | 6.3161 | 6.4812 | 0.16513 | 0 | 0.98857 | 1.1649 | 12.7098 | 0 | 36 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1366.4074 | No Control Sample | 0 | off | off | 1366.4074 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9418 | 5.6597 | 1.9749 | 79.9575 | 3.5932 | 0.2873 | 0 | 0 | 0 | 0 | 0 | 2.0792 | FALSE | 0.025 | 0.030578
1865 | 1866 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | L-PFHxS | 0.107838008150371 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Total Width | 0 | 0 | 0.0811 | 4 | 0.020275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 1314.5373 | 65258.5583 | 0.020144 | 290.6297 | 15773.2379 | 0.018425 | 4.5231 | 4.1373 | 291.2917 | 15796.4969 | 0.83269 | 0.98731 | 7.3761 | 7.3729 | 0.4919 | 0.036105 | 7.283 | 7.2221 | 7.4794 | 7.5178 | 0.1964 | 0.29575 | 0.069662 | 0.064263 | 71.4524 | 177.0438 | 0.0043346 | 0.0026437 | FALSE | 1.0004 | TRUE | [M-H]- | 132.9692 | 7.3027 | 7.4526 | 0.14985 | 7.311 | 7.446 | 0.13496 | -6.4143 | 1.0081 | 1.0436 | 13.2851 | 0 | 24 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1314.5373 | No Control Sample | 0 | off | off | 1314.5373 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9382 | 3.9648 | 1.5817 | 79.9571 | -1.118 | -0.089394 | 0 | 0 | 0 | 0 | 0 | 5.3257 | FALSE | 0.020275 | 0.020144
1867 | 1868 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | Br-PFHxS | 0.0140041969740529 | Qualifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.0189 | 4 | 0.004725 | 0 | 7.16 | 7.35 | Manual | Baseline | 231.9333 | 65258.5583 | 0.0035541 | 49.8679 | 15773.2379 | 0.0031616 | 4.651 | 4.1373 | 291.2917 | 15796.4969 | 0.14692 | 0.98731 | 7.1799 | 7.3729 | 0.27803 | 0.019907 | 7.09 | 7.2221 | 7.2825 | 7.5178 | 0.1926 | 0.29575 | 0.064682 | 0.064263 | 12.2602 | 177.0438 | 0.0059001 | 0.0026437 | TRUE | 0.97383 | TRUE | [M-H]- | 74.0963 | 7.094 | 7.2778 | 0.18385 | 7.1008 | 7.2728 | 0.17192 | -1.5277 | 1.0386 | 1.1053 | 0.17644 | 0 | 22 | 9 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 231.9333 | No Control Sample | 0 | off | off | 1314.5373 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9382 | 3.9648 | 1.5817 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5.3257 | FALSE | 0.004725 | 0.0035541
1869 | 1870 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFHpS | 0.124866864545421 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7.98 | 7.35 | Baseline | Baseline | 1240.1718 | 65258.5583 | 0.019004 | 284.8043 | 15773.2379 | 0.018056 | 4.3545 | 4.1373 | 287.0476 | 15796.4969 | 0.94768 | 0.98731 | 8.0027 | 7.3729 | 0.2841 | 0.022671 | 7.8883 | 7.2221 | 8.0925 | 7.5178 | 0.20412 | 0.29575 | 0.071365 | 0.064263 | 97.0618 | 177.0438 | 0.014261 | 0.0026437 | FALSE | 1.0854 | TRUE | [M-H]- | 124.8669 | 7.9343 | 8.0811 | 0.14677 | 7.9484 | 8.0738 | 0.12548 | 19.8983 | 1.1807 | 1.6062 | 16.108 | 0 | 16 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1240.1718 | No Control Sample | 0 | off | off | 1240.1718 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9359 | 5.4991 | 2.4687 | 79.9565 | -8.1623 | -0.65263 | 0 | 0 | 0 | 0 | 0 | 4.3923 | FALSE | 0.025 | 0.019004
1871 | 1872 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | L-PFOS | 0.0694031183045608 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 0.0788 | 4 | 0.0197 | 0 | 8.46 | 8.45 | Manual | Baseline | 759.555 | 42104.1526 | 0.01804 | 155.6473 | 8571.5049 | 0.018159 | 4.88 | 4.9121 | 156.4334 | 8574.1697 | 0.65168 | 0.99426 | 8.4977 | 8.4912 | 0.4457 | 0.037706 | 8.4096 | 8.3418 | 8.592 | 8.6713 | 0.18235 | 0.32948 | 0.078387 | 0.076989 | 37.2327 | 94.7672 | 0 | 0.00069919 | TRUE | 1.0008 | TRUE | [M-H]- | 88.075 | 8.4107 | 8.5786 | 0.16789 | 8.4312 | 8.5717 | 0.14052 | 0 | 1.007 | 1.2354 | 15.4906 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 759.555 | No Control Sample | 0 | off | off | 759.555 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9326 | 4.6988 | 2.3444 | 79.9567 | -7.5299 | -0.60207 | 0 | 0 | 0 | 0 | 0 | 4.166 | FALSE | 0.0197 | 0.01804
1873 | 1874 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | Br-PFOS | 0.0129847168717654 | Qualifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.0211 | 4 | 0.005275 | 0 | 8.3 | 8.45 | Manual | Baseline | 304.2249 | 42104.1526 | 0.0072255 | 56.4836 | 8571.5049 | 0.0065897 | 5.3861 | 4.9121 | 156.4334 | 8574.1697 | 0.26102 | 0.99426 | 8.3363 | 8.4912 | 0.43694 | 0.036266 | 8.245 | 8.3418 | 8.4096 | 8.6713 | 0.16462 | 0.32948 | 0.095891 | 0.076989 | 13.5115 | 94.7672 | 0.010487 | 0.00069919 | TRUE | 0.98175 | TRUE | [M-H]- | 61.5389 | 8.2544 | 8.4302 | 0.17581 | 8.2587 | 8.4302 | 0.17153 | 3.598 | 0.89384 | 0.82353 | 0.40053 | 0 | 12 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 304.2249 | No Control Sample | 0 | off | off | 759.555 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9326 | 4.6988 | 2.3444 | 79.9533 | -50.1944 | -4.0134 | 0 | 0 | 0 | 0 | 0 | 4.166 | FALSE | 0.005275 | 0.0072255
1875 | 1876 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFNS | 0.0993962301522027 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.85 | 8.45 | Baseline | Baseline | 927.5877 | 42104.1526 | 0.022031 | 178.592 | 8571.5049 | 0.020836 | 5.1939 | 4.9121 | 180.6304 | 8574.1697 | 0.94819 | 0.99426 | 8.892 | 8.4912 | 0.47438 | 0.041983 | 8.7596 | 8.3418 | 8.9924 | 8.6713 | 0.23278 | 0.32948 | 0.080368 | 0.076989 | 69.9903 | 94.7672 | 0.019282 | 0.00069919 | FALSE | 1.0472 | TRUE | [M-H]- | 99.3962 | 8.8005 | 8.9778 | 0.17733 | 8.8113 | 8.9664 | 0.15509 | 14.7932 | 1.0465 | 1.1006 | 20.2087 | 0 | 12 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Fail | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 927.5877 | No Control Sample | 0 | off | off | 927.5877 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9296 | 4.7279 | 2.5952 | 79.958 | 11.7002 | 0.93552 | 0 | 0 | 0 | 0 | 0 | 1.0758 | FALSE | 0.025 | 0.022031
1877 | 1878 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFDS | 0.122005744320698 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.19 | 8.45 | Baseline | Baseline | 1169.848 | 42104.1526 | 0.027785 | 291.1604 | 8571.5049 | 0.033968 | 4.0179 | 4.9121 | 291.1604 | 8574.1697 | 0.97168 | 0.99426 | 9.2201 | 8.4912 | 0.32807 | 0.030149 | 9.033 | 8.3418 | 9.3954 | 8.6713 | 0.36245 | 0.32948 | 0.063524 | 0.076989 | 72.9916 | 94.7672 | 0 | 0.00069919 | FALSE | 1.0858 | TRUE | [M-H]- | 122.0057 | 9.1533 | 9.2839 | 0.13058 | 9.1615 | 9.2775 | 0.11605 | 0 | 0.89077 | 0.78263 | 13.1512 | 0 | 20 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1169.848 | No Control Sample | 0 | off | off | 1169.848 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9267 | 4.7389 | 2.8382 | 79.956 | -11.1683 | -0.89298 | 0 | 0 | 0 | 0 | 0 | 2.9631 | FALSE | 0.025 | 0.027785
1881 | 1882 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FBSA | 0.0807755191496768 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 4.6 | 8.86 | Manual | Baseline | 1641.7634 | 77082.8566 | 0.021299 | 178.4148 | 16126.0885 | 0.011064 | 9.2019 | 4.78 | 181.1457 | 16133.3849 | 0.91263 | 0.99569 | 4.9473 | 9.0293 | 7.5494 | 0.34727 | 4.7692 | 8.8516 | 5.1082 | 9.173 | 0.33903 | 0.32145 | 0.15471 | 0.074002 | 55.1266 | 135.4519 | 0 | 0.0008441 | TRUE | 0.54791 | TRUE | [M-H]- | 80.7755 | 4.7748 | 5.1019 | 0.32713 | 4.779 | 5.0946 | 0.31568 | 0 | 1.0723 | 1.1279 | 4.8197 | 0 | 61 | 29 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1641.7634 | No Control Sample | 0 | off | off | 1641.7634 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9612 | 7.3597 | 2.1929 | 77.9669 | 19.9632 | 1.5564 | 0 | 0 | 0 | 0 | 0 | 1.0031 | FALSE | 0.025 | 0.021299
1883 | 1884 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FPeSA | 0.105846339499655 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7 | 8.86 | Baseline | Baseline | 3340.8193 | 77082.8566 | 0.043341 | 761.8214 | 16126.0885 | 0.047242 | 4.3853 | 4.78 | 764.8355 | 16133.3849 | 0.98384 | 0.99569 | 7.119 | 9.0293 | 1.7 | 0.119 | 7.002 | 8.8516 | 7.2195 | 9.173 | 0.2175 | 0.32145 | 0.067856 | 0.074002 | 198.3996 | 135.4519 | 0.0065748 | 0.0008441 | FALSE | 0.78843 | TRUE | [M-H]- | 105.8463 | 7.0383 | 7.1909 | 0.15258 | 7.0538 | 7.1798 | 0.12606 | 23.0291 | 0.88196 | 0.77487 | 3.5623 | 0 | 24 | 9 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3340.8193 | No Control Sample | 0 | off | off | 3340.8193 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9577 | 5.6155 | 1.9539 | 77.9651 | -2.6662 | -0.20787 | 0 | 0 | 0 | 0 | 0 | 6.7379 | FALSE | 0.025 | 0.043341
1885 | 1886 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FHxSA | 0.114787202618567 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7.86 | 8.86 | Baseline | Baseline | 1807.9649 | 77082.8566 | 0.023455 | 389.9795 | 16126.0885 | 0.024183 | 4.6361 | 4.78 | 393.0638 | 16133.3849 | 0.96042 | 0.99569 | 8.0034 | 9.0293 | 1.8239 | 0.14336 | 7.8866 | 8.8516 | 8.1225 | 9.173 | 0.23587 | 0.32145 | 0.076107 | 0.074002 | 104.8385 | 135.4519 | 0.014965 | 0.0008441 | FALSE | 0.88638 | TRUE | [M-H]- | 114.7872 | 7.9376 | 8.0786 | 0.14107 | 7.943 | 8.0704 | 0.12735 | 24.7436 | 1.0059 | 0.96944 | 8.4727 | 0 | 18 | 8 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1807.9649 | No Control Sample | 0 | off | off | 1807.9649 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9549 | 5.6978 | 2.2675 | 77.9643 | -13.7437 | -1.0715 | 0 | 0 | 0 | 0 | 0 | 2.2319 | FALSE | 0.025 | 0.023455
1887 | 1888 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FHpSA | 0.10628209851446 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.58 | 8.86 | Baseline | Baseline | 3467.3223 | 77082.8566 | 0.044982 | 774.595 | 16126.0885 | 0.048034 | 4.4763 | 4.78 | 780.8026 | 16133.3849 | 0.96515 | 0.99569 | 8.5889 | 9.0293 | 0.10427 | 0.0089463 | 8.4923 | 8.8516 | 8.6995 | 9.173 | 0.20722 | 0.32145 | 0.069684 | 0.074002 | 73.49 | 135.4519 | 0.0030569 | 0.0008441 | FALSE | 0.95123 | TRUE | [M-H]- | 106.2821 | 8.5162 | 8.6648 | 0.14858 | 8.5259 | 8.6551 | 0.12917 | 11.4271 | 0.92477 | 0.82869 | 3.7354 | 0 | 13 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3467.3223 | No Control Sample | 0 | off | off | 3467.3223 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9519 | 5.5051 | 2.466 | 77.965 | -5.2429 | -0.40877 | 0 | 0 | 0 | 0 | 0 | 3.7573 | FALSE | 0.025 | 0.044982
1889 | 1890 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FOSA | 0.122500067228525 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.88 | 8.86 | Baseline | Baseline | 2148.521 | 77082.8566 | 0.027873 | 414.4382 | 16126.0885 | 0.0257 | 5.1842 | 4.78 | 414.4382 | 16133.3849 | 0.98448 | 0.99569 | 9.0348 | 9.0293 | 1.7438 | 0.15485 | 8.7788 | 8.8516 | 9.2448 | 9.173 | 0.46607 | 0.32145 | 0.081702 | 0.074002 | 71.9599 | 135.4519 | 0 | 0.0008441 | FALSE | 1.0006 | TRUE | [M-H]- | 122.5001 | 8.9403 | 9.1073 | 0.16703 | 8.9546 | 9.1013 | 0.14673 | 0 | 0.92696 | 0.93641 | 6.0668 | 0 | 24 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2148.521 | No Control Sample | 0 | off | off | 2148.521 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9487 | 5.0951 | 2.5371 | 77.9648 | -7.0549 | -0.55004 | 0 | 0 | 0 | 0 | 0 | 4.4984 | FALSE | 0.025 | 0.027873
1908 | 1909 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 4:2 FTS | 0.092731925622618 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 5.94 | 5.92 | Baseline | Baseline | 892.3656 | 18952.4037 | 0.047085 | 184.371 | 3128.8458 | 0.058926 | 4.8401 | 6.0573 | 192.3335 | 3138.5067 | 0.85745 | 0.98634 | 5.9635 | 5.9561 | 0.39503 | 0.023465 | 5.8856 | 5.7955 | 6.0609 | 6.1198 | 0.1753 | 0.32435 | 0.092609 | 0.094236 | 73.4157 | 173.5977 | 0.021447 | 0.0022945 | FALSE | 1.0012 | TRUE | [M-H]- | 92.7319 | 5.889 | 6.0445 | 0.15554 | 5.8933 | 6.038 | 0.14467 | -22.5563 | 0.86945 | 0.70008 | 4.966 | 0 | 21 | 13 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Fail | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 892.3656 | No Control Sample | 0 | off | off | 892.3656 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9764 | 6.5877 | 2.154 | 306.9671 | -1.5562 | -0.4777 | 0 | 0 | 0 | 0 | 0 | 58.2284 | FALSE | 0.025 | 0.047085
1910 | 1911 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 6:2 FTS | 0.105429175367563 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7.92 | 7.92 | Manual | Baseline | 1078.1426 | 16854.7723 | 0.063967 | 264.4898 | 4164.6309 | 0.063509 | 4.0763 | 4.0471 | 265.2397 | 4165.2311 | 0.97178 | 0.99829 | 7.9452 | 7.9439 | 0.3179 | 0.025178 | 7.8687 | 7.8141 | 8.0918 | 8.0842 | 0.22309 | 0.27017 | 0.061849 | 0.062088 | 67.6792 | 122.2875 | 0.0028869 | 0.00030261 | TRUE | 1.0002 | TRUE | [M-H]- | 105.4292 | 7.8878 | 8.0621 | 0.17428 | 7.8939 | 8.0069 | 0.11302 | 3.4226 | 1.6986 | 1.4986 | 5.0092 | 0 | 18 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1078.1426 | No Control Sample | 0 | off | off | 1078.1426 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9713 | 7.9976 | 3.4147 | 406.9619 | 0.95145 | 0.3872 | 0 | 0 | 0 | 0 | 0 | 4457.172 | FALSE | 0.025 | 0.063967
1912 | 1913 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 8:2 FTS | 0.136195720651764 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.88 | 8.89 | Baseline | Baseline | 1374.9183 | 20574.9998 | 0.066825 | 293.7953 | 4206.4099 | 0.069845 | 4.6799 | 4.8913 | 293.7953 | 4206.9753 | 1 | 0.99836 | 8.8974 | 8.9 | 0.19614 | 0.017417 | 8.6838 | 8.7494 | 9.0902 | 9.0607 | 0.40633 | 0.31122 | 0.073819 | 0.076539 | 517.2122 | 175.8998 | 0 | 0.00027125 | FALSE | 0.99971 | TRUE | [M-H]- | 136.1957 | 8.8243 | 8.9695 | 0.14519 | 8.8304 | 8.9632 | 0.13277 | 0 | 0.86994 | 0.7161 | 3.6935 | 0 | 21 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1374.9183 | No Control Sample | 0 | off | off | 1374.9183 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9649 | 6.4467 | 3.3972 | 506.9575 | 3.4703 | 1.7593 | 0 | 0 | 0 | 0 | 0 | 3868.5009 | FALSE | 0.025 | 0.066825
1914 | 1915 | PFAS CS3 0.1ng/ml | 14 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 15) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 10:2 FTS | 0.0585445626769408 | Quantifiers | 10/26/2023 22:43:27 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.5 | 8.89 | Manual | Baseline | 615.2611 | 20574.9998 | 0.029903 | 157.7066 | 4206.4099 | 0.037492 | 3.9013 | 4.8913 | 158.8567 | 4206.9753 | 0.97922 | 0.99836 | 9.5134 | 8.9 | 0.14139 | 0.013432 | 9.4422 | 8.7494 | 9.5977 | 9.0607 | 0.15559 | 0.31122 | 0.063968 | 0.076539 | 2502.2739 | 175.8998 | 0.0050618 | 0.00027125 | TRUE | 1.0689 | TRUE | [M-H]- | 58.5446 | 9.4526 | 9.5856 | 0.13301 | 9.4616 | 9.5793 | 0.11774 | 5.1308 | 1.1442 | 1.3946 | 5.2011 | 0 | 13 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Fail | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 615.2611 | No Control Sample | 0 | off | off | 615.2611 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9572 | 3.2759 | 2.0538 | 606.9511 | 5.263 | 3.1944 | 0 | 0 | 0 | 0 | 0 | 9.8007 | FALSE | 0.025 | 0.029903
1975 | 1976 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFBA | 0.116791604568416 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy, Integration Quality | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 1.49 | 1.48 | Manual | Manual | 2287.406 | 109186.7898 | 0.020949 | 317.8998 | 13903.4375 | 0.022865 | 7.1954 | 7.8532 | 330.4241 | 13974.4175 | 0.6751 | 0.98462 | 1.4751 | 1.4802 | 0.9978 | 0.014867 | 1.3673 | 1.3515 | 1.635 | 1.7201 | 0.26768 | 0.36865 | 0.12003 | 0.12221 | 35.4437 | 151.0109 | 0.036831 | 0.0029696 | TRUE | 0.99658 | TRUE | [M-H]- | 116.7916 | 1.3758 | 1.6284 | 0.25252 | 1.3913 | 1.6238 | 0.23251 | -43.7405 | 1.1653 | 1.5024 | 1.4638 | 0 | 55 | 26 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Fail | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2287.406 | No Control Sample | 0 | off | off | 2287.406 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9808 | 7.3023 | 1.5552 | 168.99 | 6.1018 | 1.0311 | 0 | 0 | 0 | 0 | 0 | 50.1989 | FALSE | 0.025 | 0.020949
1977 | 1978 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFPeA | 0.102285333508335 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Integration Quality | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 3.18 | 3.21 | Manual | Baseline | 1792.663 | 78367.7279 | 0.022875 | 239.1046 | 8168.5126 | 0.029272 | 7.4974 | 9.5939 | 244.0093 | 8301.9667 | 0.77238 | 0.94934 | 3.2522 | 3.2622 | 2.2693 | 0.072163 | 3.1341 | 3.0803 | 3.4354 | 3.5126 | 0.30132 | 0.43232 | 0.097289 | 0.1471 | 48.5904 | 609.9186 | 0 | 0.0070836 | TRUE | 0.99693 | TRUE | [M-H]- | 102.2853 | 3.1472 | 3.4266 | 0.27943 | 3.1516 | 3.4222 | 0.27067 | 0 | 1.3714 | 1.7772 | 1.9414 | 0 | 52 | 19 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1792.663 | No Control Sample | 0 | off | off | 1792.663 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9778 | 6.957 | 1.8295 | 218.985 | -3.6618 | -0.80189 | 0 | 0 | 0 | 0 | 0 | 5.484 | FALSE | 0.025 | 0.022875
1979 | 1980 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFHxA | 0.107340639840685 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 2.0111 | 6.13 | 6.12 | Baseline | Baseline | 4099.4829 | 164103.1291 | 0.024981 | 814.059 | 29281.8151 | 0.027801 | 5.0359 | 5.6043 | 817.9018 | 29312.9662 | 0.9398 | 0.99299 | 6.1823 | 6.1871 | 0.85276 | 0.052274 | 6.066 | 6.0225 | 6.3053 | 6.3549 | 0.23932 | 0.33242 | 0.071993 | 0.087057 | 141.56 | 254.5099 | 0.0088255 | 0.0021205 | FALSE | 0.99921 | TRUE | [M-H]- | 107.3406 | 6.1041 | 6.2816 | 0.17742 | 6.1134 | 6.2687 | 0.1553 | 30.0207 | 1.0589 | 1.0853 | 1.5347 | 0 | 26 | 10 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Fail | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4099.4829 | No Control Sample | 203843.5852 | off | off | 4099.4829 | 203843.5852 | off | 10 | off | 0.1 | off | 0.0020111 | off | 312.9751 | 7.2669 | 2.2743 | 268.9838 | 5.3254 | 1.4324 | 0 | 0 | 0 | 0 | 0 | 1389.9954 | FALSE | 0.025 | 0.024981
1981 | 1982 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFHpA | 0.123233011349462 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 2.3641 | 7.26 | 7.27 | Baseline | Baseline | 5397.9378 | 175054.0958 | 0.030836 | 1258.612 | 40236.7964 | 0.03128 | 4.2888 | 4.3506 | 1267.9355 | 40255.4593 | 0.93173 | 0.9962 | 7.3051 | 7.3063 | 0.62114 | 0.045095 | 7.2066 | 7.1524 | 7.3897 | 7.4605 | 0.18318 | 0.3081 | 0.069864 | 0.067781 | 155.1511 | 232.7967 | 0.013689 | 0.0009184 | FALSE | 0.99984 | TRUE | [M-H]- | 123.233 | 7.2383 | 7.3721 | 0.13388 | 7.2457 | 7.3657 | 0.12003 | 94.0542 | 1.0096 | 1.0389 | 1.5201 | 0 | 22 | 10 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5397.9378 | No Control Sample | 228328.7703 | off | off | 5397.9378 | 228328.7703 | off | 10 | off | 0.1 | off | 0.0023641 | off | 362.9721 | 6.7863 | 2.4632 | 318.9816 | 7.8299 | 2.4976 | 0 | 0 | 0 | 0 | 0 | 5.8149 | FALSE | 0.025 | 0.030836
1983 | 1984 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFOA | 0.108556369931127 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 2.5837 | 7.96 | 7.94 | Baseline | Baseline | 4947.1361 | 187909.9715 | 0.026327 | 1050.034 | 43904.491 | 0.023916 | 4.7114 | 4.28 | 1065.5049 | 43931.5773 | 0.82722 | 0.99469 | 7.9851 | 7.9843 | 0.31535 | 0.025102 | 7.889 | 7.8274 | 8.0961 | 8.1353 | 0.20715 | 0.30783 | 0.074952 | 0.066535 | 59.5514 | 109.5356 | 0.0033229 | 0.0011873 | FALSE | 1.0001 | TRUE | [M-H]- | 108.5564 | 7.9146 | 8.0637 | 0.14912 | 7.9229 | 8.0562 | 0.13328 | 16.8435 | 1.0553 | 1.1368 | 1.7483 | 0 | 16 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4947.1361 | No Control Sample | 191475.4056 | off | off | 4947.1361 | 191475.4056 | off | 10 | off | 0.1 | off | 0.0025837 | off | 412.969 | 6.1363 | 2.5341 | 368.9792 | 8.2831 | 3.0563 | 0 | 0 | 0 | 0 | 0 | 10.6792 | FALSE | 0.025 | 0.026327
1985 | 1986 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFNA | 0.116938446658016 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 2.6866 | 8.46 | 8.47 | Baseline | Baseline | 4275.5589 | 157388.2311 | 0.027166 | 881.0041 | 33041.7212 | 0.026663 | 4.8531 | 4.7633 | 883.551 | 33041.7212 | 0.96521 | 0.99925 | 8.4916 | 8.4924 | 0.37405 | 0.031644 | 8.3563 | 8.2484 | 8.5989 | 8.7889 | 0.24262 | 0.54053 | 0.076249 | 0.074585 | 123.0405 | 179.8317 | 0.0026756 | 9.0726e-10 | FALSE | 0.99991 | TRUE | [M-H]- | 116.9384 | 8.4052 | 8.5659 | 0.16062 | 8.4162 | 8.5575 | 0.1413 | 9.7157 | 0.88304 | 0.76554 | 0.96511 | 0 | 14 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Fail | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4275.5589 | No Control Sample | 159143.3885 | off | off | 4275.5589 | 159143.3885 | off | 10 | off | 0.1 | off | 0.0026866 | off | 462.9657 | 5.3309 | 2.468 | 418.976 | 6.6461 | 2.7846 | 0 | 0 | 0 | 0 | 0 | 1.9471 | FALSE | 0.025 | 0.027166
1987 | 1988 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFDA | 0.0993295832850331 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.87 | 8.86 | Baseline | Baseline | 3438.0196 | 139427.9998 | 0.024658 | 668.0064 | 28675.9281 | 0.023295 | 5.1467 | 4.8622 | 668.0064 | 28675.9281 | 0.98033 | 0.99975 | 8.9058 | 8.8995 | 0.4037 | 0.035808 | 8.7027 | 8.6568 | 9.1102 | 9.1776 | 0.40748 | 0.52073 | 0.081299 | 0.076138 | 101.6363 | 147.0691 | 8.2252e-15 | 1.0508e-09 | FALSE | 1.0007 | TRUE | [M-H]- | 99.3296 | 8.8245 | 8.9914 | 0.16697 | 8.8315 | 8.9829 | 0.15134 | 1.3484e-11 | 1.0164 | 1.0159 | 1.6192 | 0 | 21 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3438.0196 | No Control Sample | 0 | off | off | 3438.0196 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9617 | 3.3064 | 1.6961 | 468.9723 | 4.833 | 2.2665 | 0 | 0 | 0 | 0 | 0 | 868.0045 | FALSE | 0.025 | 0.024658
1989 | 1990 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFUdA | 0.108794480584407 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.21 | 9.2 | Baseline | Baseline | 2883.658 | 119520.0272 | 0.024127 | 645.6487 | 26787.9987 | 0.024102 | 4.4663 | 4.4617 | 645.6501 | 26790.3979 | 0.97468 | 0.9991 | 9.2411 | 9.2377 | 0.33794 | 0.031124 | 9.0729 | 9.0811 | 9.4296 | 9.4681 | 0.35668 | 0.387 | 0.069268 | 0.070217 | 97.6085 | 200.3744 | 4.1456e-06 | 0.00022072 | FALSE | 1.0004 | TRUE | [M-H]- | 108.7945 | 9.172 | 9.315 | 0.14295 | 9.1791 | 9.3055 | 0.12643 | -0.0075042 | 0.92497 | 0.79985 | 1.4394 | 0 | 20 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2883.658 | No Control Sample | 0 | off | off | 2883.658 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9588 | 3.507 | 1.9743 | 518.9686 | 2.5245 | 1.3102 | 0 | 0 | 0 | 0 | 0 | 11.3386 | FALSE | 0.025 | 0.024127
1991 | 1992 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFDoA | 0.0981149216186649 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.47 | 9.47 | Baseline | Baseline | 2472.3717 | 97375.8485 | 0.02539 | 639.5313 | 24626.7773 | 0.025969 | 3.8659 | 3.9541 | 639.5735 | 24629.6718 | 0.96164 | 0.99836 | 9.5031 | 9.5039 | 0.34981 | 0.033127 | 9.3802 | 9.3694 | 9.6408 | 9.672 | 0.26058 | 0.3026 | 0.059961 | 0.061718 | 112.2458 | 340.9926 | 0.0001385 | 0.00024388 | FALSE | 0.99991 | TRUE | [M-H]- | 98.1149 | 9.4435 | 9.5736 | 0.13011 | 9.4502 | 9.5636 | 0.1134 | 0.33991 | 1.2363 | 1.4685 | 1.6258 | 0 | 18 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2472.3717 | No Control Sample | 0 | off | off | 2472.3717 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9562 | 4.1617 | 2.5509 | 568.9658 | 4.2206 | 2.4014 | 0 | 0 | 0 | 0 | 0 | 2663.247 | FALSE | 0.025 | 0.02539
1993 | 1994 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFTrDA | 0.0999298272399591 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.68 | 9.47 | Baseline | Baseline | 1835.5199 | 97375.8485 | 0.01885 | 517.803 | 24626.7773 | 0.021026 | 3.5448 | 3.9541 | 517.803 | 24629.6718 | 0.99556 | 0.99836 | 9.7196 | 9.5039 | 0.40957 | 0.039646 | 9.5404 | 9.3694 | 9.8629 | 9.672 | 0.32248 | 0.3026 | 0.056034 | 0.061718 | 386.677 | 340.9926 | 0 | 0.00024388 | FALSE | 1.0227 | TRUE | [M-H]- | 99.9298 | 9.6598 | 9.7742 | 0.11434 | 9.669 | 9.7666 | 0.097551 | 0 | 0.98682 | 1.0023 | 0.5454 | 0 | 28 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1835.5199 | No Control Sample | 0 | off | off | 1835.5199 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9586 | 12.327 | 8.1722 | 618.9589 | -4.4618 | -2.7617 | 0 | 0 | 0 | 0 | 0 | 14.1366 | FALSE | 0.025 | 0.01885
1995 | 1996 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFTeDA | 0.0862727200231972 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Integration Quality, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.87 | 9.86 | Baseline | Baseline | 754.6473 | 24671.5778 | 0.030588 | 194.6116 | 7481.8876 | 0.026011 | 3.8777 | 3.2975 | 195.4832 | 7482.4085 | 0.76078 | 0.99573 | 9.895 | 9.8959 | 0.25342 | 0.025013 | 9.7836 | 9.7858 | 9.9615 | 10.0226 | 0.17787 | 0.23687 | 0.061977 | 0.051644 | 83.0568 | 778.384 | 0.0072046 | 0.00014806 | FALSE | 0.99991 | TRUE | [M-H]- | 86.2727 | 9.8352 | 9.9515 | 0.11629 | 9.8468 | 9.9477 | 0.10088 | 7.8829 | 0.99045 | 1.1414 | 1.6219 | 0 | 19 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 754.6473 | No Control Sample | 0 | off | off | 754.6473 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9498 | 3.6076 | 2.572 | 668.9627 | 7.1168 | 4.7608 | 0 | 0 | 0 | 0 | 0 | 3389.696 | FALSE | 0.025 | 0.030588
2002 | 2003 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFBS | 0.0884419415411323 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 3.5 | 3.5 | Manual | Manual | 1035.319 | 44226.44 | 0.02341 | 113.8057 | 4124.9869 | 0.027589 | 9.0972 | 10.7216 | 114.4231 | 4133.1744 | 0.86746 | 0.99448 | 3.919 | 3.9429 | 11.9701 | 0.41895 | 3.7627 | 3.675 | 4.0898 | 4.2903 | 0.32706 | 0.61531 | 0.15469 | 0.16543 | 45.0512 | 563.65 | 0.0011679 | 0.00080966 | TRUE | 0.99392 | TRUE | [M-H]- | 88.4419 | 3.7689 | 4.0809 | 0.312 | 3.7719 | 4.0782 | 0.30634 | -0.40639 | 1.4188 | 1.8652 | 12.5449 | 0 | 52 | 25 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1035.319 | No Control Sample | 0 | off | off | 1035.319 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9452 | 7.3945 | 2.2105 | 79.9589 | 21.3749 | 1.7091 | 0 | 0 | 0 | 0 | 0 | 7.6663 | FALSE | 0.025 | 0.02341
2004 | 2005 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFPeS | 0.131824134532384 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 6.38 | 3.5 | Manual | Manual | 1553.7855 | 44226.44 | 0.035133 | 315.3572 | 4124.9869 | 0.07645 | 4.9271 | 10.7216 | 315.6123 | 4133.1744 | 0.98944 | 0.99448 | 6.4343 | 3.9429 | 0.85064 | 0.054271 | 6.3258 | 3.675 | 6.563 | 4.2903 | 0.23723 | 0.61531 | 0.072918 | 0.16543 | 106.5513 | 563.65 | 0 | 0.00080966 | TRUE | 1.6319 | TRUE | [M-H]- | 131.8241 | 6.3443 | 6.5231 | 0.17877 | 6.3554 | 6.5139 | 0.15853 | 0 | 1.0635 | 1.172 | 11.3306 | 0 | 27 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1553.7855 | No Control Sample | 0 | off | off | 1553.7855 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9417 | 5.37 | 1.8738 | 79.9568 | -5.4007 | -0.43182 | 0 | 0 | 0 | 0 | 0 | 2.0005 | FALSE | 0.025 | 0.035133
2006 | 2007 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | L-PFHxS | 0.096650612709631 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.0811 | 4 | 0.020275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 1174.486 | 65054.846 | 0.018054 | 338.1139 | 15779.0912 | 0.021428 | 3.4736 | 4.1229 | 339.2895 | 15779.2148 | 0.7839 | 0.99791 | 7.3794 | 7.3797 | 0.53746 | 0.039449 | 7.2891 | 7.2385 | 7.4596 | 7.6579 | 0.17045 | 0.41942 | 0.048753 | 0.063766 | 92.6304 | 200.1814 | 0.0064025 | 1.141e-05 | FALSE | 0.99997 | TRUE | [M-H]- | 119.1746 | 7.3068 | 7.4387 | 0.13191 | 7.3162 | 7.4329 | 0.11667 | -12.7004 | 0.97211 | 0.99641 | 13.6107 | 0 | 21 | 8 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1174.486 | No Control Sample | 0 | off | off | 1174.486 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9384 | 4.4871 | 1.7901 | 79.9557 | -18.625 | -1.4892 | 0 | 0 | 0 | 0 | 0 | 5.3257 | FALSE | 0.020275 | 0.018054
2008 | 2009 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | Br-PFHxS | 0.016968765178456 | Qualifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.0189 | 4 | 0.004725 | 0 | 7.16 | 7.35 | Manual | Baseline | 280.1543 | 65054.846 | 0.0043064 | 63.7263 | 15779.0912 | 0.0040387 | 4.3962 | 4.1229 | 339.2895 | 15779.2148 | 0.18699 | 0.99791 | 7.1976 | 7.3797 | 0.52447 | 0.037552 | 7.113 | 7.2385 | 7.2891 | 7.6579 | 0.1761 | 0.41942 | 0.062236 | 0.063766 | 17.4586 | 200.1814 | 0 | 1.141e-05 | TRUE | 0.97532 | TRUE | [M-H]- | 89.7818 | 7.1182 | 7.2851 | 0.16693 | 7.1237 | 7.2769 | 0.15315 | 0 | 1.0345 | 1.039 | 0.23853 | 0 | 21 | 8 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 280.1543 | No Control Sample | 0 | off | off | 1174.486 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9384 | 4.4871 | 1.7901 | 79.9593 | 25.8338 | 2.0656 | 0 | 0 | 0 | 0 | 0 | 5.3257 | FALSE | 0.004725 | 0.0043064
2010 | 2011 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFHpS | 0.122347408313867 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7.98 | 7.35 | Baseline | Baseline | 1211.3555 | 65054.846 | 0.018621 | 254.289 | 15779.0912 | 0.016116 | 4.7637 | 4.1229 | 255.1797 | 15779.2148 | 0.95734 | 0.99791 | 8.011 | 7.3797 | 0.38877 | 0.031024 | 7.878 | 7.2385 | 8.1155 | 7.6579 | 0.23755 | 0.41942 | 0.074195 | 0.063766 | 80.3054 | 200.1814 | 0.0084596 | 1.141e-05 | FALSE | 1.0856 | TRUE | [M-H]- | 122.3474 | 7.9243 | 8.0797 | 0.15533 | 7.9457 | 8.0732 | 0.12747 | -9.0557 | 0.82574 | 0.75311 | 16.12 | 0 | 18 | 6 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1211.3555 | No Control Sample | 0 | off | off | 1211.3555 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9355 | 4.718 | 2.1181 | 79.9571 | -1.0624 | -0.084948 | 0 | 0 | 0 | 0 | 0 | 5.77 | FALSE | 0.025 | 0.018621
2012 | 2013 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | L-PFOS | 0.0947390696807956 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.0788 | 4 | 0.0197 | 0 | 8.46 | 8.45 | Manual | Baseline | 1007.9187 | 40929.9354 | 0.024625 | 202.034 | 8776.6797 | 0.023019 | 4.9889 | 4.6635 | 205.0251 | 8778.6558 | 0.5739 | 0.99562 | 8.4854 | 8.4905 | 0.30026 | 0.025402 | 8.4155 | 8.331 | 8.5712 | 8.6944 | 0.15578 | 0.36342 | 0.080122 | 0.072464 | 33.8945 | 143.6911 | 0.025498 | 0.00051166 | TRUE | 0.9994 | TRUE | [M-H]- | 120.2272 | 8.3979 | 8.5599 | 0.16195 | 8.3979 | 8.5509 | 0.15302 | -33.0682 | 0.97738 | 0.84691 | 11.5341 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1007.9187 | No Control Sample | 0 | off | off | 1007.9187 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9329 | 5.3189 | 2.6537 | 79.9569 | -5.0852 | -0.4066 | 0 | 0 | 0 | 0 | 0 | 6.841 | FALSE | 0.0197 | 0.024625
2014 | 2015 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | Br-PFOS | 0.0206529819271779 | Qualifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Asymmetry, Total Width | 0 | 0 | 0.0211 | 4 | 0.005275 | 0 | 8.3 | 8.45 | Manual | Baseline | 470.3934 | 40929.9354 | 0.011493 | 83.8691 | 8776.6797 | 0.0095559 | 5.6087 | 4.6635 | 205.0251 | 8778.6558 | 0.27148 | 0.99562 | 8.3843 | 8.4905 | 1.0158 | 0.08431 | 8.2224 | 8.331 | 8.418 | 8.6944 | 0.19557 | 0.36342 | 0.10278 | 0.072464 | 14.0704 | 143.6911 | 0.0053088 | 0.00051166 | TRUE | 0.98749 | TRUE | [M-H]- | 97.8814 | 8.2367 | 8.4381 | 0.20142 | 8.2467 | 8.4381 | 0.19146 | -2.2766 | 0.70631 | 0.4436 | 0.4667 | 0 | 14 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 470.3934 | No Control Sample | 0 | off | off | 1007.9187 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9329 | 5.3189 | 2.6537 | 79.9564 | -11.1949 | -0.89511 | 0 | 0 | 0 | 0 | 0 | 6.841 | FALSE | 0.005275 | 0.011493
2016 | 2017 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFNS | 0.0805594036244996 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.85 | 8.45 | Baseline | Baseline | 730.8318 | 40929.9354 | 0.017856 | 175.1645 | 8776.6797 | 0.019958 | 4.1723 | 4.6635 | 175.1645 | 8778.6558 | 0.94338 | 0.99562 | 8.8886 | 8.4905 | 0.43573 | 0.038562 | 8.6698 | 8.331 | 9.0575 | 8.6944 | 0.38777 | 0.36342 | 0.064427 | 0.072464 | 50.1697 | 143.6911 | 2.0854e-09 | 0.00051166 | FALSE | 1.0469 | TRUE | [M-H]- | 80.5594 | 8.8124 | 8.9654 | 0.15302 | 8.8276 | 8.9524 | 0.12476 | 9.4202e-07 | 0.98388 | 0.99616 | 24.1728 | 0 | 20 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 730.8318 | No Control Sample | 0 | off | off | 730.8318 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9299 | 5.2935 | 2.9057 | 79.9572 | 1.6294 | 0.13029 | 0 | 0 | 0 | 0 | 0 | 2.5187 | FALSE | 0.025 | 0.017856
2018 | 2019 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFDS | 0.111998732170808 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.19 | 8.45 | Baseline | Baseline | 1043.9469 | 40929.9354 | 0.025506 | 234.2052 | 8776.6797 | 0.026685 | 4.4574 | 4.6635 | 239.5445 | 8778.6558 | 0.89665 | 0.99562 | 9.219 | 8.4905 | 0.31545 | 0.02899 | 9.1355 | 8.331 | 9.3318 | 8.6944 | 0.19632 | 0.36342 | 0.070075 | 0.072464 | 45.9401 | 143.6911 | 0.036536 | 0.00051166 | FALSE | 1.0858 | TRUE | [M-H]- | 111.9987 | 9.1466 | 9.2932 | 0.14668 | 9.1574 | 9.2872 | 0.12975 | -43.5871 | 1.0931 | 1.3068 | 15.1843 | 0 | 11 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Fail | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1043.9469 | No Control Sample | 0 | off | off | 1043.9469 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9262 | 3.9217 | 2.3488 | 79.9565 | -4.6306 | -0.37025 | 0 | 0 | 0 | 0 | 0 | 11.626 | FALSE | 0.025 | 0.025506
2022 | 2023 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FBSA | 0.0852615571823694 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 4.6 | 8.86 | Manual | Baseline | 1724.3844 | 76702.1975 | 0.022482 | 196.1912 | 16634.9268 | 0.011794 | 8.7893 | 4.6109 | 198.6734 | 16637.7057 | 0.95424 | 0.9977 | 5.0379 | 9.0278 | 9.5193 | 0.43789 | 4.8479 | 8.7936 | 5.239 | 9.2993 | 0.3911 | 0.50567 | 0.15496 | 0.070994 | 85.5184 | 173.116 | 0.0050887 | 0.00035082 | TRUE | 0.55804 | TRUE | [M-H]- | 85.2616 | 4.8536 | 5.236 | 0.38235 | 4.9029 | 5.2293 | 0.32645 | 2.5527 | 1.0363 | 1.4135 | 4.9172 | 0 | 71 | 29 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1724.3844 | No Control Sample | 0 | off | off | 1724.3844 | 0 | off | 0 | off | 0 | off | 0 | off | 297.961 | 6.8887 | 2.0526 | 77.9662 | 11.6355 | 0.90717 | 0 | 0 | 0 | 0 | 0 | 1.8272 | FALSE | 0.025 | 0.022482
2024 | 2025 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FPeSA | 0.0874257352071054 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7 | 8.86 | Baseline | Baseline | 2745.7844 | 76702.1975 | 0.035798 | 637.8099 | 16634.9268 | 0.038342 | 4.305 | 4.6109 | 653.1534 | 16637.7057 | 0.91896 | 0.9977 | 7.1287 | 9.0278 | 1.8381 | 0.12866 | 7.0276 | 8.7936 | 7.2169 | 9.2993 | 0.1893 | 0.50567 | 0.06481 | 0.070994 | 139.3133 | 173.116 | 0.037678 | 0.00035082 | FALSE | 0.78963 | TRUE | [M-H]- | 87.4257 | 7.0572 | 7.2042 | 0.14705 | 7.0649 | 7.1966 | 0.13169 | 126.9475 | 1.1304 | 1.2997 | 4.2084 | 0 | 21 | 9 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2745.7844 | No Control Sample | 0 | off | off | 2745.7844 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9579 | 6.2264 | 2.1665 | 77.9644 | -11.7241 | -0.91407 | 0 | 0 | 0 | 0 | 0 | 6.5438 | FALSE | 0.025 | 0.035798
2026 | 2027 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FHxSA | 0.125669487502751 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7.86 | 8.86 | Baseline | Baseline | 1969.5925 | 76702.1975 | 0.025678 | 437.0552 | 16634.9268 | 0.026273 | 4.5065 | 4.6109 | 437.7647 | 16637.7057 | 0.96886 | 0.9977 | 8.0108 | 9.0278 | 1.9189 | 0.15082 | 7.8899 | 8.7936 | 8.1296 | 9.2993 | 0.23978 | 0.50567 | 0.069369 | 0.070994 | 98.223 | 173.116 | 0.0030792 | 0.00035082 | FALSE | 0.88735 | TRUE | [M-H]- | 125.6695 | 7.937 | 8.0929 | 0.15593 | 7.9491 | 8.079 | 0.12993 | 5.6125 | 0.9893 | 0.94725 | 7.5972 | 0 | 18 | 6 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1969.5925 | No Control Sample | 0 | off | off | 1969.5925 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9546 | 5.0293 | 2.0014 | 77.9648 | -7.168 | -0.55886 | 0 | 0 | 0 | 0 | 0 | 6.3408 | FALSE | 0.025 | 0.025678
2028 | 2029 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FHpSA | 0.102784082608855 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.58 | 8.86 | Baseline | Baseline | 3336.6447 | 76702.1975 | 0.043501 | 740.7863 | 16634.9268 | 0.044532 | 4.5042 | 4.6109 | 744.3602 | 16637.7057 | 0.97359 | 0.9977 | 8.5871 | 9.0278 | 0.082482 | 0.0070769 | 8.479 | 8.7936 | 8.7045 | 9.2993 | 0.22543 | 0.50567 | 0.070278 | 0.070994 | 111.5843 | 173.116 | 0.005315 | 0.00035082 | FALSE | 0.95118 | TRUE | [M-H]- | 102.7841 | 8.5141 | 8.6648 | 0.1507 | 8.5214 | 8.6557 | 0.13425 | -17.4654 | 1.0873 | 1.1674 | 4.1979 | 0 | 14 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3336.6447 | No Control Sample | 0 | off | off | 3336.6447 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9514 | 4.4108 | 1.9758 | 77.9647 | -9.0884 | -0.70858 | 0 | 0 | 0 | 0 | 0 | 1.4988 | FALSE | 0.025 | 0.043501
2030 | 2031 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FOSA | 0.125015874351583 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.88 | 8.86 | Baseline | Baseline | 2181.8176 | 76702.1975 | 0.028445 | 482.3755 | 16634.9268 | 0.028998 | 4.5231 | 4.6109 | 485.7057 | 16637.7057 | 0.97084 | 0.9977 | 9.0245 | 9.0278 | 1.6276 | 0.14453 | 8.9092 | 8.7936 | 9.1328 | 9.2993 | 0.2236 | 0.50567 | 0.070942 | 0.070994 | 87.8085 | 173.116 | 0.011548 | 0.00035082 | FALSE | 0.99963 | TRUE | [M-H]- | 125.0159 | 8.9483 | 9.1006 | 0.15235 | 8.957 | 9.0911 | 0.13402 | -24.9123 | 1.0886 | 1.1887 | 6.4597 | 0 | 12 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2181.8176 | No Control Sample | 0 | off | off | 2181.8176 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9488 | 5.2332 | 2.6058 | 77.9647 | -8.5448 | -0.6662 | 0 | 0 | 0 | 0 | 0 | 2.5114 | FALSE | 0.025 | 0.028445
2049 | 2050 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 4:2 FTS | 0.108780222785509 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 5.94 | 5.92 | Baseline | Baseline | 1086.6991 | 19674.789 | 0.055233 | 146.9393 | 3181.8709 | 0.04618 | 7.3956 | 6.1834 | 146.9869 | 3205.5809 | 0.98006 | 0.97452 | 6.0001 | 6.0006 | 1.0125 | 0.060145 | 5.8837 | 5.861 | 6.1436 | 6.1368 | 0.25992 | 0.27577 | 0.12072 | 0.097112 | 75.2565 | 156.4247 | 0.00063945 | 0.00098043 | FALSE | 0.99993 | TRUE | [M-H]- | 108.7802 | 5.9049 | 6.11 | 0.20506 | 5.9103 | 6.1037 | 0.19335 | -0.3615 | 0.95523 | 0.89618 | 4.2357 | 0 | 30 | 15 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1086.6991 | No Control Sample | 0 | off | off | 1086.6991 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9768 | 7.5936 | 2.4829 | 306.9664 | -3.8633 | -1.1859 | 0 | 0 | 0 | 0 | 0 | 66.4841 | FALSE | 0.025 | 0.055233
2051 | 2052 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 6:2 FTS | 0.0828734393393109 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7.92 | 7.92 | Baseline | Baseline | 856.1006 | 17026.1685 | 0.050281 | 200.2539 | 4118.2237 | 0.048626 | 4.2751 | 4.1343 | 200.2567 | 4118.2237 | 0.95887 | 0.99888 | 7.9458 | 7.9468 | 0.32542 | 0.025773 | 7.8126 | 7.763 | 8.0696 | 8.1709 | 0.25695 | 0.40795 | 0.065724 | 0.062475 | 65.2003 | 112.8246 | 2.7296e-05 | 2.9596e-11 | FALSE | 0.99987 | TRUE | [M-H]- | 82.8734 | 7.8736 | 8.021 | 0.14735 | 7.8837 | 8.0141 | 0.13043 | 0.021273 | 1.0716 | 1.2219 | 6.2955 | 0 | 20 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 856.1006 | No Control Sample | 0 | off | off | 856.1006 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9701 | 5.1402 | 2.1947 | 406.9635 | 4.9337 | 2.0078 | 0 | 0 | 0 | 0 | 0 | 39.7519 | FALSE | 0.025 | 0.050281
2053 | 2054 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 8:2 FTS | 0.0860863667517304 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.88 | 8.89 | Baseline | Baseline | 901.5607 | 21344.5502 | 0.042238 | 190.7325 | 4429.7016 | 0.043058 | 4.7268 | 4.8185 | 190.7325 | 4429.7016 | 1 | 0.99802 | 8.8942 | 8.8992 | 0.15945 | 0.014159 | 8.6533 | 8.6607 | 9.0972 | 9.1654 | 0.44385 | 0.50475 | 0.073438 | 0.074187 | 159.0269 | 134.8372 | 0 | 4.1243e-14 | FALSE | 0.99943 | TRUE | [M-H]- | 86.0864 | 8.8103 | 8.9743 | 0.16398 | 8.8307 | 8.9659 | 0.13517 | 0 | 1.0057 | 1.21 | 6.702 | 0 | 23 | 5 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 901.5607 | No Control Sample | 0 | off | off | 901.5607 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9649 | 6.3535 | 3.3481 | 506.9603 | 9.0956 | 4.6111 | 0 | 0 | 0 | 0 | 0 | 3274.5822 | FALSE | 0.025 | 0.042238
2055 | 2056 | PFAS CS3 0.1ng/ml | 15 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 16) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 10:2 FTS | 0.0397000071270998 | Quantifiers | 10/26/2023 23:00:04 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.5 | 8.89 | Manual | Baseline | 432.8233 | 21344.5502 | 0.020278 | 114.8547 | 4429.7016 | 0.025928 | 3.7684 | 4.8185 | 115.749 | 4429.7016 | 0.97934 | 0.99802 | 9.5094 | 8.8992 | 0.099048 | 0.0094096 | 9.4451 | 8.6607 | 9.5934 | 9.1654 | 0.14827 | 0.50475 | 0.056639 | 0.074187 | 1412.4507 | 134.8372 | 0.0050643 | 4.1243e-14 | TRUE | 1.0686 | TRUE | [M-H]- | 39.7 | 9.4542 | 9.5814 | 0.12725 | 9.4599 | 9.5752 | 0.1152 | -3.9228 | 1.0688 | 1.1433 | 6.4452 | 0 | 12 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Fail | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 432.8233 | No Control Sample | 0 | off | off | 432.8233 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9566 | 2.4059 | 1.5084 | 606.9418 | -10.0656 | -6.1093 | 0 | 0 | 0 | 0 | 0 | 15.4063 | FALSE | 0.025 | 0.020278
2116 | 2117 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFBA | 0.113457504559125 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Integration Quality | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 1.49 | 1.48 | Manual | Baseline | 2313.9078 | 113697.5993 | 0.020351 | 325.0031 | 14438.937 | 0.022509 | 7.1196 | 7.8744 | 334.3343 | 14466.2943 | 0.70813 | 0.98796 | 1.4719 | 1.4711 | 1.2146 | 0.018098 | 1.3421 | 1.3127 | 1.6567 | 1.7175 | 0.31462 | 0.4048 | 0.097227 | 0.12536 | 35.1715 | 181.8511 | 0.015609 | 0.0055876 | TRUE | 1.0006 | TRUE | [M-H]- | 113.4575 | 1.3436 | 1.6442 | 0.30069 | 1.3509 | 1.6351 | 0.28424 | -16.1238 | 1.2055 | 1.4215 | 1.5065 | 0 | 64 | 21 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Fail | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2313.9078 | No Control Sample | 0 | off | off | 2313.9078 | 0 | off | 0 | off | 0 | off | 0 | off | 212.98 | 3.5315 | 0.75214 | 168.9878 | -6.8811 | -1.1628 | 0 | 0 | 0 | 0 | 0 | 50.7494 | FALSE | 0.025 | 0.020351
2118 | 2119 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFPeA | 0.113674934141999 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 3.18 | 3.21 | Manual | Manual | 2142.8485 | 84290.5181 | 0.025422 | 233.7425 | 8400.8616 | 0.027824 | 9.1676 | 10.0336 | 237.9259 | 8428.9512 | 0.80879 | 0.98669 | 3.2342 | 3.2299 | 1.7051 | 0.054221 | 3.0576 | 3.0075 | 3.4158 | 3.532 | 0.35821 | 0.5245 | 0.16897 | 0.14961 | 45.1521 | 587.447 | 0.0051151 | 0.0016215 | TRUE | 1.0013 | TRUE | [M-H]- | 113.6749 | 3.0615 | 3.4081 | 0.3466 | 3.0925 | 3.4012 | 0.30871 | 3.3377 | 0.99122 | 1.1459 | 1.6236 | 0 | 62 | 31 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2142.8485 | No Control Sample | 0 | off | off | 2142.8485 | 0 | off | 0 | off | 0 | off | 0 | off | 262.978 | 7.7154 | 2.029 | 218.9868 | 4.4215 | 0.96825 | 0 | 0 | 0 | 0 | 0 | 14.7993 | FALSE | 0.025 | 0.025422
2120 | 2121 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFHxA | 0.102754857970672 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 1.9147 | 6.13 | 6.12 | Baseline | Baseline | 3902.9026 | 163206.4451 | 0.023914 | 623.7884 | 29076.513 | 0.021453 | 6.2568 | 5.613 | 630.7546 | 29097.7564 | 0.9138 | 0.99553 | 6.1853 | 6.1879 | 0.90189 | 0.055286 | 6.0613 | 5.9899 | 6.3194 | 6.3971 | 0.25805 | 0.40727 | 0.10199 | 0.085243 | 98.9298 | 247.3851 | 0.015391 | 0.0014279 | FALSE | 0.99958 | TRUE | [M-H]- | 102.7549 | 6.0887 | 6.2792 | 0.19056 | 6.1032 | 6.2706 | 0.16742 | 37.2056 | 1.0069 | 1.0903 | 1.5711 | 0 | 28 | 13 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3902.9026 | No Control Sample | 203843.5852 | off | off | 3902.9026 | 203843.5852 | off | 10 | off | 0.1 | off | 0.0019147 | off | 312.9744 | 5.2189 | 1.6334 | 268.9825 | 0.22211 | 0.059743 | 0 | 0 | 0 | 0 | 0 | 1472.0968 | FALSE | 0.025 | 0.023914
2122 | 2123 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFHpA | 0.115369596948412 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 2.3829 | 7.26 | 7.27 | Baseline | Baseline | 5440.8983 | 188473.6701 | 0.028868 | 1336.174 | 44786.2345 | 0.029834 | 4.072 | 4.2083 | 1346.9795 | 44792.8061 | 0.94425 | 0.99714 | 7.3187 | 7.3188 | 0.80866 | 0.058709 | 7.2306 | 7.1668 | 7.3954 | 7.4917 | 0.16475 | 0.32485 | 0.06339 | 0.065336 | 196.7581 | 266.3912 | 0.010975 | 0.00030488 | FALSE | 0.99999 | TRUE | [M-H]- | 115.3696 | 7.2507 | 7.3851 | 0.13443 | 7.257 | 7.3788 | 0.12175 | 89.0128 | 0.99307 | 0.98453 | 1.5256 | 0 | 20 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5440.8983 | No Control Sample | 228328.7703 | off | off | 5440.8983 | 228328.7703 | off | 10 | off | 0.1 | off | 0.0023829 | off | 362.9712 | 4.3413 | 1.5758 | 318.9814 | 7.1501 | 2.2807 | 0 | 0 | 0 | 0 | 0 | 6.8145 | FALSE | 0.025 | 0.028868
2124 | 2125 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFOA | 0.10901394775755 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 2.6076 | 7.96 | 7.94 | Baseline | Baseline | 4992.8457 | 188850.1594 | 0.026438 | 1112.504 | 43674.8881 | 0.025472 | 4.4879 | 4.324 | 1121.8794 | 43680.8542 | 0.89149 | 0.99772 | 7.994 | 7.9942 | 0.4271 | 0.033997 | 7.887 | 7.8126 | 8.1103 | 8.2131 | 0.22332 | 0.4004 | 0.070636 | 0.067317 | 60.2472 | 132.8492 | 0.0019014 | 0.00030409 | FALSE | 0.99997 | TRUE | [M-H]- | 109.0139 | 7.9262 | 8.0746 | 0.14843 | 7.936 | 8.0623 | 0.12625 | -9.472 | 1.0729 | 1.128 | 1.6152 | 0 | 17 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4992.8457 | No Control Sample | 191475.4056 | off | off | 4992.8457 | 191475.4056 | off | 10 | off | 0.1 | off | 0.0026076 | off | 412.9683 | 4.6008 | 1.9 | 368.9779 | 4.9992 | 1.8446 | 0 | 0 | 0 | 0 | 0 | 2.2967 | FALSE | 0.025 | 0.026438
2126 | 2127 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFNA | 0.116558355409102 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 2.6941 | 8.46 | 8.47 | Baseline | Baseline | 4287.4747 | 158341.5313 | 0.027077 | 874.1957 | 34126.9143 | 0.025616 | 4.9045 | 4.6398 | 874.3505 | 34129.7537 | 0.97515 | 0.99853 | 8.5022 | 8.499 | 0.49831 | 0.042157 | 8.3733 | 8.3451 | 8.6821 | 8.7874 | 0.30883 | 0.44232 | 0.076764 | 0.072332 | 114.1142 | 180.6743 | 0.00028176 | 7.1873e-05 | FALSE | 1.0004 | TRUE | [M-H]- | 116.5584 | 8.418 | 8.5773 | 0.15927 | 8.4311 | 8.569 | 0.13787 | -0.79757 | 1.0387 | 1.1681 | 1.1785 | 0 | 18 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4287.4747 | No Control Sample | 159143.3885 | off | off | 4287.4747 | 159143.3885 | off | 10 | off | 0.1 | off | 0.0026941 | off | 462.9649 | 3.6217 | 1.6767 | 418.976 | 6.7021 | 2.808 | 0 | 0 | 0 | 0 | 0 | 9.8103 | FALSE | 0.025 | 0.027077
2128 | 2129 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFDA | 0.0931364069003621 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.87 | 8.86 | Baseline | Baseline | 3313.9497 | 143333.1675 | 0.023121 | 681.694 | 29680.6363 | 0.022968 | 4.8613 | 4.8292 | 682.8444 | 29705.522 | 0.9696 | 0.99583 | 8.9075 | 8.9072 | 0.42316 | 0.037534 | 8.7586 | 8.7482 | 9.0704 | 9.0598 | 0.3118 | 0.3116 | 0.073815 | 0.075255 | 78.7594 | 151.7148 | 0.0031605 | 0.0014866 | FALSE | 1 | TRUE | [M-H]- | 93.1364 | 8.8258 | 8.9975 | 0.17163 | 8.8373 | 8.9827 | 0.1454 | -6.91 | 1.0826 | 1.1434 | 1.8411 | 0 | 16 | 5 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3313.9497 | No Control Sample | 0 | off | off | 3313.9497 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9636 | 6.9038 | 3.5414 | 468.9714 | 2.9098 | 1.3646 | 0 | 0 | 0 | 0 | 0 | 743.9469 | FALSE | 0.025 | 0.023121
2130 | 2131 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFUdA | 0.108288770014107 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.21 | 9.2 | Baseline | Baseline | 2678.8123 | 111548.2222 | 0.024015 | 553.2126 | 25166.4376 | 0.021982 | 4.8423 | 4.4324 | 554.0162 | 25166.876 | 0.97316 | 0.99971 | 9.2379 | 9.2373 | 0.30309 | 0.027914 | 9.0904 | 9.0795 | 9.343 | 9.5146 | 0.2526 | 0.43507 | 0.07588 | 0.06967 | 98.9695 | 213.986 | 0.0023591 | 4.7763e-05 | FALSE | 1.0001 | TRUE | [M-H]- | 108.2888 | 9.1541 | 9.313 | 0.15891 | 9.1703 | 9.3045 | 0.13424 | 5.1667 | 1.0797 | 1.3387 | 1.7728 | 0 | 14 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2678.8123 | No Control Sample | 0 | off | off | 2678.8123 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9607 | 6.8423 | 3.8519 | 518.9672 | -0.14925 | -0.077454 | 0 | 0 | 0 | 0 | 0 | 15.6711 | FALSE | 0.025 | 0.024015
2132 | 2133 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFDoA | 0.0900916185144154 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.47 | 9.47 | Baseline | Baseline | 2090.8651 | 89683.8356 | 0.023314 | 480.9927 | 23301.4606 | 0.020642 | 4.347 | 3.8489 | 485.1389 | 23307.03 | 0.93317 | 0.99716 | 9.4986 | 9.5015 | 0.30222 | 0.02862 | 9.4121 | 9.3676 | 9.6115 | 9.6849 | 0.19947 | 0.31727 | 0.069055 | 0.059455 | 99.0899 | 352.1951 | 0.014332 | 0.00056126 | FALSE | 0.9997 | TRUE | [M-H]- | 90.0916 | 9.4315 | 9.57 | 0.13848 | 9.438 | 9.5619 | 0.12391 | -34.5609 | 1.1017 | 1.199 | 1.7721 | 0 | 14 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2090.8651 | No Control Sample | 0 | off | off | 2090.8651 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9572 | 5.7563 | 3.5283 | 568.9667 | 5.7485 | 3.2707 | 0 | 0 | 0 | 0 | 0 | 4007.318 | FALSE | 0.025 | 0.023314
2134 | 2135 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFTrDA | 0.0885949417165727 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.68 | 9.47 | Baseline | Baseline | 1498.7728 | 89683.8356 | 0.016712 | 394.4334 | 23301.4606 | 0.016927 | 3.7998 | 3.8489 | 396.3611 | 23307.03 | 0.95924 | 0.99716 | 9.7179 | 9.5015 | 0.39107 | 0.037856 | 9.6271 | 9.3676 | 9.8031 | 9.6849 | 0.176 | 0.31727 | 0.061418 | 0.059455 | 139.8259 | 352.1951 | 0.0091763 | 0.00056126 | FALSE | 1.0228 | TRUE | [M-H]- | 88.5949 | 9.659 | 9.7835 | 0.12443 | 9.6646 | 9.7708 | 0.10619 | 20.5649 | 1.0589 | 0.9953 | 0.83126 | 0 | 16 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1498.7728 | No Control Sample | 0 | off | off | 1498.7728 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9546 | 6.3056 | 4.1803 | 618.9637 | 3.3001 | 2.0426 | 0 | 0 | 0 | 0 | 0 | 1.7134 | FALSE | 0.025 | 0.016712
2136 | 2137 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFTeDA | 0.101456741134951 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Integration Quality, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.87 | 9.86 | Baseline | Baseline | 821.5501 | 22839.1301 | 0.035971 | 224.762 | 6544.3056 | 0.034345 | 3.6552 | 3.4899 | 224.7634 | 6556.6825 | 0.73735 | 0.98818 | 9.8924 | 9.892 | 0.22698 | 0.022403 | 9.7939 | 9.796 | 9.9947 | 9.9796 | 0.20082 | 0.18363 | 0.05876 | 0.05545 | 64.0034 | 613.4433 | 1.1776e-05 | 0.0037788 | FALSE | 1 | TRUE | [M-H]- | 101.4567 | 9.8288 | 9.9625 | 0.13362 | 9.8349 | 9.9557 | 0.12075 | 0.01318 | 1.0209 | 1.0347 | 0.90345 | 0 | 22 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 821.5501 | No Control Sample | 0 | off | off | 821.5501 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9504 | 4.439 | 3.1647 | 668.9559 | -3.0436 | -2.036 | 0 | 0 | 0 | 0 | 0 | 2816.6787 | FALSE | 0.025 | 0.035971
2143 | 2144 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFBS | 0.105376610479537 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 3.5 | 3.5 | Manual | Manual | 1312.6802 | 47063.1251 | 0.027892 | 166.1723 | 4161.4777 | 0.039931 | 7.8995 | 11.3092 | 167.0681 | 4169.9802 | 0.95343 | 0.99122 | 3.9067 | 3.8848 | 11.619 | 0.40667 | 3.6769 | 3.6334 | 4.1143 | 4.2105 | 0.43746 | 0.57711 | 0.16055 | 0.16118 | 68.4417 | 647.3723 | 0.0021834 | 0.00083501 | TRUE | 1.0056 | TRUE | [M-H]- | 105.3766 | 3.7018 | 4.1063 | 0.40445 | 3.7437 | 4.0564 | 0.31273 | -0.82937 | 0.97498 | 0.88911 | 9.9484 | 0 | 69 | 27 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1312.6802 | No Control Sample | 0 | off | off | 1312.6802 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9446 | 5.5076 | 1.6465 | 79.9573 | 1.0175 | 0.081359 | 0 | 0 | 0 | 0 | 0 | 5.2009 | FALSE | 0.025 | 0.027892
2145 | 2146 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFPeS | 0.108300728288991 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 6.38 | 3.5 | Manual | Manual | 1358.3956 | 47063.1251 | 0.028863 | 257.7033 | 4161.4777 | 0.061926 | 5.2712 | 11.3092 | 258.0592 | 4169.9802 | 0.96649 | 0.99122 | 6.4389 | 3.8848 | 0.92366 | 0.05893 | 6.3066 | 3.6334 | 6.5411 | 4.2105 | 0.23449 | 0.57711 | 0.077585 | 0.16118 | 85.0577 | 647.3723 | 0.0015678 | 0.00083501 | TRUE | 1.6575 | TRUE | [M-H]- | 108.3007 | 6.3226 | 6.5247 | 0.20218 | 6.362 | 6.5185 | 0.1565 | -1.723 | 0.78236 | 0.74426 | 13.4609 | 0 | 27 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1358.3956 | No Control Sample | 0 | off | off | 1358.3956 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9415 | 4.9065 | 1.7121 | 79.9579 | 8.6999 | 0.69562 | 0 | 0 | 0 | 0 | 0 | 1.2598 | FALSE | 0.025 | 0.028863
2147 | 2148 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | L-PFHxS | 0.0920597021282502 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Total Width | 0 | 0 | 0.0811 | 4 | 0.020275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 1144.7257 | 66568.4228 | 0.017196 | 281.8313 | 16601.9674 | 0.016976 | 4.0617 | 4.0097 | 283.2368 | 16616.1646 | 0.82223 | 0.99225 | 7.394 | 7.3928 | 0.73536 | 0.053976 | 7.3201 | 7.2874 | 7.4908 | 7.5351 | 0.17067 | 0.24772 | 0.063054 | 0.062043 | 87.9119 | 181.8551 | 0.0083114 | 0.00035989 | FALSE | 1.0002 | TRUE | [M-H]- | 113.5138 | 7.3334 | 7.4625 | 0.12911 | 7.3398 | 7.4547 | 0.1149 | -13.7251 | 1.1646 | 1.3467 | 14.4826 | 0 | 21 | 9 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1144.7257 | No Control Sample | 0 | off | off | 1144.7257 | 0 | off | 0 | off | 0 | off | 0 | off | 398.939 | 6.051 | 2.414 | 79.9567 | -6.1392 | -0.49088 | 0 | 0 | 0 | 0 | 0 | 241.7001 | FALSE | 0.020275 | 0.017196
2149 | 2150 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | Br-PFHxS | 0.0108102243062017 | Qualifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.0189 | 4 | 0.004725 | 0 | 7.16 | 7.35 | Manual | Baseline | 182.6293 | 66568.4228 | 0.0027435 | 45.6558 | 16601.9674 | 0.00275 | 4.0001 | 4.0097 | 283.2368 | 16616.1646 | 0.13118 | 0.99225 | 7.2054 | 7.3928 | 0.63438 | 0.045422 | 7.1537 | 7.2874 | 7.3198 | 7.5351 | 0.16607 | 0.24772 | 0.089476 | 0.062043 | 14.2415 | 181.8551 | 0.0048563 | 0.00035989 | TRUE | 0.97465 | TRUE | [M-H]- | 57.197 | 7.1574 | 7.3184 | 0.161 | 7.1601 | 7.3108 | 0.15075 | -1.3351 | 1.7294 | 2.4415 | 0.15954 | 0 | 20 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 182.6293 | No Control Sample | 0 | off | off | 1144.7257 | 0 | off | 0 | off | 0 | off | 0 | off | 398.939 | 6.051 | 2.414 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 241.7001 | FALSE | 0.004725 | 0.0027435
2151 | 2152 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFHpS | 0.0983151575841707 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7.98 | 7.35 | Baseline | Baseline | 996.061 | 66568.4228 | 0.014963 | 234.6601 | 16601.9674 | 0.014134 | 4.2447 | 4.0097 | 235.2815 | 16616.1646 | 0.9191 | 0.99225 | 8.0187 | 7.3928 | 0.4855 | 0.038743 | 7.915 | 7.2874 | 8.1297 | 7.5351 | 0.2147 | 0.24772 | 0.068551 | 0.062043 | 62.7544 | 181.8551 | 0.0049781 | 0.00035989 | FALSE | 1.0847 | TRUE | [M-H]- | 98.3152 | 7.9442 | 8.0864 | 0.14213 | 7.9574 | 8.0795 | 0.12208 | -5.4409 | 0.98334 | 1.0645 | 19.418 | 0 | 16 | 6 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 996.061 | No Control Sample | 0 | off | off | 996.061 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9356 | 4.9441 | 2.2196 | 79.9562 | -12.7967 | -1.0232 | 0 | 0 | 0 | 0 | 0 | 3.2358 | FALSE | 0.025 | 0.014963
2153 | 2154 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | L-PFOS | 0.0680677554289867 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 0.0788 | 4 | 0.0197 | 0 | 8.46 | 8.45 | Manual | Baseline | 746.5879 | 42197.2566 | 0.017693 | 153.9895 | 9010.6688 | 0.01709 | 4.8483 | 4.683 | 154.8193 | 9016.6956 | 0.60169 | 0.99441 | 8.4908 | 8.4973 | 0.36423 | 0.030814 | 8.4153 | 8.348 | 8.5948 | 8.6566 | 0.17954 | 0.30853 | 0.074488 | 0.072562 | 30.7131 | 164.8092 | 0.010104 | 1.617e-06 | TRUE | 0.99923 | TRUE | [M-H]- | 86.3804 | 8.4152 | 8.5794 | 0.16428 | 8.4152 | 8.5644 | 0.14923 | -8.6665 | 0.97903 | 0.77865 | 16.4936 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 746.5879 | No Control Sample | 0 | off | off | 746.5879 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9333 | 6.0845 | 3.0357 | 79.9564 | -11.7062 | -0.93599 | 0 | 0 | 0 | 0 | 0 | 6.5856 | FALSE | 0.0197 | 0.017693
2155 | 2156 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | Br-PFOS | 0.0122504416124025 | Qualifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.0211 | 4 | 0.005275 | 0 | 8.3 | 8.45 | Manual | Baseline | 287.6559 | 42197.2566 | 0.0068169 | 55.8693 | 9010.6688 | 0.0062003 | 5.1487 | 4.683 | 154.8193 | 9016.6956 | 0.24114 | 0.99441 | 8.3544 | 8.4973 | 0.65533 | 0.054393 | 8.2603 | 8.348 | 8.4161 | 8.6566 | 0.15583 | 0.30853 | 0.084016 | 0.072562 | 11.1431 | 164.8092 | 0.013925 | 1.617e-06 | TRUE | 0.98318 | TRUE | [M-H]- | 58.059 | 8.2729 | 8.4365 | 0.16358 | 8.2813 | 8.4365 | 0.1552 | 4.9927 | 0.95073 | 0.9987 | 0.38529 | 0 | 11 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 287.6559 | No Control Sample | 0 | off | off | 746.5879 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9333 | 6.0845 | 3.0357 | 79.9586 | 16.7221 | 1.3371 | 0 | 0 | 0 | 0 | 0 | 6.5856 | FALSE | 0.005275 | 0.0068169
2157 | 2158 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFNS | 0.104496285174823 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.85 | 8.45 | Baseline | Baseline | 977.3389 | 42197.2566 | 0.023161 | 202.5914 | 9010.6688 | 0.022484 | 4.8242 | 4.683 | 202.6426 | 9016.6956 | 0.93552 | 0.99441 | 8.8865 | 8.4973 | 0.41203 | 0.036465 | 8.763 | 8.348 | 9.0372 | 8.6566 | 0.27417 | 0.30853 | 0.076348 | 0.072562 | 51.8372 | 164.8092 | 0.00046523 | 1.617e-06 | FALSE | 1.0458 | TRUE | [M-H]- | 104.4963 | 8.8115 | 8.97 | 0.15849 | 8.8215 | 8.9604 | 0.13892 | -0.34377 | 1.0287 | 1.072 | 18.8377 | 0 | 14 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 977.3389 | No Control Sample | 0 | off | off | 977.3389 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9294 | 4.336 | 2.3801 | 79.9573 | 2.9141 | 0.233 | 0 | 0 | 0 | 0 | 0 | 5.3315 | FALSE | 0.025 | 0.023161
2159 | 2160 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | PFDS | 0.102788582978293 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.19 | 8.45 | Baseline | Baseline | 987.7643 | 42197.2566 | 0.023408 | 239.6908 | 9010.6688 | 0.026601 | 4.121 | 4.683 | 242.9549 | 9016.6956 | 0.94498 | 0.99441 | 9.2167 | 8.4973 | 0.29045 | 0.026693 | 9.1138 | 8.348 | 9.3119 | 8.6566 | 0.19803 | 0.30853 | 0.064771 | 0.072562 | 108.7736 | 164.8092 | 0.025872 | 1.617e-06 | FALSE | 1.0847 | TRUE | [M-H]- | 102.7886 | 9.1528 | 9.2901 | 0.13725 | 9.1588 | 9.2828 | 0.12392 | 31.3139 | 1.159 | 1.3291 | 15.9508 | 0 | 11 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Fail | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 987.7643 | No Control Sample | 0 | off | off | 987.7643 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9268 | 5.0186 | 3.0057 | 79.9568 | -1.4218 | -0.11368 | 0 | 0 | 0 | 0 | 0 | 1.272 | FALSE | 0.025 | 0.023408
2163 | 2164 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FBSA | 0.0801410461959236 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 4.6 | 8.86 | Manual | Baseline | 1665.8737 | 78834.0919 | 0.021131 | 173.1213 | 16816.2914 | 0.010295 | 9.6226 | 4.688 | 175.4229 | 16816.2937 | 0.93411 | 0.99975 | 4.9851 | 9.0306 | 8.3727 | 0.38514 | 4.7962 | 8.7712 | 5.1922 | 9.3684 | 0.39604 | 0.59725 | 0.14237 | 0.072817 | 56.6941 | 168.9909 | 0.0050919 | 3.1657e-07 | TRUE | 0.55203 | TRUE | [M-H]- | 80.141 | 4.797 | 5.1869 | 0.38995 | 4.803 | 5.183 | 0.38009 | -2.2258 | 1.2719 | 1.5802 | 5.0917 | 0 | 72 | 27 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1665.8737 | No Control Sample | 0 | off | off | 1665.8737 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9615 | 8.5199 | 2.5386 | 77.9646 | -8.6545 | -0.67475 | 0 | 0 | 0 | 0 | 0 | 1.8436 | FALSE | 0.025 | 0.021131
2165 | 2166 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FPeSA | 0.0995440044297835 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7 | 8.86 | Baseline | Baseline | 3213.2796 | 78834.0919 | 0.04076 | 757.1425 | 16816.2914 | 0.045024 | 4.244 | 4.688 | 760.3711 | 16816.2937 | 0.96904 | 0.99975 | 7.1416 | 9.0306 | 2.0226 | 0.14159 | 7.0325 | 8.7712 | 7.2505 | 9.3684 | 0.218 | 0.59725 | 0.064428 | 0.072817 | 172.8027 | 168.9909 | 0.0079491 | 3.1657e-07 | FALSE | 0.79082 | TRUE | [M-H]- | 99.544 | 7.0713 | 7.2213 | 0.15 | 7.0793 | 7.2111 | 0.13179 | 27.6084 | 1.0017 | 0.97286 | 3.6224 | 0 | 24 | 8 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3213.2796 | No Control Sample | 0 | off | off | 3213.2796 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9576 | 5.2749 | 1.8354 | 77.9651 | -2.549 | -0.19873 | 0 | 0 | 0 | 0 | 0 | 4.0788 | FALSE | 0.025 | 0.04076
2167 | 2168 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FHxSA | 0.11316679266834 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7.86 | 8.86 | Baseline | Baseline | 1822.9376 | 78834.0919 | 0.023124 | 389.7255 | 16816.2914 | 0.023175 | 4.6775 | 4.688 | 389.7255 | 16816.2937 | 1 | 0.99975 | 8.0211 | 9.0306 | 2.05 | 0.16113 | 7.847 | 8.7712 | 8.2871 | 9.3684 | 0.44005 | 0.59725 | 0.072131 | 0.072817 | 80.553 | 168.9909 | 0 | 3.1657e-07 | FALSE | 0.88822 | TRUE | [M-H]- | 113.1668 | 7.9515 | 8.0995 | 0.14795 | 7.9624 | 8.0886 | 0.12619 | 0 | 1.1858 | 1.4516 | 7.5435 | 0 | 31 | 6 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1822.9376 | No Control Sample | 0 | off | off | 1822.9376 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9553 | 6.8316 | 2.7186 | 77.9663 | 11.3388 | 0.88403 | 0 | 0 | 0 | 0 | 0 | 2.2611 | FALSE | 0.025 | 0.023124
2169 | 2170 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FHpSA | 0.0935337651327231 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.58 | 8.86 | Baseline | Baseline | 3120.7486 | 78834.0919 | 0.039586 | 691.1952 | 16816.2914 | 0.041103 | 4.515 | 4.688 | 693.6406 | 16816.2937 | 0.96881 | 0.99975 | 8.5916 | 9.0306 | 0.13481 | 0.011566 | 8.4964 | 8.7712 | 8.722 | 9.3684 | 0.22562 | 0.59725 | 0.07079 | 0.072817 | 75.1362 | 168.9909 | 0.0052708 | 3.1657e-07 | FALSE | 0.95138 | TRUE | [M-H]- | 93.5338 | 8.5238 | 8.6699 | 0.14616 | 8.5323 | 8.6601 | 0.12778 | -16.1476 | 1.256 | 1.5741 | 4.2405 | 0 | 14 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3120.7486 | No Control Sample | 0 | off | off | 3120.7486 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9518 | 5.3384 | 2.3913 | 77.9649 | -6.0417 | -0.47105 | 0 | 0 | 0 | 0 | 0 | 2.395 | FALSE | 0.025 | 0.039586
2171 | 2172 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | FOSA | 0.114348893257895 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.88 | 8.86 | Baseline | Baseline | 2051.1221 | 78834.0919 | 0.026018 | 460.8804 | 16816.2914 | 0.027407 | 4.4504 | 4.688 | 460.8804 | 16816.2937 | 0.97921 | 0.99975 | 9.0323 | 9.0306 | 1.7149 | 0.15228 | 8.7821 | 8.7712 | 9.1903 | 9.3684 | 0.40817 | 0.59725 | 0.068462 | 0.072817 | 78.1089 | 168.9909 | 1.8516e-09 | 3.1657e-07 | FALSE | 1.0002 | TRUE | [M-H]- | 114.3489 | 8.952 | 9.104 | 0.15197 | 8.9615 | 9.0931 | 0.13155 | 2.0907e-06 | 0.91154 | 0.78014 | 6.2143 | 0 | 21 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2051.1221 | No Control Sample | 0 | off | off | 2051.1221 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9482 | 4.0119 | 1.9977 | 77.9651 | -3.6134 | -0.28172 | 0 | 0 | 0 | 0 | 0 | 3.3858 | FALSE | 0.025 | 0.026018
2190 | 2191 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 4:2 FTS | 0.0953408531680624 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 5.94 | 5.92 | Baseline | Baseline | 893.8632 | 18464.7219 | 0.048409 | 146.0186 | 3184.8839 | 0.045847 | 6.1216 | 5.7976 | 149.8814 | 3204.7717 | 0.91906 | 0.97431 | 5.9973 | 5.9953 | 0.96423 | 0.057275 | 5.8962 | 5.8651 | 6.1109 | 6.1791 | 0.21472 | 0.31407 | 0.094697 | 0.088983 | 68.1364 | 176.959 | 0.046005 | 0.00216 | FALSE | 1.0003 | TRUE | [M-H]- | 95.3409 | 5.9088 | 6.0962 | 0.18741 | 5.9129 | 6.0864 | 0.17344 | 31.2857 | 0.94746 | 0.82966 | 4.9823 | 0 | 25 | 12 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Fail | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 893.8632 | No Control Sample | 0 | off | off | 893.8632 | 0 | off | 0 | off | 0 | off | 0 | off | 326.977 | 8.3045 | 2.7154 | 306.9691 | 4.9638 | 1.5237 | 0 | 0 | 0 | 0 | 0 | 73.474 | FALSE | 0.025 | 0.048409
2192 | 2193 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 6:2 FTS | 0.105655692113928 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 7.92 | 7.92 | Baseline | Baseline | 1090.691 | 17014.3868 | 0.064104 | 287.246 | 4124.0698 | 0.069651 | 3.7971 | 4.1256 | 287.246 | 4124.0698 | 1 | 1 | 7.9549 | 7.9581 | 0.4403 | 0.034872 | 7.8105 | 7.7608 | 8.1314 | 8.201 | 0.32092 | 0.4402 | 0.058626 | 0.060224 | 319.3793 | 84.8081 | 0 | 0 | FALSE | 0.99959 | TRUE | [M-H]- | 105.6557 | 7.8965 | 8.0262 | 0.12972 | 7.9045 | 8.0164 | 0.11196 | 0 | 1.1651 | 1.3469 | 5.1248 | 0 | 24 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1090.691 | No Control Sample | 0 | off | off | 1090.691 | 0 | off | 0 | off | 0 | off | 0 | off | 426.97 | 4.9544 | 2.1154 | 406.963 | 3.776 | 1.5367 | 0 | 0 | 0 | 0 | 0 | 40.2076 | FALSE | 0.025 | 0.064104
2194 | 2195 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 8:2 FTS | 0.10938138814005 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 8.88 | 8.89 | Baseline | Baseline | 1228.0271 | 22881.8392 | 0.053668 | 241.1417 | 4851.0248 | 0.049709 | 5.0926 | 4.7169 | 241.1417 | 4851.0248 | 1 | 0.99832 | 8.9006 | 8.9012 | 0.23183 | 0.020586 | 8.6874 | 8.6956 | 9.1147 | 9.1238 | 0.42735 | 0.42822 | 0.081098 | 0.072829 | 447.2282 | 172.6798 | 0 | 6.2046e-09 | FALSE | 0.99993 | TRUE | [M-H]- | 109.3814 | 8.8228 | 8.9831 | 0.16031 | 8.8312 | 8.9732 | 0.14192 | 0 | 1.188 | 1.4034 | 4.8018 | 0 | 22 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1228.0271 | No Control Sample | 0 | off | off | 1228.0271 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9634 | 3.6404 | 1.9184 | 506.9604 | 9.1913 | 4.6596 | 0 | 0 | 0 | 0 | 0 | 35.6343 | FALSE | 0.025 | 0.053668
2196 | 2197 | PFAS CS3 0.1ng/ml | 16 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 17) | CS3 | 0 | 0 | Standard | 0 | 0 | 0 | 4 | 1 | 10:2 FTS | 0.0391778067601568 | Quantifiers | 10/26/2023 23:16:42 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 0.1 | 4 | 0.025 | 0 | 9.5 | 8.89 | Manual | Baseline | 457.8931 | 22881.8392 | 0.020011 | 135.3906 | 4851.0248 | 0.02791 | 3.382 | 4.7169 | 136.7652 | 4851.0248 | 0.97291 | 0.99832 | 9.5104 | 8.9012 | 0.10972 | 0.010423 | 9.4444 | 8.6956 | 9.5755 | 9.1238 | 0.13109 | 0.42822 | 0.053454 | 0.072829 | 12663.7772 | 172.6798 | 0 | 6.2046e-09 | TRUE | 1.0684 | TRUE | [M-H]- | 39.1778 | 9.4533 | 9.5656 | 0.11227 | 9.4596 | 9.5584 | 0.098781 | 0 | 0.92928 | 0.82668 | 6.9117 | 0 | 11 | 5 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Fail | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 457.8931 | No Control Sample | 0 | off | off | 457.8931 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9594 | 6.8127 | 4.2712 | 606.9522 | 7.0644 | 4.2877 | 0 | 0 | 0 | 0 | 0 | 17.59 | FALSE | 0.025 | 0.020011
2257 | 2258 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFBA | 0.256560187304429 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 1.49 | 1.48 | Manual | Baseline | 4873.9687 | 105908.7739 | 0.04602 | 722.977 | 15624.2363 | 0.046273 | 6.7415 | 6.7785 | 736.2747 | 15741.1972 | 0.84171 | 0.96717 | 1.4637 | 1.4661 | 1.7638 | 0.026281 | 1.3609 | 1.3444 | 1.6503 | 1.663 | 0.28938 | 0.3186 | 0.10702 | 0.10098 | 61.0159 | 217.4638 | 0.0051188 | 0.011403 | TRUE | 0.9984 | TRUE | [M-H]- | 102.6241 | 1.378 | 1.6281 | 0.25007 | 1.3834 | 1.5946 | 0.2112 | -12.7886 | 1.4448 | 1.6031 | 1.6361 | 0 | 59 | 24 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4873.9687 | No Control Sample | 0 | off | off | 4873.9687 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9808 | 7.7047 | 1.6409 | 168.9898 | 4.7585 | 0.80414 | 0 | 0 | 0 | 0 | 0 | 32.4582 | FALSE | 0.0625 | 0.04602
2259 | 2260 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFPeA | 0.243981968319225 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 3.18 | 3.21 | Manual | Baseline | 4319.9939 | 79173.0487 | 0.054564 | 522.6175 | 8569.7366 | 0.060984 | 8.2661 | 9.2387 | 527.9205 | 8663.4495 | 0.85585 | 0.96378 | 3.195 | 3.1981 | 0.47281 | 0.015035 | 3.0303 | 3.0075 | 3.3792 | 3.4869 | 0.34891 | 0.4794 | 0.13281 | 0.14074 | 96.5319 | 670.3698 | 0 | 0.010133 | TRUE | 0.99903 | TRUE | [M-H]- | 97.5928 | 3.0633 | 3.3667 | 0.30342 | 3.0716 | 3.3499 | 0.27826 | 0 | 1.1665 | 1.2857 | 1.9172 | 0 | 62 | 25 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4319.9939 | No Control Sample | 0 | off | off | 4319.9939 | 0 | off | 0 | off | 0 | off | 0 | off | 262.978 | 7.5816 | 1.9938 | 218.9868 | 4.6846 | 1.0259 | 0 | 0 | 0 | 0 | 0 | 0.53211 | FALSE | 0.0625 | 0.054564
2261 | 2262 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFHxA | 0.263607944499271 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 4.5921 | 6.13 | 6.12 | Baseline | Baseline | 9360.7013 | 152581.4898 | 0.061349 | 1797.1281 | 28439.2359 | 0.063192 | 5.2087 | 5.3652 | 1803.8386 | 28475.0694 | 0.96381 | 0.99157 | 6.1366 | 6.1328 | 0.10821 | 0.0066333 | 6.0086 | 5.9511 | 6.2941 | 6.2976 | 0.2855 | 0.34643 | 0.079942 | 0.083112 | 196.7938 | 314.0599 | 0.0053386 | 0.0019129 | FALSE | 1.0006 | TRUE | [M-H]- | 105.4432 | 6.0452 | 6.2292 | 0.18397 | 6.0567 | 6.2172 | 0.16043 | -33.6049 | 0.97052 | 0.92645 | 1.4799 | 0 | 31 | 10 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9360.7013 | No Control Sample | 203843.5852 | off | off | 9360.7013 | 203843.5852 | off | 4 | off | 0.25 | off | 0.01148 | off | 312.9745 | 5.3757 | 1.6824 | 268.9837 | 4.9919 | 1.3427 | 0 | 0 | 0 | 0 | 0 | 738.1465 | FALSE | 0.0625 | 0.061349
2263 | 2264 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFHpA | 0.267258325914552 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 5.1179 | 7.26 | 7.27 | Baseline | Baseline | 11685.6945 | 174740.9698 | 0.066874 | 2834.0272 | 42451.2027 | 0.06676 | 4.1234 | 4.1163 | 2849.8677 | 42512.0073 | 0.96865 | 0.99185 | 7.2842 | 7.2828 | 0.33355 | 0.024216 | 7.178 | 7.1521 | 7.3971 | 7.3998 | 0.21917 | 0.24765 | 0.062619 | 0.063887 | 148.6642 | 224.2528 | 0.011053 | 0.0024153 | FALSE | 1.0002 | TRUE | [M-H]- | 106.9033 | 7.2143 | 7.3609 | 0.14663 | 7.2272 | 7.3472 | 0.12003 | 142.9237 | 1.0031 | 0.99413 | 1.5497 | 0 | 26 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Pass | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 11685.6945 | No Control Sample | 228328.7703 | off | off | 11685.6945 | 228328.7703 | off | 4 | off | 0.25 | off | 0.012795 | off | 362.9717 | 5.7526 | 2.088 | 318.9809 | 5.5652 | 1.7752 | 0 | 0 | 0 | 0 | 0 | 3.7562 | FALSE | 0.0625 | 0.066874
2265 | 2266 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFOA | 0.28984513529093 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 6.237 | 7.96 | 7.94 | Baseline | Baseline | 11942.2421 | 169891.2985 | 0.070293 | 2994.6386 | 39966.579 | 0.074929 | 3.9879 | 4.2508 | 3002.282 | 39968.0729 | 0.95938 | 0.99873 | 7.9747 | 7.9724 | 0.18529 | 0.014749 | 7.8613 | 7.827 | 8.1275 | 8.1826 | 0.26613 | 0.35555 | 0.061622 | 0.064594 | 119.0112 | 133.3571 | 0.0014366 | 8.8546e-05 | FALSE | 1.0003 | TRUE | [M-H]- | 115.9381 | 7.9142 | 8.0497 | 0.13546 | 7.9203 | 8.0398 | 0.11954 | 16.1647 | 1.1539 | 1.2708 | 1.5941 | 0 | 20 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 11942.2421 | No Control Sample | 191475.4056 | off | off | 11942.2421 | 191475.4056 | off | 4 | off | 0.25 | off | 0.015592 | off | 412.9683 | 4.6528 | 1.9214 | 368.9779 | 4.7868 | 1.7662 | 0 | 0 | 0 | 0 | 0 | 3.491 | FALSE | 0.0625 | 0.070293
2267 | 2268 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFNA | 0.252739148809234 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 5.4677 | 8.46 | 8.47 | Baseline | Baseline | 8701.4356 | 148202.3447 | 0.058713 | 1823.1159 | 32039.3006 | 0.056902 | 4.7728 | 4.6256 | 1823.2115 | 32051.4454 | 0.99321 | 0.99643 | 8.488 | 8.4867 | 0.33147 | 0.028042 | 8.3035 | 8.3465 | 8.6653 | 8.6723 | 0.36175 | 0.32582 | 0.073569 | 0.071672 | 141.366 | 157.8278 | 9.5785e-05 | 0.00044711 | FALSE | 1.0002 | TRUE | [M-H]- | 101.0957 | 8.4032 | 8.5638 | 0.16059 | 8.4145 | 8.5537 | 0.13921 | 0.48273 | 0.86493 | 0.70819 | 1.1642 | 0 | 21 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8701.4356 | No Control Sample | 159143.3885 | off | off | 8701.4356 | 159143.3885 | off | 4 | off | 0.25 | off | 0.013669 | off | 462.9659 | 5.7127 | 2.6448 | 418.9764 | 7.7495 | 3.2468 | 0 | 0 | 0 | 0 | 0 | 7.6234 | FALSE | 0.0625 | 0.058713
2269 | 2270 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFDA | 0.256614539226277 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.87 | 8.86 | Baseline | Baseline | 8014.0233 | 125802.5949 | 0.063703 | 1665.0403 | 25953.6283 | 0.064154 | 4.8131 | 4.8472 | 1665.7878 | 25956.2959 | 0.99071 | 0.99898 | 8.9034 | 8.8984 | 0.37705 | 0.033444 | 8.7601 | 8.6922 | 9.0344 | 9.1587 | 0.27438 | 0.4665 | 0.07532 | 0.075305 | 116.758 | 158.9559 | 0.00082448 | 0.00022817 | FALSE | 1.0006 | TRUE | [M-H]- | 102.6458 | 8.8254 | 8.9855 | 0.16016 | 8.8335 | 8.9719 | 0.13845 | 5.0032 | 0.98499 | 0.89163 | 1.6948 | 0 | 14 | 5 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8014.0233 | No Control Sample | 0 | off | off | 8014.0233 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9623 | 4.4585 | 2.287 | 468.9734 | 7.32 | 3.4329 | 0 | 0 | 0 | 0 | 0 | 331.0178 | FALSE | 0.0625 | 0.063703
2271 | 2272 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFUdA | 0.247636639809318 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.21 | 9.2 | Baseline | Baseline | 5759.5202 | 104875.777 | 0.054918 | 1299.5058 | 22809.1957 | 0.056973 | 4.4321 | 4.598 | 1302.2129 | 22809.1972 | 0.96803 | 0.99977 | 9.2347 | 9.2366 | 0.26816 | 0.024698 | 9.1113 | 9.0435 | 9.3795 | 9.5147 | 0.26825 | 0.47118 | 0.069774 | 0.072526 | 115.9826 | 175.2443 | 0.0045479 | 1.5836e-07 | FALSE | 0.9998 | TRUE | [M-H]- | 99.0547 | 9.1668 | 9.3108 | 0.14403 | 9.1748 | 9.3028 | 0.12803 | 22.0317 | 1.1476 | 1.3375 | 1.99 | 0 | 15 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5759.5202 | No Control Sample | 0 | off | off | 5759.5202 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9602 | 5.9879 | 3.3709 | 518.9698 | 4.8161 | 2.4994 | 0 | 0 | 0 | 0 | 0 | 1.5731 | FALSE | 0.0625 | 0.054918
2273 | 2274 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFDoA | 0.205867461077026 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.47 | 9.47 | Baseline | Baseline | 5083.4056 | 95420.0438 | 0.053274 | 1293.368 | 23869.7608 | 0.054184 | 3.9304 | 3.9975 | 1293.368 | 23880.9981 | 0.98958 | 0.99553 | 9.5027 | 9.5062 | 0.34572 | 0.03274 | 9.2905 | 9.3696 | 9.6812 | 9.6457 | 0.39068 | 0.27612 | 0.060183 | 0.062735 | 166.0231 | 431.113 | 0 | 0.0009765 | FALSE | 0.99963 | TRUE | [M-H]- | 82.347 | 9.437 | 9.5724 | 0.1354 | 9.4475 | 9.5626 | 0.11512 | 0 | 1.1414 | 1.3598 | 1.6822 | 0 | 26 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5083.4056 | No Control Sample | 0 | off | off | 5083.4056 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9563 | 4.2657 | 2.6147 | 568.9669 | 6.2191 | 3.5385 | 0 | 0 | 0 | 0 | 0 | 1504.6142 | FALSE | 0.0625 | 0.053274
2275 | 2276 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFTrDA | 0.200507244181624 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy, Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.68 | 9.47 | Baseline | Baseline | 3608.9626 | 95420.0438 | 0.037822 | 1045.3412 | 23869.7608 | 0.043794 | 3.4524 | 3.9975 | 1045.6466 | 23880.9981 | 0.98676 | 0.99553 | 9.7177 | 9.5062 | 0.38997 | 0.037749 | 9.6126 | 9.3696 | 9.8133 | 9.6457 | 0.20073 | 0.27612 | 0.053799 | 0.062735 | 261.8583 | 431.113 | 0.00054086 | 0.0009765 | FALSE | 1.0222 | TRUE | [M-H]- | 80.2029 | 9.6618 | 9.7797 | 0.11786 | 9.6701 | 9.7711 | 0.10095 | 2.8166 | 1.048 | 1.1053 | 0.91465 | 0 | 18 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3608.9626 | No Control Sample | 0 | off | off | 3608.9626 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9524 | 2.9132 | 1.9313 | 618.9657 | 6.5092 | 4.0289 | 0 | 0 | 0 | 0 | 0 | 5.9593 | FALSE | 0.0625 | 0.037822
2277 | 2278 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFTeDA | 0.197121597277602 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.87 | 9.86 | Baseline | Baseline | 1445.6329 | 20684.7463 | 0.069889 | 422.9289 | 6057.0501 | 0.069824 | 3.4181 | 3.415 | 423.5681 | 6058.1969 | 0.78879 | 0.99702 | 9.9 | 9.8964 | 0.3041 | 0.030015 | 9.8142 | 9.786 | 9.9788 | 10.0029 | 0.16463 | 0.21687 | 0.052073 | 0.052802 | 127.4381 | 742.1003 | 0.0028832 | 0.00036914 | FALSE | 1.0004 | TRUE | [M-H]- | 78.8486 | 9.8414 | 9.9582 | 0.11682 | 9.8481 | 9.9505 | 0.10236 | 7.4068 | 1.0972 | 1.2007 | 1.7583 | 0 | 18 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1445.6329 | No Control Sample | 0 | off | off | 1445.6329 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9528 | 7.7758 | 5.5437 | 668.9582 | 0.43149 | 0.28865 | 0 | 0 | 0 | 0 | 0 | 1426.6338 | FALSE | 0.0625 | 0.069889
2284 | 2285 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFBS | 0.267000698761521 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 3.5 | 3.5 | Manual | Manual | 3058.1448 | 43272.4717 | 0.070672 | 319.8489 | 3998.7265 | 0.079988 | 9.5612 | 10.8216 | 320.8737 | 4008.0624 | 0.94206 | 0.99009 | 3.8365 | 3.8365 | 9.6144 | 0.3365 | 3.6523 | 3.5849 | 4.03 | 4.1547 | 0.37766 | 0.56972 | 0.15238 | 0.16861 | 96.8256 | 760.2128 | 0.0035466 | 0.00096181 | TRUE | 0.99999 | TRUE | [M-H]- | 106.8003 | 3.6907 | 4.0225 | 0.33173 | 3.6975 | 4.018 | 0.32049 | -3.0037 | 1.1122 | 1.2515 | 10.9132 | 0 | 60 | 25 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3058.1448 | No Control Sample | 0 | off | off | 3058.1448 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9446 | 5.3144 | 1.5887 | 79.9571 | -0.84123 | -0.067263 | 0 | 0 | 0 | 0 | 0 | 5.2009 | FALSE | 0.0625 | 0.070672
2286 | 2287 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFPeS | 0.290672478140449 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 6.38 | 3.5 | Baseline | Manual | 3352.1987 | 43272.4717 | 0.077467 | 640.0768 | 3998.7265 | 0.16007 | 5.2372 | 10.8216 | 640.0784 | 4008.0624 | 0.98468 | 0.99009 | 6.3873 | 3.8365 | 0.11366 | 0.0072515 | 6.2263 | 3.5849 | 6.5311 | 4.1547 | 0.30473 | 0.56972 | 0.083138 | 0.16861 | 135.944 | 760.2128 | 4.1352e-06 | 0.00096181 | FALSE | 1.6649 | TRUE | [M-H]- | 116.269 | 6.3017 | 6.4766 | 0.17491 | 6.3222 | 6.4603 | 0.13806 | 0.0086859 | 0.89599 | 0.79105 | 12.9034 | 0 | 33 | 11 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3352.1987 | No Control Sample | 0 | off | off | 3352.1987 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9418 | 5.7253 | 1.9978 | 79.9568 | -5.4193 | -0.43332 | 0 | 0 | 0 | 0 | 0 | 0.66552 | FALSE | 0.0625 | 0.077467
2288 | 2289 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | L-PFHxS | 0.206613469802866 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.20275 | 4 | 0.050688 | 0 | 7.34 | 7.35 | Baseline | Baseline | 2471.8868 | 64048.1159 | 0.038594 | 567.3024 | 15728.6454 | 0.036068 | 4.3573 | 4.0721 | 573.3185 | 15733.3795 | 0.78234 | 0.99376 | 7.3592 | 7.3594 | 0.26143 | 0.019189 | 7.2632 | 7.2194 | 7.4589 | 7.5087 | 0.19568 | 0.28932 | 0.067843 | 0.062709 | 66.0955 | 182.7028 | 0.021176 | 0.00058845 | FALSE | 0.99998 | TRUE | [M-H]- | 101.9055 | 7.2818 | 7.4361 | 0.1543 | 7.2979 | 7.4263 | 0.12841 | -61.3909 | 0.92327 | 0.90275 | 13.4843 | 0 | 24 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2471.8868 | No Control Sample | 0 | off | off | 2471.8868 | 0 | off | 0 | off | 0 | off | 0 | off | 398.938 | 3.5878 | 1.4313 | 79.9563 | -11.8749 | -0.94948 | 0 | 0 | 0 | 0 | 0 | 5.3257 | FALSE | 0.050688 | 0.038594
2290 | 2291 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | Br-PFHxS | 0.0344394348971285 | Qualifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.04725 | 4 | 0.011813 | 0 | 7.16 | 7.35 | Manual | Baseline | 559.796 | 64048.1159 | 0.0087402 | 140.8035 | 15728.6454 | 0.008952 | 3.9757 | 4.0721 | 573.3185 | 15733.3795 | 0.17757 | 0.99376 | 7.1852 | 7.3594 | 0.35244 | 0.025235 | 7.0754 | 7.2194 | 7.2659 | 7.5087 | 0.19049 | 0.28932 | 0.065542 | 0.062709 | 16.4048 | 182.7028 | 0.0034959 | 0.00058845 | TRUE | 0.97634 | TRUE | [M-H]- | 72.8877 | 7.0888 | 7.2718 | 0.18306 | 7.1154 | 7.2582 | 0.14277 | 2.5841 | 0.912 | 0.93707 | 0.22647 | 0 | 23 | 9 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 559.796 | No Control Sample | 0 | off | off | 2471.8868 | 0 | off | 0 | off | 0 | off | 0 | off | 398.938 | 3.5878 | 1.4313 | 79.9583 | 13.7479 | 1.0992 | 0 | 0 | 0 | 0 | 0 | 5.3257 | FALSE | 0.011813 | 0.0087402
2292 | 2293 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFHpS | 0.281185700177615 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7.98 | 7.35 | Baseline | Baseline | 2740.9226 | 64048.1159 | 0.042795 | 661.0972 | 15728.6454 | 0.042031 | 4.146 | 4.0721 | 661.0972 | 15733.3795 | 1 | 0.99376 | 7.9977 | 7.3594 | 0.22222 | 0.017733 | 7.824 | 7.2194 | 8.2424 | 7.5087 | 0.41837 | 0.28932 | 0.063982 | 0.062709 | 100.3441 | 182.7028 | 0 | 0.00058845 | FALSE | 1.0867 | TRUE | [M-H]- | 112.4743 | 7.9403 | 8.0777 | 0.13737 | 7.9461 | 8.0644 | 0.11832 | 0 | 1.1284 | 1.149 | 17.3521 | 0 | 30 | 6 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2740.9226 | No Control Sample | 0 | off | off | 2740.9226 | 0 | off | 0 | off | 0 | off | 0 | off | 448.936 | 5.6869 | 2.553 | 79.9564 | -10.2626 | -0.82057 | 0 | 0 | 0 | 0 | 0 | 0.65097 | FALSE | 0.0625 | 0.042795
2294 | 2295 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | L-PFOS | 0.223700555052968 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 0.197 | 4 | 0.04925 | 0 | 8.46 | 8.45 | Manual | Baseline | 2356.4429 | 40526.0745 | 0.058146 | 485.0766 | 8379.2377 | 0.05789 | 4.8579 | 4.8365 | 490.305 | 8380.2381 | 0.69612 | 0.99673 | 8.4885 | 8.4858 | 0.33714 | 0.028522 | 8.3992 | 8.3307 | 8.5847 | 8.6943 | 0.18551 | 0.36355 | 0.075301 | 0.075584 | 54.3801 | 141.642 | 0.010159 | 0.00026632 | TRUE | 1.0003 | TRUE | [M-H]- | 113.5536 | 8.3978 | 8.5626 | 0.16481 | 8.398 | 8.5513 | 0.15332 | -26.5634 | 0.99441 | 0.85558 | 11.3996 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2356.4429 | No Control Sample | 0 | off | off | 2356.4429 | 0 | off | 0 | off | 0 | off | 0 | off | 498.933 | 5.5559 | 2.772 | 79.9571 | -2.3298 | -0.18629 | 0 | 0 | 0 | 0 | 0 | 2.9033 | FALSE | 0.04925 | 0.058146
2296 | 2297 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | Br-PFOS | 0.035193988263395 | Qualifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.05275 | 4 | 0.013187 | 0 | 8.3 | 8.45 | Manual | Baseline | 793.6708 | 40526.0745 | 0.019584 | 148.6919 | 8379.2377 | 0.017745 | 5.3377 | 4.8365 | 490.305 | 8380.2381 | 0.23446 | 0.99673 | 8.3481 | 8.4858 | 0.57992 | 0.048134 | 8.2315 | 8.3307 | 8.4085 | 8.6943 | 0.17704 | 0.36355 | 0.086716 | 0.075584 | 16.6693 | 141.642 | 0.0063283 | 0.00026632 | TRUE | 0.98378 | TRUE | [M-H]- | 66.7185 | 8.2471 | 8.4166 | 0.16958 | 8.2568 | 8.4166 | 0.15986 | 5.3151 | 0.74817 | 0.54288 | 0.33681 | 0 | 12 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 793.6708 | No Control Sample | 0 | off | off | 2356.4429 | 0 | off | 0 | off | 0 | off | 0 | off | 498.933 | 5.5559 | 2.772 | 79.9612 | 48.8565 | 3.9064 | 0 | 0 | 0 | 0 | 0 | 2.9033 | FALSE | 0.013187 | 0.019584
2298 | 2299 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFNS | 0.303994676130786 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy, Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.85 | 8.45 | Baseline | Baseline | 2730.6161 | 40526.0745 | 0.067379 | 513.5619 | 8379.2377 | 0.06129 | 5.317 | 4.8365 | 517.2851 | 8380.2381 | 0.96753 | 0.99673 | 8.8833 | 8.4858 | 0.3763 | 0.033303 | 8.7644 | 8.3307 | 8.9992 | 8.6943 | 0.23482 | 0.36355 | 0.084912 | 0.075584 | 90.0526 | 141.642 | 0.01242 | 0.00026632 | FALSE | 1.0468 | TRUE | [M-H]- | 121.5979 | 8.8027 | 8.9714 | 0.16865 | 8.8105 | 8.9611 | 0.15053 | 27.1637 | 0.965 | 0.8922 | 15.5525 | 0 | 12 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Fail | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2730.6161 | No Control Sample | 0 | off | off | 2730.6161 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9294 | 4.3375 | 2.381 | 79.9568 | -3.5515 | -0.28397 | 0 | 0 | 0 | 0 | 0 | 2.1116 | FALSE | 0.0625 | 0.067379
2300 | 2301 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFDS | 0.223741691976483 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.19 | 8.45 | Baseline | Baseline | 2064.9317 | 40526.0745 | 0.050953 | 444.4474 | 8379.2377 | 0.053042 | 4.6461 | 4.8365 | 451.2224 | 8380.2381 | 0.90171 | 0.99673 | 9.2157 | 8.4858 | 0.27992 | 0.025724 | 9.1156 | 8.3307 | 9.3139 | 8.6943 | 0.19827 | 0.36355 | 0.075671 | 0.075584 | 90.675 | 141.642 | 0.0015779 | 0.00026632 | FALSE | 1.086 | TRUE | [M-H]- | 89.4967 | 9.1463 | 9.292 | 0.14572 | 9.1562 | 9.286 | 0.12985 | 3.537 | 1.0782 | 1.2502 | 17.7504 | 0 | 11 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Fail | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2064.9317 | No Control Sample | 0 | off | off | 2064.9317 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9266 | 4.6749 | 2.7999 | 79.9575 | 6.9589 | 0.55642 | 0 | 0 | 0 | 0 | 0 | 3.3116 | FALSE | 0.0625 | 0.050953
2304 | 2305 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FBSA | 0.188939023651373 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 4.6 | 8.86 | Manual | Baseline | 3654.7102 | 73359.8217 | 0.049819 | 333.0169 | 15966.0971 | 0.020858 | 10.9745 | 4.5947 | 339.0496 | 15966.2197 | 0.93844 | 0.99888 | 4.9376 | 9.0263 | 7.3391 | 0.3376 | 4.744 | 8.8355 | 5.1494 | 9.2595 | 0.4054 | 0.42407 | 0.18082 | 0.07107 | 89.4586 | 180.7225 | 0.0051162 | 1.6656e-05 | TRUE | 0.54703 | TRUE | [M-H]- | 75.5756 | 4.7593 | 5.1418 | 0.38258 | 4.7827 | 5.1263 | 0.34362 | -4.2027 | 1.1195 | 1.3306 | 5.8327 | 0 | 73 | 34 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3654.7102 | No Control Sample | 0 | off | off | 3654.7102 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9606 | 5.3303 | 1.5882 | 77.9654 | 1.1408 | 0.088943 | 0 | 0 | 0 | 0 | 0 | 1.7679 | FALSE | 0.0625 | 0.049819
2306 | 2307 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FPeSA | 0.247500652035828 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7 | 8.86 | Baseline | Baseline | 7434.5375 | 73359.8217 | 0.10134 | 1591.7731 | 15966.0971 | 0.099697 | 4.6706 | 4.5947 | 1591.7732 | 15966.2197 | 0.99588 | 0.99888 | 7.1003 | 9.0263 | 1.4329 | 0.1003 | 6.947 | 8.8355 | 7.2538 | 9.2595 | 0.30673 | 0.42407 | 0.072952 | 0.07107 | 256.809 | 180.7225 | 3.7131e-08 | 1.6656e-05 | FALSE | 0.78663 | TRUE | [M-H]- | 99.0003 | 7.0196 | 7.1817 | 0.16212 | 7.0355 | 7.1734 | 0.13791 | 0.00019269 | 0.97276 | 1.0455 | 3.7904 | 0 | 33 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 7434.5375 | No Control Sample | 0 | off | off | 7434.5375 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9577 | 5.6 | 1.9486 | 77.9648 | -5.831 | -0.45461 | 0 | 0 | 0 | 0 | 0 | 1.4958 | FALSE | 0.0625 | 0.10134
2308 | 2309 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FHxSA | 0.256397285634158 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7.86 | 8.86 | Baseline | Baseline | 3843.3543 | 73359.8217 | 0.05239 | 887.1228 | 15966.0971 | 0.055563 | 4.3324 | 4.5947 | 887.1228 | 15966.2197 | 0.9923 | 0.99888 | 7.9936 | 9.0263 | 1.7002 | 0.13364 | 7.8095 | 8.8355 | 8.2407 | 9.2595 | 0.4311 | 0.42407 | 0.064742 | 0.07107 | 90.8674 | 180.7225 | 0 | 1.6656e-05 | FALSE | 0.8856 | TRUE | [M-H]- | 102.5589 | 7.9229 | 8.074 | 0.15118 | 7.9342 | 8.0627 | 0.12851 | 0 | 0.9989 | 0.99743 | 9.6953 | 0 | 31 | 6 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3843.3543 | No Control Sample | 0 | off | off | 3843.3543 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9549 | 5.7047 | 2.2702 | 77.9643 | -14.724 | -1.148 | 0 | 0 | 0 | 0 | 0 | 3.4577 | FALSE | 0.0625 | 0.05239
2310 | 2311 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FHpSA | 0.236717396363273 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.58 | 8.86 | Baseline | Baseline | 7349.6175 | 73359.8217 | 0.10019 | 1583.9028 | 15966.0971 | 0.099204 | 4.6402 | 4.5947 | 1585.9013 | 15966.2197 | 0.9901 | 0.99888 | 8.5825 | 9.0263 | 0.028739 | 0.0024658 | 8.4576 | 8.8355 | 8.7043 | 9.2595 | 0.2467 | 0.42407 | 0.073148 | 0.07107 | 125.4926 | 180.7225 | 0.0022317 | 1.6656e-05 | FALSE | 0.95083 | TRUE | [M-H]- | 94.687 | 8.5111 | 8.6645 | 0.15334 | 8.519 | 8.6531 | 0.13418 | 14.3286 | 1.0619 | 1.0851 | 4.2687 | 0 | 15 | 7 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 7349.6175 | No Control Sample | 0 | off | off | 7349.6175 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9519 | 5.534 | 2.479 | 77.9656 | 2.3157 | 0.18054 | 0 | 0 | 0 | 0 | 0 | 0.039922 | FALSE | 0.0625 | 0.10019
2312 | 2313 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FOSA | 0.230481691208917 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.88 | 8.86 | Baseline | Baseline | 3847.1593 | 73359.8217 | 0.052442 | 801.2257 | 15966.0971 | 0.050183 | 4.8016 | 4.5947 | 804.1432 | 15966.2197 | 0.96494 | 0.99888 | 9.0274 | 9.0263 | 1.6604 | 0.14745 | 8.9093 | 8.8355 | 9.1331 | 9.2595 | 0.22378 | 0.42407 | 0.076714 | 0.07107 | 105.0218 | 180.7225 | 0.0064178 | 1.6656e-05 | FALSE | 1.0001 | TRUE | [M-H]- | 92.1927 | 8.9493 | 9.1034 | 0.15416 | 8.957 | 9.093 | 0.13599 | 22.9782 | 1.1147 | 1.2151 | 8.6253 | 0 | 12 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3847.1593 | No Control Sample | 0 | off | off | 3847.1593 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9487 | 4.9495 | 2.4646 | 77.9651 | -3.4712 | -0.27064 | 0 | 0 | 0 | 0 | 0 | 2.5601 | FALSE | 0.0625 | 0.052442
2331 | 2332 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 4:2 FTS | 0.227071754165129 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 5.94 | 5.92 | Baseline | Baseline | 2128.9517 | 18465.1754 | 0.1153 | 377.9647 | 2965.4323 | 0.12746 | 5.6327 | 6.2268 | 390.7784 | 2990.4606 | 0.91398 | 0.96728 | 5.9367 | 5.9374 | 0.055016 | 0.003268 | 5.7916 | 5.8094 | 6.03 | 6.0889 | 0.2384 | 0.27953 | 0.084968 | 0.098259 | 118.1564 | 169.8312 | 0.050586 | 0.007842 | FALSE | 0.9999 | TRUE | [M-H]- | 90.8287 | 5.8371 | 6.0175 | 0.18044 | 5.8473 | 6.0111 | 0.16383 | 80.1997 | 0.79712 | 0.59131 | 5.2013 | 0 | 29 | 12 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Fail | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2128.9517 | No Control Sample | 0 | off | off | 2128.9517 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9761 | 5.5718 | 1.8218 | 306.9701 | 8.0945 | 2.4848 | 0 | 0 | 0 | 0 | 0 | 27.2938 | FALSE | 0.0625 | 0.1153
2333 | 2334 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 6:2 FTS | 0.223200173318821 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7.92 | 7.92 | Baseline | Baseline | 2220.9084 | 16399.991 | 0.13542 | 580.4765 | 4002.6283 | 0.14502 | 3.826 | 4.0973 | 583.7616 | 4002.6283 | 0.97617 | 0.99732 | 7.9393 | 7.9352 | 0.24354 | 0.019288 | 7.8385 | 7.748 | 8.0365 | 8.1864 | 0.198 | 0.43843 | 0.060425 | 0.062426 | 114.9554 | 106.119 | 0.010526 | 9.1287e-14 | FALSE | 1.0005 | TRUE | [M-H]- | 89.2801 | 7.8802 | 8.0028 | 0.12265 | 7.8865 | 7.9955 | 0.10897 | 30.8599 | 0.99391 | 0.96847 | 5.2822 | 0 | 16 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2220.9084 | No Control Sample | 0 | off | off | 2220.9084 | 0 | off | 0 | off | 0 | off | 0 | off | 426.97 | 4.8358 | 2.0647 | 406.9644 | 7.1521 | 2.9106 | 0 | 0 | 0 | 0 | 0 | 13.7728 | FALSE | 0.0625 | 0.13542
2335 | 2336 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 8:2 FTS | 0.246998061019286 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.88 | 8.89 | Baseline | Baseline | 2605.6278 | 21500.3351 | 0.12119 | 518.6928 | 4276.1513 | 0.1213 | 5.0235 | 5.028 | 518.6928 | 4276.2996 | 1 | 0.99875 | 8.8925 | 8.8993 | 0.14044 | 0.012471 | 8.6705 | 8.716 | 9.1165 | 9.0676 | 0.44597 | 0.35155 | 0.07759 | 0.078071 | 191.6859 | 124.5943 | 0 | 6.4724e-05 | FALSE | 0.99924 | TRUE | [M-H]- | 98.7992 | 8.8134 | 8.9865 | 0.17311 | 8.8241 | 8.974 | 0.14991 | 0 | 1.1038 | 1.2135 | 4.9597 | 0 | 23 | 5 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2605.6278 | No Control Sample | 0 | off | off | 2605.6278 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9645 | 5.6314 | 2.9675 | 506.9577 | 3.8987 | 1.9765 | 0 | 0 | 0 | 0 | 0 | 1652.4161 | FALSE | 0.0625 | 0.12119
2337 | 2338 | PFAS CS4 0.25ng/ml | 17 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 18) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 10:2 FTS | 0.157370744429157 | Quantifiers | 10/26/2023 23:33:17 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.5 | 8.89 | Manual | Baseline | 1728.2329 | 21500.3351 | 0.080382 | 442.9686 | 4276.1513 | 0.10359 | 3.9015 | 5.028 | 446.4644 | 4276.2996 | 0.97941 | 0.99875 | 9.5173 | 8.8993 | 0.18262 | 0.017349 | 9.4504 | 8.716 | 9.594 | 9.0676 | 0.14357 | 0.35155 | 0.062554 | 0.078071 | 1691.1546 | 124.5943 | 0.0050648 | 6.4724e-05 | TRUE | 1.0695 | TRUE | [M-H]- | 62.9483 | 9.4559 | 9.5827 | 0.12685 | 9.4631 | 9.5773 | 0.11424 | -15.6268 | 1.0941 | 1.2484 | 4.0241 | 0 | 12 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Fail | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1728.2329 | No Control Sample | 0 | off | off | 1728.2329 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9579 | 4.483 | 2.8106 | 606.9494 | 2.5453 | 1.5449 | 0 | 0 | 0 | 0 | 0 | 2.9997 | FALSE | 0.0625 | 0.080382
2398 | 2399 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFBA | 0.276246720265624 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 1.49 | 1.48 | Manual | Baseline | 5303.6655 | 107032.9424 | 0.049552 | 823.5922 | 15293.2171 | 0.053853 | 6.4397 | 6.9987 | 838.6534 | 15402.8707 | 0.80989 | 0.97115 | 1.4733 | 1.4709 | 1.1192 | 0.016676 | 1.3661 | 1.323 | 1.6573 | 1.6623 | 0.2912 | 0.3393 | 0.10322 | 0.10692 | 61.8212 | 218.5314 | 0.0051379 | 0.01585 | TRUE | 1.0016 | TRUE | [M-H]- | 110.4987 | 1.3818 | 1.6372 | 0.25539 | 1.3887 | 1.6017 | 0.21303 | 14.5312 | 1.3907 | 1.5083 | 1.5636 | 0 | 60 | 23 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5303.6655 | No Control Sample | 0 | off | off | 5303.6655 | 0 | off | 0 | off | 0 | off | 0 | off | 212.981 | 8.3162 | 1.7712 | 168.99 | 6.0751 | 1.0266 | 0 | 0 | 0 | 0 | 0 | 33.5598 | FALSE | 0.0625 | 0.049552
2400 | 2401 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFPeA | 0.265039093484323 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 3.18 | 3.21 | Manual | Baseline | 4484.4225 | 75656.9024 | 0.059273 | 500.0521 | 7963.5398 | 0.062793 | 8.9679 | 9.5004 | 508.9481 | 8161.7273 | 0.86669 | 0.93087 | 3.2057 | 3.2039 | 0.80896 | 0.025725 | 3.0276 | 3.0392 | 3.3955 | 3.4404 | 0.36784 | 0.40117 | 0.12676 | 0.15323 | 85.4149 | 476.2256 | 0.0051145 | 0.014459 | TRUE | 1.0006 | TRUE | [M-H]- | 106.0156 | 3.0513 | 3.3909 | 0.33961 | 3.0617 | 3.3855 | 0.32378 | 6.9528 | 1.115 | 1.2814 | 1.7519 | 0 | 66 | 24 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4484.4225 | No Control Sample | 0 | off | off | 4484.4225 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9782 | 8.3215 | 2.1884 | 218.9856 | -0.82592 | -0.18087 | 0 | 0 | 0 | 0 | 0 | 5.877 | FALSE | 0.0625 | 0.059273
2402 | 2403 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFHxA | 0.287383385429158 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 4.8995 | 6.13 | 6.12 | Baseline | Baseline | 9987.3597 | 149327.9125 | 0.066882 | 1699.9229 | 26825.5468 | 0.06337 | 5.8752 | 5.5666 | 1701.6887 | 26858.9744 | 0.97325 | 0.99158 | 6.1501 | 6.1491 | 0.32798 | 0.020105 | 6.0197 | 5.9697 | 6.2967 | 6.3181 | 0.27707 | 0.3484 | 0.08759 | 0.083976 | 124.436 | 251.9065 | 0.0015681 | 0.002393 | FALSE | 1.0002 | TRUE | [M-H]- | 114.9534 | 6.05 | 6.2534 | 0.20335 | 6.062 | 6.2373 | 0.17529 | -9.6208 | 1.0051 | 0.96589 | 1.328 | 0 | 30 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9987.3597 | No Control Sample | 203843.5852 | off | off | 9987.3597 | 203843.5852 | off | 4 | off | 0.25 | off | 0.012249 | off | 312.9746 | 5.6386 | 1.7647 | 268.9842 | 6.7977 | 1.8285 | 0 | 0 | 0 | 0 | 0 | 564.6592 | FALSE | 0.0625 | 0.066882
2404 | 2405 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFHpA | 0.296526237067773 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 5.6235 | 7.26 | 7.27 | Baseline | Baseline | 12840.1384 | 173052.564 | 0.074198 | 3010.3015 | 41367.165 | 0.07277 | 4.2654 | 4.1833 | 3020.6333 | 41393.0832 | 0.97156 | 0.99505 | 7.2955 | 7.2946 | 0.48836 | 0.035455 | 7.1905 | 7.1457 | 7.4181 | 7.4465 | 0.22752 | 0.3008 | 0.066473 | 0.065017 | 190.0367 | 214.5127 | 0.0026842 | 0.0012074 | FALSE | 1.0001 | TRUE | [M-H]- | 118.6105 | 7.2293 | 7.3678 | 0.13842 | 7.2378 | 7.3598 | 0.12206 | -35.5143 | 0.99183 | 0.98973 | 1.4618 | 0 | 27 | 10 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 12840.1384 | No Control Sample | 228328.7703 | off | off | 12840.1384 | 228328.7703 | off | 4 | off | 0.25 | off | 0.014059 | off | 362.9713 | 4.5037 | 1.6347 | 318.9803 | 3.8457 | 1.2267 | 0 | 0 | 0 | 0 | 0 | 4.3798 | FALSE | 0.0625 | 0.074198
2406 | 2407 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFOA | 0.287111980528968 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 6.0547 | 7.96 | 7.94 | Baseline | Baseline | 11593.2926 | 166497.1342 | 0.069631 | 2705.414 | 39195.6545 | 0.069023 | 4.2852 | 4.2478 | 2715.5444 | 39199.5954 | 0.94484 | 0.99855 | 7.9789 | 7.9784 | 0.23688 | 0.018856 | 7.887 | 7.8252 | 8.158 | 8.2132 | 0.27102 | 0.38798 | 0.06652 | 0.066188 | 113.8629 | 133.4458 | 0.0039742 | 0.00024947 | FALSE | 1.0001 | TRUE | [M-H]- | 114.8448 | 7.9165 | 8.0605 | 0.14404 | 7.9238 | 8.0481 | 0.12435 | 39.6722 | 1.0769 | 1.0876 | 1.6337 | 0 | 20 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 11593.2926 | No Control Sample | 191475.4056 | off | off | 11593.2926 | 191475.4056 | off | 4 | off | 0.25 | off | 0.015137 | off | 412.9683 | 4.6249 | 1.9099 | 368.9777 | 4.3753 | 1.6144 | 0 | 0 | 0 | 0 | 0 | 2.7493 | FALSE | 0.0625 | 0.069631
2408 | 2409 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFNA | 0.289408357108277 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 5.8656 | 8.46 | 8.47 | Baseline | Baseline | 9334.652 | 138842.9236 | 0.067232 | 1915.793 | 29279.3416 | 0.065432 | 4.8725 | 4.742 | 1915.793 | 29289.412 | 0.98531 | 0.99653 | 8.4883 | 8.487 | 0.3341 | 0.028265 | 8.2704 | 8.2787 | 8.7204 | 8.7282 | 0.44997 | 0.44958 | 0.075837 | 0.074117 | 119.3644 | 172.4796 | 1.9072e-13 | 0.00073829 | FALSE | 1.0001 | TRUE | [M-H]- | 115.7633 | 8.4114 | 8.5704 | 0.15903 | 8.4176 | 8.5597 | 0.1421 | 8.1204e-10 | 0.94789 | 0.83084 | 1.2666 | 0 | 26 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9334.652 | No Control Sample | 159143.3885 | off | off | 9334.652 | 159143.3885 | off | 4 | off | 0.25 | off | 0.014664 | off | 462.9664 | 6.7944 | 3.1456 | 418.9745 | 3.1283 | 1.3107 | 0 | 0 | 0 | 0 | 0 | 3.7905 | FALSE | 0.0625 | 0.067232
2410 | 2411 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFDA | 0.275803440462963 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.87 | 8.86 | Baseline | Baseline | 8982.2022 | 131190.8076 | 0.068467 | 1748.9229 | 28113.555 | 0.062209 | 5.1358 | 4.6665 | 1751.6711 | 28120.2 | 0.98802 | 0.99759 | 8.8986 | 8.8968 | 0.32199 | 0.028561 | 8.7596 | 8.7491 | 9.0342 | 9.0993 | 0.27457 | 0.35018 | 0.080762 | 0.07219 | 154.8946 | 153.8626 | 0.0028559 | 0.00053632 | FALSE | 1.0002 | TRUE | [M-H]- | 110.3214 | 8.8191 | 8.9862 | 0.16703 | 8.8275 | 8.9735 | 0.14601 | 18.1916 | 0.9576 | 0.85076 | 1.5154 | 0 | 14 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8982.2022 | No Control Sample | 0 | off | off | 8982.2022 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9627 | 5.141 | 2.6371 | 468.9729 | 6.2168 | 2.9155 | 0 | 0 | 0 | 0 | 0 | 355.4995 | FALSE | 0.0625 | 0.068467
2412 | 2413 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFUdA | 0.243514603478808 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.21 | 9.2 | Baseline | Baseline | 6439.2611 | 119238.0358 | 0.054003 | 1386.6779 | 26130.4438 | 0.053068 | 4.6437 | 4.5632 | 1386.8188 | 26130.735 | 0.99088 | 0.99898 | 9.2322 | 9.2336 | 0.24062 | 0.022161 | 9.0728 | 9.0621 | 9.3797 | 9.4355 | 0.30687 | 0.3734 | 0.071961 | 0.071908 | 130.2907 | 130.3945 | 0.0001845 | 2.3447e-05 | FALSE | 0.99985 | TRUE | [M-H]- | 97.4058 | 9.156 | 9.3136 | 0.15758 | 9.1662 | 9.3037 | 0.13757 | 0.83371 | 1.0726 | 1.1737 | 1.6839 | 0 | 17 | 5 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 6439.2611 | No Control Sample | 0 | off | off | 6439.2611 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9599 | 5.3621 | 3.0186 | 518.9702 | 5.6496 | 2.932 | 0 | 0 | 0 | 0 | 0 | 1.8314 | FALSE | 0.0625 | 0.054003
2414 | 2415 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFDoA | 0.214578790358128 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.47 | 9.47 | Baseline | Baseline | 4577.4657 | 82434.8518 | 0.055528 | 1195.8905 | 20720.9419 | 0.057714 | 3.8277 | 3.9783 | 1203.0775 | 20729.7113 | 0.96785 | 0.99548 | 9.501 | 9.5004 | 0.32727 | 0.030992 | 9.4143 | 9.3525 | 9.583 | 9.6464 | 0.1687 | 0.29388 | 0.061397 | 0.061712 | 202.5468 | 268.8995 | 7.656e-05 | 0.0008169 | FALSE | 1.0001 | TRUE | [M-H]- | 85.8315 | 9.4371 | 9.5607 | 0.12364 | 9.446 | 9.5529 | 0.10688 | -0.54272 | 1.0393 | 1.1158 | 1.9039 | 0 | 12 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4577.4657 | No Control Sample | 0 | off | off | 4577.4657 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9559 | 3.6733 | 2.2516 | 568.9668 | 6.0212 | 3.4258 | 0 | 0 | 0 | 0 | 0 | 1631.0071 | FALSE | 0.0625 | 0.055528
2416 | 2417 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFTrDA | 0.24727897793101 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.68 | 9.47 | Baseline | Baseline | 3845.1276 | 82434.8518 | 0.046644 | 1068.3288 | 20720.9419 | 0.051558 | 3.5992 | 3.9783 | 1071.4209 | 20729.7113 | 0.96166 | 0.99548 | 9.7121 | 9.5004 | 0.3321 | 0.032148 | 9.6279 | 9.3525 | 9.8143 | 9.6464 | 0.18642 | 0.29388 | 0.055423 | 0.061712 | 256.9513 | 268.8995 | 0.0016239 | 0.0008169 | FALSE | 1.0223 | TRUE | [M-H]- | 98.9116 | 9.6526 | 9.7725 | 0.11986 | 9.6601 | 9.7629 | 0.10277 | -9.3063 | 1.1018 | 1.1937 | 1.2104 | 0 | 17 | 6 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3845.1276 | No Control Sample | 0 | off | off | 3845.1276 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9527 | 3.4045 | 2.257 | 618.9612 | -0.86717 | -0.53675 | 0 | 0 | 0 | 0 | 0 | 1.3342 | FALSE | 0.0625 | 0.046644
2418 | 2419 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFTeDA | 0.204840990085806 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.87 | 9.86 | Baseline | Baseline | 1609.3726 | 22159.8149 | 0.072626 | 455.0754 | 6538.3802 | 0.069601 | 3.5365 | 3.3892 | 455.0754 | 6539.4391 | 0.91543 | 0.99616 | 9.8932 | 9.8912 | 0.23471 | 0.023166 | 9.7748 | 9.7669 | 10.0233 | 10.0242 | 0.24848 | 0.25732 | 0.053688 | 0.052369 | 294.7573 | 793.3515 | 0 | 0.00034385 | FALSE | 1.0002 | TRUE | [M-H]- | 81.9364 | 9.8379 | 9.9583 | 0.12038 | 9.8431 | 9.9493 | 0.10617 | 0 | 1.042 | 1.0197 | 1.563 | 0 | 28 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1609.3726 | No Control Sample | 0 | off | off | 1609.3726 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9499 | 3.6702 | 2.6166 | 668.9629 | 7.4972 | 5.0153 | 0 | 0 | 0 | 0 | 0 | 1499.9176 | FALSE | 0.0625 | 0.072626
2425 | 2426 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFBS | 0.284706565800884 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 3.5 | 3.5 | Manual | Manual | 3254.6649 | 43189.1701 | 0.075358 | 352.5444 | 4284.6379 | 0.082281 | 9.2319 | 10.08 | 353.7347 | 4293.331 | 0.96577 | 0.98874 | 3.8354 | 3.8378 | 9.5823 | 0.33538 | 3.6488 | 3.6137 | 4.1054 | 4.1488 | 0.45655 | 0.5351 | 0.17475 | 0.14983 | 113.7225 | 914.5739 | 0.0019763 | 0.0007844 | TRUE | 0.99937 | TRUE | [M-H]- | 113.8826 | 3.6885 | 4.0628 | 0.3743 | 3.7052 | 4.055 | 0.34972 | 1.5261 | 1.4223 | 2.046 | 10.5885 | 0 | 72 | 29 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3254.6649 | No Control Sample | 0 | off | off | 3254.6649 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9446 | 5.4281 | 1.6227 | 79.9571 | -1.2785 | -0.10223 | 0 | 0 | 0 | 0 | 0 | 5.2009 | FALSE | 0.0625 | 0.075358
2427 | 2428 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFPeS | 0.266958101955256 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 6.38 | 3.5 | Baseline | Manual | 3072.7845 | 43189.1701 | 0.071147 | 660.4008 | 4284.6379 | 0.15413 | 4.6529 | 10.08 | 662.4184 | 4293.331 | 0.97194 | 0.98874 | 6.4018 | 3.8378 | 0.3414 | 0.021781 | 6.2882 | 3.6137 | 6.5046 | 4.1488 | 0.21633 | 0.5351 | 0.072261 | 0.14983 | 195.6586 | 914.5739 | 0.0021925 | 0.0007844 | FALSE | 1.6681 | TRUE | [M-H]- | 106.7832 | 6.3189 | 6.4887 | 0.16977 | 6.3354 | 6.4818 | 0.14641 | 6.6932 | 1.0341 | 1.232 | 14.1267 | 0 | 24 | 9 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3072.7845 | No Control Sample | 0 | off | off | 3072.7845 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9417 | 5.3554 | 1.8687 | 79.9569 | -3.6368 | -0.29079 | 0 | 0 | 0 | 0 | 0 | 0.7669 | FALSE | 0.0625 | 0.071147
2429 | 2430 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | L-PFHxS | 0.253510178754264 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.20275 | 4 | 0.050688 | 0 | 7.34 | 7.35 | Baseline | Baseline | 2928.0156 | 61832.1558 | 0.047354 | 750.7558 | 14858.3743 | 0.050527 | 3.9001 | 4.1614 | 761.0269 | 14867.6298 | 0.77478 | 0.99253 | 7.3707 | 7.37 | 0.41862 | 0.030726 | 7.2837 | 7.251 | 7.4542 | 7.5186 | 0.17057 | 0.26755 | 0.060352 | 0.064704 | 92.2036 | 204.3123 | 0.012296 | 0.0011831 | FALSE | 1.0001 | TRUE | [M-H]- | 125.0358 | 7.2997 | 7.4341 | 0.13434 | 7.3182 | 7.4282 | 0.10997 | -54.1205 | 0.86262 | 0.8513 | 13.2258 | 0 | 21 | 9 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2928.0156 | No Control Sample | 0 | off | off | 2928.0156 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9385 | 4.7584 | 1.8983 | 79.957 | -2.255 | -0.1803 | 0 | 0 | 0 | 0 | 0 | 96.5941 | FALSE | 0.050688 | 0.047354
2431 | 2432 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | Br-PFHxS | 0.0414652619792557 | Qualifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.04725 | 4 | 0.011813 | 0 | 7.16 | 7.35 | Manual | Baseline | 650.6781 | 61832.1558 | 0.010523 | 150.0154 | 14858.3743 | 0.010096 | 4.3374 | 4.1614 | 761.0269 | 14867.6298 | 0.17283 | 0.99253 | 7.191 | 7.37 | 0.43362 | 0.031047 | 7.0834 | 7.251 | 7.2797 | 7.5186 | 0.19624 | 0.26755 | 0.066004 | 0.064704 | 18.424 | 204.3123 | 0.0058099 | 0.0011831 | TRUE | 0.97571 | TRUE | [M-H]- | 87.7572 | 7.1062 | 7.2837 | 0.1775 | 7.1267 | 7.2707 | 0.14397 | 4.4414 | 1.0304 | 1.1942 | 0.22222 | 0 | 23 | 9 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 650.6781 | No Control Sample | 0 | off | off | 2928.0156 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9385 | 4.7584 | 1.8983 | 79.959 | 22.1532 | 1.7713 | 0 | 0 | 0 | 0 | 0 | 96.5941 | FALSE | 0.011813 | 0.010523
2433 | 2434 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFHpS | 0.27293124380898 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7.98 | 7.35 | Baseline | Baseline | 2568.4128 | 61832.1558 | 0.041538 | 630.8611 | 14858.3743 | 0.042458 | 4.0713 | 4.1614 | 633.0684 | 14867.6298 | 0.96853 | 0.99253 | 8.0039 | 7.37 | 0.29891 | 0.023853 | 7.8896 | 7.251 | 8.0979 | 7.5186 | 0.20823 | 0.26755 | 0.065028 | 0.064704 | 105.0002 | 204.3123 | 0.0063408 | 0.0011831 | FALSE | 1.086 | TRUE | [M-H]- | 109.1725 | 7.9454 | 8.0774 | 0.13194 | 7.953 | 8.0689 | 0.11589 | 19.2101 | 1.2132 | 1.4743 | 18.0756 | 0 | 16 | 6 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2568.4128 | No Control Sample | 0 | off | off | 2568.4128 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9356 | 4.8997 | 2.1996 | 79.9572 | -0.22568 | -0.018045 | 0 | 0 | 0 | 0 | 0 | 0.57767 | FALSE | 0.0625 | 0.041538
2435 | 2436 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | L-PFOS | 0.250086827636408 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Asymmetry, Total Width | 0 | 0 | 0.197 | 4 | 0.04925 | 0 | 8.46 | 8.45 | Manual | Baseline | 2570.1552 | 39537.8635 | 0.065005 | 540.6683 | 8615.5087 | 0.062755 | 4.7537 | 4.5892 | 548.0315 | 8631.3343 | 0.71804 | 0.98977 | 8.4887 | 8.4871 | 0.33976 | 0.028744 | 8.399 | 8.3296 | 8.572 | 8.6246 | 0.173 | 0.29497 | 0.073518 | 0.071147 | 57.6191 | 157.4436 | 0.020374 | 0.0033562 | TRUE | 1.0002 | TRUE | [M-H]- | 126.9476 | 8.3968 | 8.556 | 0.15917 | 8.3968 | 8.5479 | 0.15104 | -63.674 | 0.77431 | 0.46955 | 10.4791 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2570.1552 | No Control Sample | 0 | off | off | 2570.1552 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9332 | 6.0459 | 3.0165 | 79.9562 | -13.8448 | -1.107 | 0 | 0 | 0 | 0 | 0 | 0.9033 | FALSE | 0.04925 | 0.065005
2437 | 2438 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | Br-PFOS | 0.0377682716951786 | Qualifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.05275 | 4 | 0.013187 | 0 | 8.3 | 8.45 | Manual | Baseline | 830.9553 | 39537.8635 | 0.021017 | 168.8957 | 8615.5087 | 0.019604 | 4.9199 | 4.5892 | 548.0315 | 8631.3343 | 0.23215 | 0.98977 | 8.3347 | 8.4871 | 0.41778 | 0.034676 | 8.2345 | 8.3296 | 8.4024 | 8.6246 | 0.16791 | 0.29497 | 0.077755 | 0.071147 | 17.9992 | 157.4436 | 0.011224 | 0.0033562 | TRUE | 0.98204 | TRUE | [M-H]- | 71.5986 | 8.2555 | 8.4156 | 0.1601 | 8.2672 | 8.4156 | 0.14844 | 11.2899 | 0.941 | 1.0221 | 0.32331 | 0 | 12 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 830.9553 | No Control Sample | 0 | off | off | 2570.1552 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9332 | 6.0459 | 3.0165 | 79.957 | -4.0194 | -0.32138 | 0 | 0 | 0 | 0 | 0 | 0.9033 | FALSE | 0.013187 | 0.021017
2439 | 2440 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFNS | 0.253342935443999 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.85 | 8.45 | Baseline | Baseline | 2220.149 | 39537.8635 | 0.056152 | 472.6188 | 8615.5087 | 0.054857 | 4.6975 | 4.5892 | 472.6188 | 8631.3343 | 0.96173 | 0.98977 | 8.8915 | 8.4871 | 0.46861 | 0.041472 | 8.6518 | 8.3296 | 9.0954 | 8.6246 | 0.44363 | 0.29497 | 0.073211 | 0.071147 | 70.7018 | 157.4436 | 0 | 0.0033562 | FALSE | 1.0476 | TRUE | [M-H]- | 101.3372 | 8.8084 | 8.9684 | 0.16001 | 8.819 | 8.9568 | 0.13775 | 0 | 0.98215 | 0.94511 | 19.5338 | 0 | 23 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2220.149 | No Control Sample | 0 | off | off | 2220.149 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9299 | 5.1612 | 2.8331 | 79.9565 | -7.8508 | -0.62773 | 0 | 0 | 0 | 0 | 0 | 1.5956 | FALSE | 0.0625 | 0.056152
2441 | 2442 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFDS | 0.271737419667198 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.19 | 8.45 | Baseline | Baseline | 2446.7347 | 39537.8635 | 0.061883 | 504.0562 | 8615.5087 | 0.058506 | 4.8541 | 4.5892 | 504.1063 | 8631.3343 | 0.9644 | 0.98977 | 9.2126 | 8.4871 | 0.24617 | 0.022623 | 9.0772 | 8.3296 | 9.3495 | 8.6246 | 0.27228 | 0.29497 | 0.076423 | 0.071147 | 88.8023 | 157.4436 | 0.00017602 | 0.0033562 | FALSE | 1.0855 | TRUE | [M-H]- | 108.695 | 9.1365 | 9.2937 | 0.15721 | 9.1438 | 9.2868 | 0.14295 | 0.32586 | 1.0166 | 1.0429 | 16.2713 | 0 | 15 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2446.7347 | No Control Sample | 0 | off | off | 2446.7347 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9264 | 4.3179 | 2.5861 | 79.9569 | 0.052196 | 0.0041734 | 0 | 0 | 0 | 0 | 0 | 1.3492 | FALSE | 0.0625 | 0.061883
2445 | 2446 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FBSA | 0.180237271260891 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 4.6 | 8.86 | Manual | Baseline | 3688.9807 | 77622.7051 | 0.047525 | 321.2108 | 17220.657 | 0.018653 | 11.4846 | 4.5075 | 325.9427 | 17220.7452 | 0.93963 | 0.99917 | 4.9459 | 9.0259 | 7.52 | 0.34592 | 4.7532 | 8.7722 | 5.1624 | 9.2995 | 0.40919 | 0.52732 | 0.20223 | 0.069166 | 71.482 | 167.2825 | 0.020397 | 1.0298e-05 | TRUE | 0.54797 | TRUE | [M-H]- | 72.0949 | 4.7645 | 5.1411 | 0.37662 | 4.7703 | 5.1297 | 0.35939 | -16.0112 | 0.94283 | 0.85334 | 5.7822 | 0 | 74 | 38 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3688.9807 | No Control Sample | 0 | off | off | 3688.9807 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9607 | 5.7023 | 1.699 | 77.9653 | 0.58233 | 0.045402 | 0 | 0 | 0 | 0 | 0 | 3.9742 | FALSE | 0.0625 | 0.047525
2447 | 2448 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FPeSA | 0.259749506114422 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7 | 8.86 | Baseline | Baseline | 8255.87 | 77622.7051 | 0.10636 | 1849.4457 | 17220.657 | 0.1074 | 4.464 | 4.5075 | 1854.8659 | 17220.7452 | 0.98202 | 0.99917 | 7.1177 | 9.0259 | 1.6815 | 0.11771 | 6.9914 | 8.7722 | 7.2294 | 9.2995 | 0.23805 | 0.52732 | 0.070118 | 0.069166 | 222.7022 | 167.2825 | 0.0050492 | 1.0298e-05 | FALSE | 0.78859 | TRUE | [M-H]- | 103.8998 | 7.043 | 7.1981 | 0.15507 | 7.0514 | 7.1846 | 0.13317 | 39.2278 | 0.94667 | 0.81128 | 3.52 | 0 | 26 | 9 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8255.87 | No Control Sample | 0 | off | off | 8255.87 | 0 | off | 0 | off | 0 | off | 0 | off | 347.958 | 6.4182 | 2.2333 | 77.965 | -3.9409 | -0.30726 | 0 | 0 | 0 | 0 | 0 | 2.1631 | FALSE | 0.0625 | 0.10636
2449 | 2450 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FHxSA | 0.26931485845801 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7.86 | 8.86 | Baseline | Baseline | 4271.5728 | 77622.7051 | 0.05503 | 945.414 | 17220.657 | 0.0549 | 4.5182 | 4.5075 | 949.2593 | 17220.7452 | 0.96077 | 0.99917 | 8.0054 | 9.0259 | 1.8504 | 0.14544 | 7.9132 | 8.7722 | 8.1279 | 9.2995 | 0.21467 | 0.52732 | 0.071776 | 0.069166 | 105.1533 | 167.2825 | 0.0051421 | 1.0298e-05 | FALSE | 0.88694 | TRUE | [M-H]- | 107.7259 | 7.9399 | 8.0798 | 0.13991 | 7.9475 | 8.0721 | 0.12461 | -22.6465 | 0.94381 | 0.87282 | 8.9001 | 0 | 16 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4271.5728 | No Control Sample | 0 | off | off | 4271.5728 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9549 | 5.7243 | 2.278 | 77.9656 | 3.0505 | 0.23783 | 0 | 0 | 0 | 0 | 0 | 1.3476 | FALSE | 0.0625 | 0.05503
2451 | 2452 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FHpSA | 0.239722115347006 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.58 | 8.86 | Baseline | Baseline | 7875.41 | 77622.7051 | 0.10146 | 1810.8976 | 17220.657 | 0.10516 | 4.3489 | 4.5075 | 1810.8976 | 17220.7452 | 1 | 0.99917 | 8.586 | 9.0259 | 0.070102 | 0.0060147 | 8.3388 | 8.7722 | 8.8663 | 9.2995 | 0.52743 | 0.52732 | 0.068223 | 0.069166 | 130.7134 | 167.2825 | 0 | 1.0298e-05 | FALSE | 0.95126 | TRUE | [M-H]- | 95.8888 | 8.5137 | 8.6559 | 0.14217 | 8.5217 | 8.6465 | 0.12478 | 0 | 1.03 | 1.045 | 4.0498 | 0 | 29 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 7875.41 | No Control Sample | 0 | off | off | 7875.41 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9515 | 4.7687 | 2.1361 | 77.9652 | -2.1989 | -0.17144 | 0 | 0 | 0 | 0 | 0 | 3.8745 | FALSE | 0.0625 | 0.10146
2453 | 2454 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FOSA | 0.24613747573262 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.88 | 8.86 | Baseline | Baseline | 4347.2237 | 77622.7051 | 0.056005 | 922.1328 | 17220.657 | 0.053548 | 4.7143 | 4.5075 | 922.1369 | 17220.7452 | 0.98487 | 0.99917 | 9.0259 | 9.0259 | 1.6429 | 0.14589 | 8.8672 | 8.7722 | 9.1729 | 9.2995 | 0.30572 | 0.52732 | 0.073665 | 0.069166 | 112.2157 | 167.2825 | 8.7074e-06 | 1.0298e-05 | FALSE | 1 | TRUE | [M-H]- | 98.455 | 8.9456 | 9.1028 | 0.15723 | 8.9555 | 9.0914 | 0.13588 | 0.026264 | 1.0842 | 1.1706 | 7.3458 | 0 | 16 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4347.2237 | No Control Sample | 0 | off | off | 4347.2237 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9484 | 4.4678 | 2.2247 | 77.9651 | -4.1975 | -0.32726 | 0 | 0 | 0 | 0 | 0 | 0.11551 | FALSE | 0.0625 | 0.056005
2472 | 2473 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 4:2 FTS | 0.259634240198729 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 5.94 | 5.92 | Baseline | Baseline | 2466.2167 | 18707.6833 | 0.13183 | 460.9876 | 3175.9562 | 0.14515 | 5.3499 | 5.8904 | 460.9876 | 3179.6415 | 0.99092 | 0.99178 | 5.959 | 5.957 | 0.31996 | 0.019006 | 5.7932 | 5.7941 | 6.1256 | 6.1094 | 0.33233 | 0.3153 | 0.074681 | 0.091126 | 186.7546 | 158.5666 | 7.9262e-13 | 0.0015542 | FALSE | 1.0003 | TRUE | [M-H]- | 103.8537 | 5.8695 | 6.0641 | 0.19453 | 5.8781 | 6.0561 | 0.17804 | 1.0995e-09 | 1.0707 | 1.1638 | 4.3253 | 0 | 39 | 11 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2466.2167 | No Control Sample | 0 | off | off | 2466.2167 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9763 | 6.1431 | 2.0086 | 306.9681 | 1.6347 | 0.5018 | 0 | 0 | 0 | 0 | 0 | 32.5593 | FALSE | 0.0625 | 0.13183
2474 | 2475 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 6:2 FTS | 0.242794511086302 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7.92 | 7.92 | Baseline | Baseline | 2447.6701 | 16615.8095 | 0.14731 | 653.0489 | 3905.6831 | 0.1672 | 3.7481 | 4.2543 | 653.0489 | 3905.7682 | 1 | 0.99875 | 7.9415 | 7.9401 | 0.27149 | 0.021502 | 7.7678 | 7.8025 | 8.1314 | 8.1056 | 0.3636 | 0.30312 | 0.056373 | 0.064719 | 92.7996 | 83.1962 | 0 | 4.5359e-05 | FALSE | 1.0002 | TRUE | [M-H]- | 97.1178 | 7.8702 | 8.0082 | 0.13802 | 7.8819 | 7.9981 | 0.1162 | 0 | 0.98414 | 0.98713 | 4.9134 | 0 | 27 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2447.6701 | No Control Sample | 0 | off | off | 2447.6701 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9706 | 6.2387 | 2.6637 | 406.9639 | 5.8428 | 2.3778 | 0 | 0 | 0 | 0 | 0 | 18.6936 | FALSE | 0.0625 | 0.14731
2476 | 2477 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 8:2 FTS | 0.244347082508946 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.88 | 8.89 | Baseline | Baseline | 2446.5912 | 20407.0697 | 0.11989 | 506.4339 | 3997.8718 | 0.12668 | 4.831 | 5.1045 | 506.4339 | 4001.0833 | 1 | 0.99543 | 8.8934 | 8.8949 | 0.15119 | 0.013426 | 8.67 | 8.7347 | 9.1163 | 9.0298 | 0.44632 | 0.29508 | 0.074527 | 0.079824 | 148.1368 | 176.8863 | 0 | 0.0013839 | FALSE | 0.99983 | TRUE | [M-H]- | 97.7388 | 8.8117 | 8.9778 | 0.16615 | 8.8233 | 8.9701 | 0.14679 | 0 | 1.0394 | 1.1506 | 5.2688 | 0 | 23 | 5 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2446.5912 | No Control Sample | 0 | off | off | 2446.5912 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9646 | 5.8906 | 3.1041 | 506.9575 | 3.4532 | 1.7506 | 0 | 0 | 0 | 0 | 0 | 1599.3473 | FALSE | 0.0625 | 0.11989
2478 | 2479 | PFAS CS4 0.25ng/ml | 18 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 19) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 10:2 FTS | 0.178115846693011 | Quantifiers | 10/26/2023 23:49:51 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.5 | 8.89 | Baseline | Baseline | 1856.591 | 20407.0697 | 0.090978 | 449.5896 | 3997.8718 | 0.11246 | 4.1295 | 5.1045 | 449.5896 | 4001.0833 | 0.98226 | 0.99543 | 9.5135 | 8.8949 | 0.1423 | 0.013518 | 9.368 | 8.7347 | 9.672 | 9.0298 | 0.30405 | 0.29508 | 0.064383 | 0.079824 | 96.7886 | 176.8863 | 1.0281e-07 | 0.0013839 | FALSE | 1.0695 | TRUE | [M-H]- | 71.2463 | 9.4511 | 9.5882 | 0.13701 | 9.4567 | 9.575 | 0.1183 | 0.00015201 | 1.0873 | 1.0598 | 3.706 | 0 | 21 | 7 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1856.591 | No Control Sample | 0 | off | off | 1856.591 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9578 | 4.2682 | 2.676 | 606.9558 | 13.0561 | 7.9244 | 0 | 0 | 0 | 0 | 0 | 6.687 | FALSE | 0.0625 | 0.090978
2539 | 2540 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFBA | 0.27503433454432 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Integration Quality | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 1.49 | 1.48 | Manual | Manual | 5090.7226 | 103188.4271 | 0.049334 | 809.7045 | 14365.1823 | 0.056366 | 6.2871 | 7.1832 | 833.6141 | 14401.4814 | 0.79904 | 0.99062 | 1.4649 | 1.4672 | 1.6814 | 0.025053 | 1.3676 | 1.3403 | 1.6205 | 1.7012 | 0.25289 | 0.36091 | 0.095804 | 0.11732 | 69.0785 | 186.2794 | 0 | 0.00093668 | TRUE | 0.99846 | TRUE | [M-H]- | 110.0137 | 1.3791 | 1.6143 | 0.23513 | 1.3869 | 1.5759 | 0.18904 | 0 | 1.2452 | 1.1812 | 1.5521 | 0 | 52 | 21 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Fail | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5090.7226 | No Control Sample | 0 | off | off | 5090.7226 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9809 | 7.9316 | 1.6893 | 168.9887 | -1.4956 | -0.25275 | 0 | 0 | 0 | 0 | 0 | 25.7019 | FALSE | 0.0625 | 0.049334
2541 | 2542 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFPeA | 0.258382894994192 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 3.18 | 3.21 | Manual | Manual | 4604.2352 | 79679.337 | 0.057785 | 486.9197 | 8292.11 | 0.058721 | 9.4558 | 9.6091 | 492.7371 | 8317.9874 | 0.89231 | 0.98526 | 3.2005 | 3.1973 | 0.64571 | 0.020534 | 3.0334 | 2.9752 | 3.4149 | 3.4779 | 0.38149 | 0.50274 | 0.13276 | 0.1511 | 96.5439 | 586.4978 | 0.015255 | 0.0016212 | TRUE | 1.001 | TRUE | [M-H]- | 103.3532 | 3.0511 | 3.4022 | 0.35111 | 3.0583 | 3.381 | 0.32268 | 19.4716 | 1.1407 | 1.1999 | 1.786 | 0 | 68 | 26 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4604.2352 | No Control Sample | 0 | off | off | 4604.2352 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9778 | 6.883 | 1.8101 | 218.9869 | 5.2421 | 1.148 | 0 | 0 | 0 | 0 | 0 | 3.4246 | FALSE | 0.0625 | 0.057785
2543 | 2544 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFHxA | 0.252194245470497 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 4.414 | 6.13 | 6.12 | Baseline | Baseline | 8997.7088 | 153302.3179 | 0.058693 | 1670.2337 | 26696.7003 | 0.062563 | 5.3871 | 5.7424 | 1680.6193 | 26771.1628 | 0.94433 | 0.98745 | 6.1217 | 6.1225 | 0.13539 | 0.0082995 | 5.9929 | 5.9449 | 6.2454 | 6.2695 | 0.25253 | 0.32468 | 0.083569 | 0.089552 | 194.6978 | 297.2905 | 0.0038095 | 0.0044904 | FALSE | 0.99988 | TRUE | [M-H]- | 100.8777 | 6.0339 | 6.2149 | 0.18102 | 6.0424 | 6.2011 | 0.15871 | 25.1954 | 1.012 | 0.96092 | 1.5076 | 0 | 28 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8997.7088 | No Control Sample | 203843.5852 | off | off | 8997.7088 | 203843.5852 | off | 4 | off | 0.25 | off | 0.011035 | off | 312.9746 | 5.8471 | 1.83 | 268.9841 | 6.4306 | 1.7297 | 0 | 0 | 0 | 0 | 0 | 626.809 | FALSE | 0.0625 | 0.058693
2545 | 2546 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFHpA | 0.276775794417576 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 5.138 | 7.26 | 7.27 | Baseline | Baseline | 11731.5939 | 169394.9169 | 0.069256 | 2825.3269 | 37707.8252 | 0.074927 | 4.1523 | 4.4923 | 2829.1412 | 37723.315 | 0.97409 | 0.99639 | 7.2799 | 7.2788 | 0.27433 | 0.019916 | 7.1225 | 7.1251 | 7.3907 | 7.445 | 0.26827 | 0.31987 | 0.066084 | 0.07042 | 176.5908 | 240.8869 | 0.0021921 | 0.00081704 | FALSE | 1.0001 | TRUE | [M-H]- | 110.7103 | 7.2106 | 7.3501 | 0.13942 | 7.2194 | 7.3394 | 0.11992 | 23.0866 | 1.0031 | 0.97557 | 1.614 | 0 | 31 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Pass | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 11731.5939 | No Control Sample | 228328.7703 | off | off | 11731.5939 | 228328.7703 | off | 4 | off | 0.25 | off | 0.012845 | off | 362.9714 | 5.0053 | 1.8168 | 318.981 | 6.0745 | 1.9376 | 0 | 0 | 0 | 0 | 0 | 3.6242 | FALSE | 0.0625 | 0.069256
2547 | 2548 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFOA | 0.272541861838495 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 6.0105 | 7.96 | 7.94 | Baseline | Baseline | 11508.5743 | 174116.364 | 0.066097 | 2610.8605 | 41264.1402 | 0.063272 | 4.408 | 4.2196 | 2627.2157 | 41268.0981 | 0.92794 | 0.9985 | 7.9684 | 7.9685 | 0.10585 | 0.0084259 | 7.8496 | 7.8163 | 8.1768 | 8.2319 | 0.32715 | 0.41558 | 0.067933 | 0.064449 | 104.845 | 121.1294 | 0.0010791 | 0.00024732 | FALSE | 0.99999 | TRUE | [M-H]- | 109.0167 | 7.9039 | 8.0475 | 0.14352 | 7.9111 | 8.0383 | 0.12721 | -8.6123 | 1.217 | 1.4537 | 1.7367 | 0 | 24 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 11508.5743 | No Control Sample | 191475.4056 | off | off | 11508.5743 | 191475.4056 | off | 4 | off | 0.25 | off | 0.015026 | off | 412.9681 | 4.1485 | 1.7132 | 368.9772 | 3.112 | 1.1482 | 0 | 0 | 0 | 0 | 0 | 1.1595 | FALSE | 0.0625 | 0.066097
2549 | 2550 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFNA | 0.258579298814706 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 5.3688 | 8.46 | 8.47 | Baseline | Baseline | 8544.1387 | 142236.5531 | 0.06007 | 1782.7574 | 29032.5314 | 0.061406 | 4.7927 | 4.8992 | 1782.8119 | 29044.1551 | 0.98802 | 0.99567 | 8.4766 | 8.477 | 0.19564 | 0.016551 | 8.289 | 8.3292 | 8.7412 | 8.6736 | 0.45212 | 0.34432 | 0.074587 | 0.077151 | 125.9312 | 179.6437 | 4.9752e-05 | 0.00037733 | FALSE | 0.99994 | TRUE | [M-H]- | 103.4317 | 8.3983 | 8.5511 | 0.15279 | 8.4067 | 8.5429 | 0.13621 | -0.19618 | 0.96909 | 0.93411 | 1.4335 | 0 | 26 | 5 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8544.1387 | No Control Sample | 159143.3885 | off | off | 8544.1387 | 159143.3885 | off | 4 | off | 0.25 | off | 0.013422 | off | 462.9661 | 6.2577 | 2.8971 | 418.9753 | 5.0483 | 2.1151 | 0 | 0 | 0 | 0 | 0 | 0.83956 | FALSE | 0.0625 | 0.06007
2551 | 2552 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFDA | 0.228410460349265 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.87 | 8.86 | Baseline | Baseline | 7967.0366 | 140508.0097 | 0.056702 | 1600.1594 | 28888.6033 | 0.055391 | 4.9789 | 4.8638 | 1601.1586 | 28892.0249 | 0.99292 | 0.99908 | 8.8879 | 8.8845 | 0.20182 | 0.017902 | 8.7613 | 8.7126 | 9.0542 | 9.0814 | 0.2929 | 0.3688 | 0.077173 | 0.075978 | 157.1204 | 144.1936 | 0.0010812 | 0.00026077 | FALSE | 1.0004 | TRUE | [M-H]- | 91.3642 | 8.809 | 8.9812 | 0.17221 | 8.82 | 8.9673 | 0.14726 | -5.907 | 1.1061 | 1.2032 | 1.7911 | 0 | 15 | 5 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 7967.0366 | No Control Sample | 0 | off | off | 7967.0366 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9625 | 4.799 | 2.4617 | 468.9738 | 8.0026 | 3.753 | 0 | 0 | 0 | 0 | 0 | 356.2369 | FALSE | 0.0625 | 0.056702
2553 | 2554 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFUdA | 0.270166737970712 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.21 | 9.2 | Baseline | Baseline | 5990.5852 | 99986.4534 | 0.059914 | 1266.5091 | 22145.4575 | 0.05719 | 4.73 | 4.515 | 1267.2838 | 22151.6036 | 0.98729 | 0.99738 | 9.2209 | 9.2196 | 0.11843 | 0.010908 | 9.055 | 9.0822 | 9.3458 | 9.4359 | 0.29072 | 0.3537 | 0.073593 | 0.070602 | 151.2753 | 189.6079 | 0.0010115 | 8.0208e-05 | FALSE | 1.0001 | TRUE | [M-H]- | 108.0667 | 9.1435 | 9.3006 | 0.1571 | 9.1552 | 9.2887 | 0.13354 | 4.4067 | 0.90608 | 0.77951 | 1.7873 | 0 | 16 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5990.5852 | No Control Sample | 0 | off | off | 5990.5852 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9594 | 4.4827 | 2.5236 | 518.9693 | 3.9212 | 2.035 | 0 | 0 | 0 | 0 | 0 | 5.0367 | FALSE | 0.0625 | 0.059914
2555 | 2556 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFDoA | 0.228604072451246 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.47 | 9.47 | Baseline | Baseline | 5213.7446 | 88132.9586 | 0.059158 | 1235.2305 | 23306.0149 | 0.053001 | 4.2209 | 3.7816 | 1235.2305 | 23310.1559 | 0.99278 | 0.99828 | 9.489 | 9.4878 | 0.20066 | 0.019002 | 9.2712 | 9.3358 | 9.669 | 9.6982 | 0.39782 | 0.36248 | 0.066507 | 0.058882 | 153.3668 | 412.6724 | 9.8524e-11 | 0.00041905 | FALSE | 1.0001 | TRUE | [M-H]- | 91.4416 | 9.4197 | 9.5536 | 0.1339 | 9.4284 | 9.5461 | 0.11765 | 3.0592e-07 | 0.86738 | 0.71968 | 1.4899 | 0 | 26 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5213.7446 | No Control Sample | 0 | off | off | 5213.7446 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9561 | 4.0281 | 2.469 | 568.9698 | 11.2045 | 6.375 | 0 | 0 | 0 | 0 | 0 | 1475.4157 | FALSE | 0.0625 | 0.059158
2557 | 2558 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFTrDA | 0.241638171483857 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.68 | 9.47 | Baseline | Baseline | 4017.1365 | 88132.9586 | 0.04558 | 1117.333 | 23306.0149 | 0.047942 | 3.5953 | 3.7816 | 1122.2217 | 23310.1559 | 0.96788 | 0.99828 | 9.704 | 9.4878 | 0.24763 | 0.023971 | 9.584 | 9.3358 | 9.7943 | 9.6982 | 0.21023 | 0.36248 | 0.056045 | 0.058882 | 222.1233 | 412.6724 | 0.0051329 | 0.00041905 | FALSE | 1.0228 | TRUE | [M-H]- | 96.6553 | 9.6442 | 9.764 | 0.11982 | 9.6516 | 9.7542 | 0.10265 | 27.2799 | 0.95529 | 0.8566 | 1.1919 | 0 | 18 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4017.1365 | No Control Sample | 0 | off | off | 4017.1365 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9536 | 4.7356 | 3.1395 | 618.9607 | -1.5647 | -0.96852 | 0 | 0 | 0 | 0 | 0 | 1.7982 | FALSE | 0.0625 | 0.04558
2559 | 2560 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFTeDA | 0.222274723758813 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.87 | 9.86 | Baseline | Baseline | 1844.0602 | 23399.7604 | 0.078807 | 527.275 | 7276.1335 | 0.072466 | 3.4973 | 3.216 | 527.275 | 7280.81 | 0.83712 | 0.99335 | 9.8785 | 9.8784 | 0.08651 | 0.0085385 | 9.7751 | 9.7872 | 10.0025 | 9.9956 | 0.22748 | 0.20842 | 0.055113 | 0.049887 | 129.4618 | 804.6343 | 6.9278e-10 | 0.00045496 | FALSE | 1 | TRUE | [M-H]- | 88.9099 | 9.8275 | 9.9408 | 0.11335 | 9.8332 | 9.9329 | 0.09977 | -1.6058e-06 | 0.97541 | 0.90311 | 1.4883 | 0 | 25 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1844.0602 | No Control Sample | 0 | off | off | 1844.0602 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9517 | 6.1636 | 4.3943 | 668.9606 | 4.0087 | 2.6817 | 0 | 0 | 0 | 0 | 0 | 1427.1356 | FALSE | 0.0625 | 0.078807
2566 | 2567 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFBS | 0.261731596912623 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 3.5 | 3.5 | Manual | Manual | 2973.9042 | 42927.626 | 0.069277 | 365.066 | 4060.4239 | 0.089908 | 8.1462 | 10.5722 | 366.2976 | 4075.3396 | 0.95935 | 0.9858 | 3.8363 | 3.8419 | 9.6074 | 0.33626 | 3.6192 | 3.5925 | 4.0642 | 4.1372 | 0.445 | 0.54471 | 0.10111 | 0.16183 | 120.0617 | 1014.5705 | 0.00076011 | 0.00081103 | TRUE | 0.99853 | TRUE | [M-H]- | 104.6926 | 3.6729 | 4.0528 | 0.37987 | 3.6822 | 4.0122 | 0.33001 | 0.62358 | 1.3432 | 1.4975 | 10.9322 | 0 | 70 | 18 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2973.9042 | No Control Sample | 0 | off | off | 2973.9042 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9445 | 4.9054 | 1.4664 | 79.957 | -2.6355 | -0.21073 | 0 | 0 | 0 | 0 | 0 | 5.2009 | FALSE | 0.0625 | 0.069277
2568 | 2569 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFPeS | 0.257117183041301 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 6.38 | 3.5 | Baseline | Manual | 2941.5898 | 42927.626 | 0.068524 | 596.7119 | 4060.4239 | 0.14696 | 4.9297 | 10.5722 | 597.4446 | 4075.3396 | 0.96268 | 0.9858 | 6.3776 | 3.8419 | 0.036945 | 0.0023571 | 6.2283 | 3.5925 | 6.4954 | 4.1372 | 0.26703 | 0.54471 | 0.078715 | 0.16183 | 124.0436 | 1014.5705 | 0.0021204 | 0.00081103 | FALSE | 1.66 | TRUE | [M-H]- | 102.8469 | 6.2924 | 6.4703 | 0.17792 | 6.3042 | 6.4571 | 0.15289 | 4.7383 | 0.98852 | 0.9538 | 14.8182 | 0 | 29 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2941.5898 | No Control Sample | 0 | off | off | 2941.5898 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9418 | 5.8034 | 2.025 | 79.9572 | -0.33222 | -0.026563 | 0 | 0 | 0 | 0 | 0 | 1.8594 | FALSE | 0.0625 | 0.068524
2570 | 2571 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | L-PFHxS | 0.215402969702884 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.20275 | 4 | 0.050688 | 0 | 7.34 | 7.35 | Baseline | Baseline | 2507.4725 | 62319.0598 | 0.040236 | 654.8919 | 15847.8092 | 0.041324 | 3.8288 | 3.9323 | 670.2093 | 15860.0126 | 0.74467 | 0.98984 | 7.3556 | 7.3535 | 0.21199 | 0.01556 | 7.2732 | 7.2208 | 7.4268 | 7.4888 | 0.1536 | 0.26802 | 0.06065 | 0.059861 | 92.1246 | 216.1784 | 0.008672 | 0.0015457 | FALSE | 1.0003 | TRUE | [M-H]- | 106.2407 | 7.2877 | 7.4159 | 0.1282 | 7.2975 | 7.41 | 0.11256 | -36.9742 | 0.90152 | 0.83594 | 15.6265 | 0 | 19 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2507.4725 | No Control Sample | 0 | off | off | 2507.4725 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9387 | 5.3544 | 2.1361 | 79.9566 | -7.7387 | -0.61877 | 0 | 0 | 0 | 0 | 0 | 88.4479 | FALSE | 0.050688 | 0.040236
2572 | 2573 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | Br-PFHxS | 0.0375187746668398 | Qualifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.04725 | 4 | 0.011813 | 0 | 7.16 | 7.35 | Manual | Baseline | 593.3855 | 62319.0598 | 0.0095217 | 116.7497 | 15847.8092 | 0.0073669 | 5.0825 | 3.9323 | 670.2093 | 15860.0126 | 0.17622 | 0.98984 | 7.1817 | 7.3535 | 0.30287 | 0.021685 | 7.0643 | 7.2208 | 7.2675 | 7.4888 | 0.20324 | 0.26802 | 0.06226 | 0.059861 | 16.4234 | 216.1784 | 0.0047837 | 0.0015457 | TRUE | 0.97664 | TRUE | [M-H]- | 79.4048 | 7.0729 | 7.2732 | 0.20038 | 7.078 | 7.2732 | 0.19524 | 2.7479 | 1.0143 | 1.0851 | 0.23665 | 0 | 24 | 9 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 593.3855 | No Control Sample | 0 | off | off | 2507.4725 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9387 | 5.3544 | 2.1361 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 88.4479 | FALSE | 0.011813 | 0.0095217
2574 | 2575 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFHpS | 0.297769741558781 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy, Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7.98 | 7.35 | Baseline | Baseline | 2824.2208 | 62319.0598 | 0.045319 | 700.2933 | 15847.8092 | 0.044189 | 4.0329 | 3.9323 | 701.8197 | 15860.0126 | 0.98134 | 0.98984 | 7.9887 | 7.3535 | 0.10961 | 0.0087467 | 7.9066 | 7.2208 | 8.101 | 7.4888 | 0.19435 | 0.26802 | 0.063782 | 0.059861 | 99.6082 | 216.1784 | 0.0047522 | 0.0015457 | FALSE | 1.0864 | TRUE | [M-H]- | 119.1079 | 7.9338 | 8.0696 | 0.13581 | 7.9403 | 8.0594 | 0.11912 | 17.1235 | 1.2273 | 1.4382 | 17.6042 | 0 | 15 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2824.2208 | No Control Sample | 0 | off | off | 2824.2208 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9357 | 5.0648 | 2.2738 | 79.9567 | -6.2034 | -0.49601 | 0 | 0 | 0 | 0 | 0 | 0.63325 | FALSE | 0.0625 | 0.045319
2576 | 2577 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | L-PFOS | 0.190926536039779 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 0.197 | 4 | 0.04925 | 0 | 8.46 | 8.45 | Manual | Baseline | 2006.7695 | 40436.7155 | 0.049627 | 430.9908 | 8486.0914 | 0.050788 | 4.6562 | 4.7651 | 437.6308 | 8497.2176 | 0.66788 | 0.98764 | 8.4803 | 8.4763 | 0.23941 | 0.020254 | 8.399 | 8.3509 | 8.5657 | 8.6265 | 0.16672 | 0.27563 | 0.074346 | 0.075253 | 51.7116 | 180.738 | 0.02041 | 0.0012301 | TRUE | 1.0005 | TRUE | [M-H]- | 96.917 | 8.3804 | 8.5503 | 0.16992 | 8.4079 | 8.543 | 0.13504 | -52.7617 | 0.84038 | 0.83459 | 14.4055 | 0 | 11 | 5 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2006.7695 | No Control Sample | 0 | off | off | 2006.7695 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9334 | 6.4503 | 3.2182 | 79.9576 | 3.7059 | 0.29632 | 0 | 0 | 0 | 0 | 0 | 1.877 | FALSE | 0.04925 | 0.049627
2578 | 2579 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | Br-PFOS | 0.0342840629082005 | Qualifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.05275 | 4 | 0.013187 | 0 | 8.3 | 8.45 | Manual | Baseline | 771.446 | 40436.7155 | 0.019078 | 140.9064 | 8486.0914 | 0.016604 | 5.4749 | 4.7651 | 437.6308 | 8497.2176 | 0.25675 | 0.98764 | 8.3154 | 8.4763 | 0.18587 | 0.015427 | 8.2238 | 8.3509 | 8.4007 | 8.6265 | 0.17688 | 0.27563 | 0.089319 | 0.075253 | 16.9064 | 180.738 | 0.015293 | 0.0012301 | TRUE | 0.98102 | TRUE | [M-H]- | 64.9935 | 8.2464 | 8.4179 | 0.17151 | 8.2519 | 8.4179 | 0.16599 | 12.1832 | 1.3803 | 1.9322 | 0.38442 | 0 | 13 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 771.446 | No Control Sample | 0 | off | off | 2006.7695 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9334 | 6.4503 | 3.2182 | 79.9566 | -8.555 | -0.68403 | 0 | 0 | 0 | 0 | 0 | 1.877 | FALSE | 0.013187 | 0.019078
2580 | 2581 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFNS | 0.33286024941184 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy, Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.85 | 8.45 | Baseline | Baseline | 2983.3069 | 40436.7155 | 0.073777 | 641.8662 | 8486.0914 | 0.075637 | 4.6479 | 4.7651 | 646.7132 | 8497.2176 | 0.95768 | 0.98764 | 8.8736 | 8.4763 | 0.26623 | 0.023561 | 8.7656 | 8.3509 | 8.9776 | 8.6265 | 0.21193 | 0.27563 | 0.073397 | 0.075253 | 85.263 | 180.738 | 0.013139 | 0.0012301 | FALSE | 1.0469 | TRUE | [M-H]- | 133.1441 | 8.7975 | 8.952 | 0.15444 | 8.8085 | 8.9437 | 0.13522 | 39.7925 | 1.0609 | 1.1872 | 14.7312 | 0 | 11 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Fail | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2983.3069 | No Control Sample | 0 | off | off | 2983.3069 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9296 | 4.6468 | 2.5508 | 79.9572 | 1.8688 | 0.14942 | 0 | 0 | 0 | 0 | 0 | 1.7507 | FALSE | 0.0625 | 0.073777
2582 | 2583 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFDS | 0.288664783564811 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy, Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.19 | 8.45 | Baseline | Baseline | 2658.2383 | 40436.7155 | 0.065738 | 607.7733 | 8486.0914 | 0.07162 | 4.3737 | 4.7651 | 608.6006 | 8497.2176 | 0.95213 | 0.98764 | 9.1997 | 8.4763 | 0.10556 | 0.0097011 | 9.0974 | 8.3509 | 9.3328 | 8.6265 | 0.23538 | 0.27563 | 0.068181 | 0.075253 | 85.5069 | 180.738 | 0.0022731 | 0.0012301 | FALSE | 1.0853 | TRUE | [M-H]- | 115.4659 | 9.1324 | 9.2795 | 0.14708 | 9.1403 | 9.2698 | 0.12949 | -5.8692 | 1.1816 | 1.3839 | 14.5354 | 0 | 13 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Fail | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2658.2383 | No Control Sample | 0 | off | off | 2658.2383 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9262 | 3.8787 | 2.323 | 79.9567 | -2.0093 | -0.16066 | 0 | 0 | 0 | 0 | 0 | 2.7919 | FALSE | 0.0625 | 0.065738
2586 | 2587 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FBSA | 0.209363038979274 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 4.6 | 8.86 | Manual | Baseline | 3759.153 | 68095.2756 | 0.055204 | 443.2425 | 14493.7742 | 0.030582 | 8.481 | 4.6982 | 447.7424 | 14500.9241 | 0.94167 | 0.99455 | 4.9289 | 9.0134 | 7.149 | 0.32885 | 4.7213 | 8.8157 | 5.1263 | 9.1621 | 0.40504 | 0.34637 | 0.14839 | 0.073526 | 97.8288 | 163.1356 | 0 | 0.00087036 | TRUE | 0.54684 | TRUE | [M-H]- | 83.7452 | 4.7298 | 5.1203 | 0.39052 | 4.7589 | 5.0743 | 0.31543 | 0 | 0.99743 | 0.89253 | 5.6345 | 0 | 74 | 28 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3759.153 | No Control Sample | 0 | off | off | 3759.153 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9608 | 6.1727 | 1.8392 | 77.9643 | -12.852 | -1.002 | 0 | 0 | 0 | 0 | 0 | 1.9644 | FALSE | 0.0625 | 0.055204
2588 | 2589 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FPeSA | 0.247777766951008 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7 | 8.86 | Baseline | Baseline | 6908.737 | 68095.2756 | 0.10146 | 1556.7569 | 14493.7742 | 0.10741 | 4.4379 | 4.6982 | 1559.6241 | 14500.9241 | 0.98477 | 0.99455 | 7.0947 | 9.0134 | 1.3533 | 0.09473 | 6.9694 | 8.8157 | 7.2087 | 9.1621 | 0.2393 | 0.34637 | 0.067554 | 0.073526 | 215.5316 | 163.1356 | 0.00046233 | 0.00087036 | FALSE | 0.78713 | TRUE | [M-H]- | 99.1111 | 7.0216 | 7.1775 | 0.15585 | 7.0312 | 7.1665 | 0.13523 | -3.0077 | 1.0625 | 1.1223 | 4.1176 | 0 | 26 | 9 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 6908.737 | No Control Sample | 0 | off | off | 6908.737 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9576 | 5.1922 | 1.8067 | 77.965 | -3.659 | -0.28527 | 0 | 0 | 0 | 0 | 0 | 1.0782 | FALSE | 0.0625 | 0.10146
2590 | 2591 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FHxSA | 0.312514956795288 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7.86 | 8.86 | Baseline | Baseline | 4348.3703 | 68095.2756 | 0.063857 | 998.3636 | 14493.7742 | 0.068882 | 4.3555 | 4.6982 | 998.3636 | 14500.9241 | 0.98969 | 0.99455 | 7.9906 | 9.0134 | 1.6613 | 0.13058 | 7.7854 | 8.8157 | 8.2263 | 9.1621 | 0.44087 | 0.34637 | 0.069223 | 0.073526 | 137.2516 | 163.1356 | 0 | 0.00087036 | FALSE | 0.88653 | TRUE | [M-H]- | 125.006 | 7.9307 | 8.0732 | 0.14246 | 7.9407 | 8.0641 | 0.12339 | 0 | 1.256 | 1.6424 | 8.6653 | 0 | 32 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4348.3703 | No Control Sample | 0 | off | off | 4348.3703 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9552 | 6.4538 | 2.5683 | 77.9643 | -14.4974 | -1.1303 | 0 | 0 | 0 | 0 | 0 | 2.7395 | FALSE | 0.0625 | 0.063857
2592 | 2593 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FHpSA | 0.256260745766771 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.58 | 8.86 | Baseline | Baseline | 7385.4228 | 68095.2756 | 0.10846 | 1698.5612 | 14493.7742 | 0.11719 | 4.348 | 4.6982 | 1702.4868 | 14500.9241 | 0.98342 | 0.99455 | 8.5717 | 9.0134 | 0.097296 | 0.008348 | 8.4587 | 8.8157 | 8.7056 | 9.1621 | 0.24687 | 0.34637 | 0.068471 | 0.073526 | 166.7102 | 163.1356 | 0.0016588 | 0.00087036 | FALSE | 0.95099 | TRUE | [M-H]- | 102.5043 | 8.5034 | 8.649 | 0.14566 | 8.5121 | 8.6402 | 0.12805 | -11.4136 | 1.1378 | 1.317 | 4.3855 | 0 | 15 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 7385.4228 | No Control Sample | 0 | off | off | 7385.4228 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9517 | 5.1771 | 2.3191 | 77.9649 | -6.5317 | -0.50924 | 0 | 0 | 0 | 0 | 0 | 0.48468 | FALSE | 0.0625 | 0.10846
2594 | 2595 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FOSA | 0.354382306528073 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy, Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.88 | 8.86 | Baseline | Baseline | 5490.7856 | 68095.2756 | 0.080634 | 1175.6537 | 14493.7742 | 0.081114 | 4.6704 | 4.6982 | 1175.6537 | 14500.9241 | 0.99706 | 0.99455 | 9.0183 | 9.0134 | 1.5575 | 0.13831 | 8.7848 | 8.8157 | 9.2929 | 9.1621 | 0.5081 | 0.34637 | 0.071815 | 0.073526 | 117.2193 | 163.1356 | 0 | 0.00087036 | FALSE | 1.0005 | TRUE | [M-H]- | 141.7529 | 8.9336 | 9.0936 | 0.15997 | 8.9434 | 9.0838 | 0.14039 | 0 | 0.93307 | 0.84842 | 6.0034 | 0 | 26 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5490.7856 | No Control Sample | 0 | off | off | 5490.7856 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9489 | 5.3591 | 2.6685 | 77.9649 | -6.9277 | -0.54012 | 0 | 0 | 0 | 0 | 0 | 0.54609 | FALSE | 0.0625 | 0.080634
2613 | 2614 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 4:2 FTS | 0.249338799052115 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 5.94 | 5.92 | Manual | Baseline | 2266.4734 | 17902.4103 | 0.1266 | 426.4955 | 3257.3859 | 0.13093 | 5.3142 | 5.4959 | 428.607 | 3269.5115 | 0.97747 | 0.98121 | 5.9377 | 5.9316 | 0.038915 | 0.0023116 | 5.8002 | 5.775 | 6.0799 | 6.1098 | 0.27974 | 0.33482 | 0.081685 | 0.08678 | 95.7367 | 190.282 | 0.0014211 | 0.003664 | TRUE | 1.001 | TRUE | [M-H]- | 99.7355 | 5.8276 | 6.0425 | 0.21487 | 5.8441 | 6.0035 | 0.15938 | 2.1666 | 0.99518 | 0.7426 | 4.9478 | 0 | 35 | 12 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2266.4734 | No Control Sample | 0 | off | off | 2266.4734 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9761 | 5.4337 | 1.7767 | 306.9686 | 3.2416 | 0.99506 | 0 | 0 | 0 | 0 | 0 | 26.3838 | FALSE | 0.0625 | 0.1266
2615 | 2616 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 6:2 FTS | 0.308074483027197 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7.92 | 7.92 | Baseline | Baseline | 2943.1836 | 15745.9586 | 0.18692 | 728.9018 | 4010.1837 | 0.18176 | 4.0378 | 3.9265 | 728.9034 | 4013.7391 | 0.98775 | 0.99111 | 7.9293 | 7.931 | 0.11705 | 0.0092701 | 7.788 | 7.8063 | 8.1186 | 8.077 | 0.33055 | 0.27072 | 0.060233 | 0.059063 | 94.5702 | 90.4991 | 3.6723e-06 | 0.0017245 | FALSE | 0.99978 | TRUE | [M-H]- | 123.2298 | 7.8625 | 8.0058 | 0.14335 | 7.8719 | 7.9944 | 0.12249 | -0.0080978 | 1.032 | 1.0416 | 4.1963 | 0 | 25 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2943.1836 | No Control Sample | 0 | off | off | 2943.1836 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9705 | 5.9749 | 2.5511 | 406.9627 | 2.8484 | 1.1592 | 0 | 0 | 0 | 0 | 0 | 18.8073 | FALSE | 0.0625 | 0.18692
2617 | 2618 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 8:2 FTS | 0.24869845955715 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.88 | 8.89 | Baseline | Baseline | 2504.9636 | 20528.3818 | 0.12202 | 513.3192 | 4289.2406 | 0.11968 | 4.8799 | 4.786 | 513.3192 | 4298.4114 | 1 | 0.98943 | 8.879 | 8.8818 | 0.010704 | 0.00095054 | 8.6545 | 8.7555 | 9.0983 | 9.0311 | 0.44377 | 0.27557 | 0.077538 | 0.074826 | 205.4866 | 168.3078 | 0 | 0.00077569 | FALSE | 0.99969 | TRUE | [M-H]- | 99.4794 | 8.8007 | 8.9583 | 0.15761 | 8.8118 | 8.9508 | 0.13904 | 0 | 1.1036 | 1.3058 | 4.9135 | 0 | 23 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2504.9636 | No Control Sample | 0 | off | off | 2504.9636 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9647 | 5.9568 | 3.139 | 506.9592 | 6.8368 | 3.4659 | 0 | 0 | 0 | 0 | 0 | 1660.7331 | FALSE | 0.0625 | 0.12202
2619 | 2620 | PFAS CS4 0.25ng/ml | 19 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 20) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 10:2 FTS | 0.223103296088878 | Quantifiers | 10/27/2023 00:06:26 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.5 | 8.89 | Baseline | Baseline | 2339.3421 | 20528.3818 | 0.11396 | 621.4425 | 4289.2406 | 0.14488 | 3.7644 | 4.786 | 621.4425 | 4298.4114 | 1 | 0.98943 | 9.5042 | 8.8818 | 0.043763 | 0.0041574 | 9.334 | 8.7555 | 9.6847 | 9.0311 | 0.3507 | 0.27557 | 0.056945 | 0.074826 | 2283.0101 | 168.3078 | 0 | 0.00077569 | FALSE | 1.0701 | TRUE | [M-H]- | 89.2413 | 9.4425 | 9.5824 | 0.13989 | 9.452 | 9.5645 | 0.11256 | 0 | 1.2205 | 1.3505 | 3.1345 | 0 | 24 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2339.3421 | No Control Sample | 0 | off | off | 2339.3421 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9576 | 3.9939 | 2.504 | 606.9541 | 10.2012 | 6.1916 | 0 | 0 | 0 | 0 | 0 | 15.4023 | FALSE | 0.0625 | 0.11396
2680 | 2681 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFBA | 0.256666727122996 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 1.49 | 1.48 | Manual | Manual | 5001.8777 | 108643.0531 | 0.04604 | 677.9664 | 14641.291 | 0.046305 | 7.3778 | 7.4203 | 695.252 | 14680.4883 | 0.82078 | 0.98811 | 1.4685 | 1.4676 | 1.4412 | 0.021475 | 1.3492 | 1.3488 | 1.6309 | 1.6956 | 0.2818 | 0.34683 | 0.11443 | 0.12126 | 54.7148 | 203.6375 | 0 | 0.0011394 | TRUE | 1.0006 | TRUE | [M-H]- | 102.6667 | 1.3644 | 1.621 | 0.25661 | 1.3939 | 1.6124 | 0.21851 | 0 | 1.4381 | 2.6574 | 1.5029 | 0 | 58 | 24 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5001.8777 | No Control Sample | 0 | off | off | 5001.8777 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9808 | 7.5416 | 1.6062 | 168.9883 | -4.1882 | -0.70775 | 0 | 0 | 0 | 0 | 0 | 18.1532 | FALSE | 0.0625 | 0.04604
2682 | 2683 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFPeA | 0.249781088143814 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 3.18 | 3.21 | Manual | Baseline | 4298.0299 | 76941.7127 | 0.055861 | 517.5992 | 8291.8826 | 0.062422 | 8.3038 | 9.2792 | 523.4947 | 8386.907 | 0.886 | 0.95796 | 3.2022 | 3.2045 | 0.69753 | 0.022181 | 3.0314 | 2.9937 | 3.4442 | 3.4349 | 0.41285 | 0.44118 | 0.14292 | 0.14003 | 87.2386 | 626.827 | 0.010165 | 0.016671 | TRUE | 0.99928 | TRUE | [M-H]- | 99.9124 | 3.0581 | 3.4278 | 0.36966 | 3.0943 | 3.3838 | 0.2895 | -12.7438 | 1.4226 | 2.0874 | 1.9114 | 0 | 72 | 27 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4298.0299 | No Control Sample | 0 | off | off | 4298.0299 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9779 | 7.0631 | 1.8574 | 218.9856 | -0.79822 | -0.1748 | 0 | 0 | 0 | 0 | 0 | 5.1163 | FALSE | 0.0625 | 0.055861
2684 | 2685 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFHxA | 0.251694477389734 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 4.4505 | 6.13 | 6.12 | Baseline | Baseline | 9072.0939 | 154876.6017 | 0.058576 | 1724.9814 | 28122.7857 | 0.061338 | 5.2592 | 5.5072 | 1743.7698 | 28159.5714 | 0.93792 | 0.99282 | 6.1481 | 6.1491 | 0.29557 | 0.018119 | 6.0108 | 5.9702 | 6.2766 | 6.2998 | 0.2658 | 0.32955 | 0.081051 | 0.084883 | 152.0988 | 287.0879 | 0.0018181 | 0.002378 | FALSE | 0.99983 | TRUE | [M-H]- | 100.6778 | 6.0578 | 6.239 | 0.18122 | 6.0718 | 6.226 | 0.15423 | 11.7991 | 1.0735 | 1.1881 | 1.585 | 0 | 29 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9072.0939 | No Control Sample | 203843.5852 | off | off | 9072.0939 | 203843.5852 | off | 4 | off | 0.25 | off | 0.011126 | off | 312.9747 | 5.9552 | 1.8638 | 268.9836 | 4.3534 | 1.171 | 0 | 0 | 0 | 0 | 0 | 735.9642 | FALSE | 0.0625 | 0.058576
2686 | 2687 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFHpA | 0.300702066167903 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy, Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 5.5381 | 7.26 | 7.27 | Baseline | Baseline | 12645.0309 | 168056.3558 | 0.075243 | 3076.8644 | 39893.4887 | 0.077127 | 4.1097 | 4.2126 | 3084.9925 | 39905.04 | 0.96856 | 0.99754 | 7.29 | 7.29 | 0.41259 | 0.029954 | 7.1806 | 7.1453 | 7.4001 | 7.48 | 0.21945 | 0.33462 | 0.064394 | 0.064986 | 205.0573 | 244.9052 | 0.0051296 | 0.00064311 | FALSE | 1 | TRUE | [M-H]- | 120.2808 | 7.2227 | 7.3614 | 0.13868 | 7.2327 | 7.3511 | 0.1184 | 71.9214 | 1.0329 | 1.0727 | 1.6202 | 0 | 26 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Fail | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 12645.0309 | No Control Sample | 228328.7703 | off | off | 12645.0309 | 228328.7703 | off | 4 | off | 0.25 | off | 0.013845 | off | 362.9719 | 6.1638 | 2.2373 | 318.9799 | 2.4418 | 0.77887 | 0 | 0 | 0 | 0 | 0 | 4.459 | FALSE | 0.0625 | 0.075243
2688 | 2689 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFOA | 0.246115377729719 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 5.5345 | 7.96 | 7.94 | Baseline | Baseline | 10597.2411 | 177543.7329 | 0.059688 | 2457.1809 | 41731.6033 | 0.058881 | 4.3128 | 4.2544 | 2463.6617 | 41746.5449 | 0.93206 | 0.99657 | 7.9772 | 7.9777 | 0.2158 | 0.017178 | 7.8855 | 7.8376 | 8.0885 | 8.1276 | 0.20303 | 0.29002 | 0.066212 | 0.06511 | 103.2669 | 139.2131 | 0.00090021 | 0.00067903 | FALSE | 0.99993 | TRUE | [M-H]- | 98.4462 | 7.9126 | 8.0564 | 0.14372 | 7.9198 | 8.0477 | 0.12799 | -10.8947 | 1.0417 | 1.069 | 1.8683 | 0 | 16 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10597.2411 | No Control Sample | 191475.4056 | off | off | 10597.2411 | 191475.4056 | off | 4 | off | 0.25 | off | 0.013836 | off | 412.9689 | 6.0968 | 2.5178 | 368.9786 | 6.6459 | 2.4522 | 0 | 0 | 0 | 0 | 0 | 4.5823 | FALSE | 0.0625 | 0.059688
2690 | 2691 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFNA | 0.25875309524291 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 5.3747 | 8.46 | 8.47 | Baseline | Baseline | 8553.5226 | 142297.1285 | 0.06011 | 1838.5088 | 29782.5104 | 0.061731 | 4.6524 | 4.7779 | 1838.5088 | 29809.6544 | 0.99302 | 0.99354 | 8.4885 | 8.4845 | 0.33631 | 0.028452 | 8.2802 | 8.3204 | 8.7159 | 8.6345 | 0.43563 | 0.3141 | 0.072319 | 0.075171 | 99.4176 | 108.9373 | 1.3229e-10 | 0.0017346 | FALSE | 1.0005 | TRUE | [M-H]- | 103.5012 | 8.4092 | 8.567 | 0.15779 | 8.4193 | 8.5549 | 0.13559 | 5.5831e-07 | 0.991 | 0.9509 | 1.3986 | 0 | 25 | 5 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8553.5226 | No Control Sample | 159143.3885 | off | off | 8553.5226 | 159143.3885 | off | 4 | off | 0.25 | off | 0.013437 | off | 462.9657 | 5.2348 | 2.4235 | 418.9745 | 3.0678 | 1.2853 | 0 | 0 | 0 | 0 | 0 | 6.2712 | FALSE | 0.0625 | 0.06011
2692 | 2693 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFDA | 0.285562853236248 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.87 | 8.86 | Baseline | Baseline | 8911.1322 | 125704.6757 | 0.070889 | 1809.6202 | 25405.4956 | 0.071229 | 4.9243 | 4.9479 | 1809.6217 | 25407.146 | 0.99322 | 0.99914 | 8.8988 | 8.8946 | 0.32479 | 0.028809 | 8.6794 | 8.7064 | 9.1634 | 9.1707 | 0.48402 | 0.46438 | 0.075303 | 0.077741 | 101.0613 | 156.2505 | 1.4791e-06 | 0.00016321 | FALSE | 1.0005 | TRUE | [M-H]- | 114.2251 | 8.8177 | 8.9838 | 0.16609 | 8.8254 | 8.9714 | 0.14591 | -0.00553 | 0.99046 | 0.91719 | 1.4951 | 0 | 25 | 5 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8911.1322 | No Control Sample | 0 | off | off | 8911.1322 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9625 | 4.8277 | 2.4764 | 468.9729 | 6.22 | 2.917 | 0 | 0 | 0 | 0 | 0 | 369.7722 | FALSE | 0.0625 | 0.070889
2694 | 2695 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFUdA | 0.261720703389458 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.21 | 9.2 | Baseline | Baseline | 6174.0802 | 106374.609 | 0.058041 | 1311.3097 | 23892.8476 | 0.054883 | 4.7083 | 4.4522 | 1315.4217 | 23903.15 | 0.96294 | 0.99589 | 9.2298 | 9.2298 | 0.21539 | 0.019838 | 9.1043 | 9.0744 | 9.3395 | 9.3805 | 0.23522 | 0.30613 | 0.074887 | 0.069907 | 119.0259 | 168.3603 | 0.0060722 | 0.0008349 | FALSE | 1 | TRUE | [M-H]- | 104.6883 | 9.156 | 9.3103 | 0.15426 | 9.1662 | 9.2996 | 0.13342 | 33.8521 | 1.1609 | 1.3732 | 1.8528 | 0 | 13 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Fail | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 6174.0802 | No Control Sample | 0 | off | off | 6174.0802 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9598 | 5.2261 | 2.9421 | 518.9688 | 2.8614 | 1.485 | 0 | 0 | 0 | 0 | 0 | 3.367 | FALSE | 0.0625 | 0.058041
2696 | 2697 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFDoA | 0.264151843409734 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.47 | 9.47 | Baseline | Baseline | 5297.3347 | 77495.4723 | 0.068357 | 1258.7541 | 19989.524 | 0.062971 | 4.2084 | 3.8768 | 1259.2929 | 19998.1369 | 0.99457 | 0.99526 | 9.4991 | 9.4973 | 0.30727 | 0.029098 | 9.3921 | 9.3641 | 9.6089 | 9.6429 | 0.2168 | 0.27883 | 0.067273 | 0.06019 | 176.7657 | 346.1472 | 0.00087465 | 0.0008815 | FALSE | 1.0002 | TRUE | [M-H]- | 105.6607 | 9.4329 | 9.564 | 0.13107 | 9.4424 | 9.5571 | 0.11468 | 5.0783 | 1.1174 | 1.3322 | 1.7591 | 0 | 15 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 5297.3347 | No Control Sample | 0 | off | off | 5297.3347 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9568 | 5.0755 | 3.1111 | 568.9658 | 4.2706 | 2.4298 | 0 | 0 | 0 | 0 | 0 | 1398.0854 | FALSE | 0.0625 | 0.068357
2698 | 2699 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFTrDA | 0.244615860650832 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.68 | 9.47 | Baseline | Baseline | 3575.8035 | 77495.4723 | 0.046142 | 940.1091 | 19989.524 | 0.04703 | 3.8036 | 3.8768 | 942.894 | 19998.1369 | 0.97454 | 0.99526 | 9.7123 | 9.4973 | 0.3333 | 0.032263 | 9.6098 | 9.3641 | 9.792 | 9.6429 | 0.18225 | 0.27883 | 0.06082 | 0.06019 | 234.9491 | 346.1472 | 0.0054611 | 0.0008815 | FALSE | 1.0226 | TRUE | [M-H]- | 97.8463 | 9.6498 | 9.77 | 0.12026 | 9.6575 | 9.764 | 0.10655 | 28.1702 | 1.0921 | 1.2499 | 1.2504 | 0 | 16 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3575.8035 | No Control Sample | 0 | off | off | 3575.8035 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9546 | 6.2228 | 4.1254 | 618.9635 | 2.8528 | 1.7658 | 0 | 0 | 0 | 0 | 0 | 9.5045 | FALSE | 0.0625 | 0.046142
2700 | 2701 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFTeDA | 0.248630612772541 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.87 | 9.86 | Baseline | Baseline | 1635.8936 | 18557.8147 | 0.088151 | 535.9738 | 5815.8656 | 0.092157 | 3.0522 | 3.1909 | 540.3898 | 5821.6652 | 0.88489 | 0.99195 | 9.8878 | 9.8886 | 0.18065 | 0.01783 | 9.8131 | 9.795 | 9.9612 | 10.0031 | 0.14818 | 0.20812 | 0.044821 | 0.048258 | 156.9908 | 594.1657 | 0.015535 | 0.0022479 | FALSE | 0.99992 | TRUE | [M-H]- | 99.4522 | 9.8328 | 9.9423 | 0.10956 | 9.8381 | 9.9336 | 0.09547 | 56.1907 | 1.0814 | 1.1072 | 1.4359 | 0 | 16 | 6 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1635.8936 | No Control Sample | 0 | off | off | 1635.8936 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9502 | 4.1186 | 2.9364 | 668.9587 | 1.2289 | 0.82208 | 0 | 0 | 0 | 0 | 0 | 1384.1648 | FALSE | 0.0625 | 0.088151
2707 | 2708 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFBS | 0.260051746295685 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 3.5 | 3.5 | Manual | Manual | 2941.7029 | 42737.1046 | 0.068833 | 274.0406 | 3996.4486 | 0.068571 | 10.7346 | 10.6938 | 275.3678 | 4008.4661 | 0.96994 | 0.98644 | 3.8585 | 3.8538 | 10.2416 | 0.35845 | 3.683 | 3.5981 | 4.1243 | 4.1492 | 0.44126 | 0.55111 | 0.16066 | 0.16208 | 123.1812 | 965.6517 | 0.00086272 | 0.0044324 | TRUE | 1.0012 | TRUE | [M-H]- | 104.0207 | 3.6893 | 4.0836 | 0.39432 | 3.6961 | 4.0074 | 0.31134 | -0.53579 | 1.2008 | 0.97839 | 11.3339 | 0 | 70 | 26 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2941.7029 | No Control Sample | 0 | off | off | 2941.7029 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9445 | 5.0506 | 1.5098 | 79.9576 | 5.5423 | 0.44314 | 0 | 0 | 0 | 0 | 0 | 5.2009 | FALSE | 0.0625 | 0.068833
2709 | 2710 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFPeS | 0.265462012277998 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 6.38 | 3.5 | Baseline | Manual | 3023.5811 | 42737.1046 | 0.070748 | 636.7107 | 3996.4486 | 0.15932 | 4.7488 | 10.6938 | 639.7683 | 4008.4661 | 0.9627 | 0.98644 | 6.3994 | 3.8538 | 0.30402 | 0.019397 | 6.2784 | 3.5981 | 6.5142 | 4.1492 | 0.23583 | 0.55111 | 0.076886 | 0.16208 | 186.2719 | 965.6517 | 0.0094357 | 0.0044324 | FALSE | 1.6606 | TRUE | [M-H]- | 106.1848 | 6.3042 | 6.4775 | 0.17327 | 6.3294 | 6.4701 | 0.14067 | -25.4748 | 0.89208 | 0.95545 | 14.5707 | 0 | 26 | 11 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3023.5811 | No Control Sample | 0 | off | off | 3023.5811 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9417 | 5.5271 | 1.9286 | 79.9569 | -3.5604 | -0.28468 | 0 | 0 | 0 | 0 | 0 | 0.50645 | FALSE | 0.0625 | 0.070748
2711 | 2712 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | L-PFHxS | 0.204132561827572 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.20275 | 4 | 0.050688 | 0 | 7.34 | 7.35 | Baseline | Baseline | 2388.9424 | 62651.2588 | 0.038131 | 596.0778 | 15492.1136 | 0.038476 | 4.0078 | 4.0441 | 603.044 | 15497.6118 | 0.74395 | 0.99507 | 7.3694 | 7.364 | 0.4004 | 0.029389 | 7.2744 | 7.2504 | 7.4705 | 7.6053 | 0.19613 | 0.35488 | 0.059271 | 0.061517 | 66.684 | 158.5798 | 0.022123 | 5.6498e-05 | FALSE | 1.0007 | TRUE | [M-H]- | 100.6819 | 7.3026 | 7.4357 | 0.13308 | 7.3114 | 7.4271 | 0.11578 | -67.2336 | 1.0136 | 1.035 | 15.7431 | 0 | 24 | 9 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2388.9424 | No Control Sample | 0 | off | off | 2388.9424 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9385 | 4.7782 | 1.9062 | 79.9577 | 6.0848 | 0.48652 | 0 | 0 | 0 | 0 | 0 | 99.0701 | FALSE | 0.050688 | 0.038131
2713 | 2714 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | Br-PFHxS | 0.0411649678731867 | Qualifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.04725 | 4 | 0.011813 | 0 | 7.16 | 7.35 | Manual | Baseline | 654.5231 | 62651.2588 | 0.010447 | 115.6229 | 15492.1136 | 0.0074633 | 5.6608 | 4.0441 | 603.044 | 15497.6118 | 0.20383 | 0.99507 | 7.1706 | 7.364 | 0.14872 | 0.010648 | 7.0802 | 7.2504 | 7.2762 | 7.6053 | 0.196 | 0.35488 | 0.091369 | 0.061517 | 12.9349 | 158.5798 | 0 | 5.6498e-05 | TRUE | 0.97375 | TRUE | [M-H]- | 87.1216 | 7.1076 | 7.283 | 0.17535 | 7.1124 | 7.273 | 0.1606 | 0 | 1.8923 | 2.8636 | 0.27398 | 0 | 23 | 12 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 654.5231 | No Control Sample | 0 | off | off | 2388.9424 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9385 | 4.7782 | 1.9062 | 79.9622 | 62.5977 | 5.0051 | 0 | 0 | 0 | 0 | 0 | 99.0701 | FALSE | 0.011813 | 0.010447
2715 | 2716 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFHpS | 0.291693148725076 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy, Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7.98 | 7.35 | Baseline | Baseline | 2781.3344 | 62651.2588 | 0.044394 | 680.7293 | 15492.1136 | 0.04394 | 4.0858 | 4.0441 | 681.4442 | 15497.6118 | 0.9899 | 0.99507 | 8.0024 | 7.364 | 0.2806 | 0.022392 | 7.9016 | 7.2504 | 8.1079 | 7.6053 | 0.20627 | 0.35488 | 0.062615 | 0.061517 | 130.1625 | 158.5798 | 0.00225 | 5.6498e-05 | FALSE | 1.0867 | TRUE | [M-H]- | 116.6773 | 7.9378 | 8.0764 | 0.13855 | 7.9456 | 8.069 | 0.12336 | 7.4256 | 1.1899 | 1.4462 | 17.4054 | 0 | 16 | 6 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2781.3344 | No Control Sample | 0 | off | off | 2781.3344 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9357 | 4.9889 | 2.2397 | 79.9569 | -3.1671 | -0.25323 | 0 | 0 | 0 | 0 | 0 | 1.4853 | FALSE | 0.0625 | 0.044394
2717 | 2718 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | L-PFOS | 0.20412066993843 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 0.197 | 4 | 0.04925 | 0 | 8.46 | 8.45 | Manual | Baseline | 2158.4832 | 40682.3813 | 0.053057 | 479.6171 | 8692.5543 | 0.055176 | 4.5004 | 4.6801 | 480.8908 | 8692.5659 | 0.70434 | 0.99889 | 8.4769 | 8.4836 | 0.19988 | 0.016909 | 8.3936 | 8.3056 | 8.56 | 8.7084 | 0.16636 | 0.40287 | 0.071559 | 0.073415 | 49.7971 | 143.5714 | 0.0050385 | 3.0398e-06 | TRUE | 0.99921 | TRUE | [M-H]- | 103.6146 | 8.3894 | 8.5503 | 0.16089 | 8.4095 | 8.5429 | 0.13341 | -14.5258 | 0.97022 | 1.1228 | 14.3473 | 0 | 10 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2158.4832 | No Control Sample | 0 | off | off | 2158.4832 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9337 | 6.9657 | 3.4754 | 79.9572 | -0.8552 | -0.068379 | 0 | 0 | 0 | 0 | 0 | 1.6151 | FALSE | 0.04925 | 0.053057
2719 | 2720 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | Br-PFOS | 0.0336960704267419 | Qualifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.05275 | 4 | 0.013187 | 0 | 8.3 | 8.45 | Manual | Baseline | 762.8216 | 40682.3813 | 0.018751 | 162.6159 | 8692.5543 | 0.018707 | 4.6909 | 4.6801 | 480.8908 | 8692.5659 | 0.24892 | 0.99889 | 8.3239 | 8.4836 | 0.2883 | 0.023929 | 8.2475 | 8.3056 | 8.3911 | 8.7084 | 0.14356 | 0.40287 | 0.07611 | 0.073415 | 16.8839 | 143.5714 | 0.0047045 | 3.0398e-06 | TRUE | 0.98118 | TRUE | [M-H]- | 63.8788 | 8.2555 | 8.4082 | 0.15274 | 8.2612 | 8.4082 | 0.14703 | 5.329 | 0.98005 | 1.0361 | 0.35341 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 762.8216 | No Control Sample | 0 | off | off | 2158.4832 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9337 | 6.9657 | 3.4754 | 79.9574 | 1.5221 | 0.1217 | 0 | 0 | 0 | 0 | 0 | 1.6151 | FALSE | 0.013187 | 0.018751
2721 | 2722 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFNS | 0.254613550920069 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.85 | 8.45 | Baseline | Baseline | 2295.8738 | 40682.3813 | 0.056434 | 516.5913 | 8692.5543 | 0.059429 | 4.4443 | 4.6801 | 517.9112 | 8692.5659 | 0.9731 | 0.99889 | 8.8843 | 8.4836 | 0.38794 | 0.034332 | 8.7803 | 8.3056 | 9.0507 | 8.7084 | 0.27043 | 0.40287 | 0.068834 | 0.073415 | 101.6159 | 143.5714 | 0.0038811 | 3.0398e-06 | FALSE | 1.0472 | TRUE | [M-H]- | 101.8454 | 8.806 | 8.9617 | 0.15569 | 8.8151 | 8.9461 | 0.131 | -7.4138 | 0.98546 | 0.87251 | 18.3139 | 0 | 14 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2295.8738 | No Control Sample | 0 | off | off | 2295.8738 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9297 | 4.872 | 2.6744 | 79.9578 | 8.3954 | 0.67127 | 0 | 0 | 0 | 0 | 0 | 2.7988 | FALSE | 0.0625 | 0.056434
2723 | 2724 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | PFDS | 0.261306800897151 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.19 | 8.45 | Baseline | Baseline | 2420.9249 | 40682.3813 | 0.059508 | 552.8121 | 8692.5543 | 0.063596 | 4.3793 | 4.6801 | 552.875 | 8692.5659 | 0.96346 | 0.99889 | 9.2086 | 8.4836 | 0.20249 | 0.018609 | 9.0896 | 8.3056 | 9.4129 | 8.7084 | 0.3233 | 0.40287 | 0.066253 | 0.073415 | 95.4377 | 143.5714 | 0.00017084 | 3.0398e-06 | FALSE | 1.0855 | TRUE | [M-H]- | 104.5227 | 9.1438 | 9.2997 | 0.15593 | 9.152 | 9.2849 | 0.13294 | -0.29212 | 1.2369 | 1.4247 | 15.8705 | 0 | 18 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2420.9249 | No Control Sample | 0 | off | off | 2420.9249 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9262 | 3.8962 | 2.3335 | 79.9567 | -2.4197 | -0.19347 | 0 | 0 | 0 | 0 | 0 | 4.906 | FALSE | 0.0625 | 0.059508
2727 | 2728 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FBSA | 0.189974663489252 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 4.6 | 8.86 | Manual | Baseline | 3711.4822 | 74093.2583 | 0.050092 | 414.4834 | 16194.4595 | 0.025594 | 8.9545 | 4.5752 | 420.0155 | 16205.5983 | 0.94257 | 0.99406 | 4.9682 | 9.0208 | 8.0039 | 0.36818 | 4.7799 | 8.8304 | 5.1807 | 9.1542 | 0.40085 | 0.32387 | 0.15183 | 0.070703 | 99.0227 | 178.3493 | 0.0050922 | 0.0010529 | TRUE | 0.55075 | TRUE | [M-H]- | 75.9899 | 4.7924 | 5.1624 | 0.37001 | 4.7991 | 5.1556 | 0.35653 | -5.2654 | 1.3272 | 1.6871 | 5.7333 | 0 | 72 | 28 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3711.4822 | No Control Sample | 0 | off | off | 3711.4822 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9608 | 6.2674 | 1.8674 | 77.9649 | -4.5578 | -0.35535 | 0 | 0 | 0 | 0 | 0 | 7.4004 | FALSE | 0.0625 | 0.050092
2729 | 2730 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FPeSA | 0.239453233566322 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7 | 8.86 | Baseline | Baseline | 7264.7177 | 74093.2583 | 0.098048 | 1512.9153 | 16194.4595 | 0.093422 | 4.8018 | 4.5752 | 1513.0416 | 16205.5983 | 0.99169 | 0.99406 | 7.1115 | 9.0208 | 1.5929 | 0.1115 | 6.9706 | 8.8304 | 7.2649 | 9.1542 | 0.29433 | 0.32387 | 0.075303 | 0.070703 | 192.3495 | 178.3493 | 0.00017401 | 0.0010529 | FALSE | 0.78834 | TRUE | [M-H]- | 95.7813 | 7.0389 | 7.1949 | 0.156 | 7.0481 | 7.1856 | 0.1375 | 0.89444 | 1.1658 | 1.3836 | 3.8647 | 0 | 32 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 7264.7177 | No Control Sample | 0 | off | off | 7264.7177 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9578 | 5.7296 | 1.9937 | 77.9652 | -1.2532 | -0.097708 | 0 | 0 | 0 | 0 | 0 | 2.3069 | FALSE | 0.0625 | 0.098048
2731 | 2732 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FHxSA | 0.293364858940237 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7.86 | 8.86 | Baseline | Baseline | 4441.4575 | 74093.2583 | 0.059944 | 967.5237 | 16194.4595 | 0.059744 | 4.5905 | 4.5752 | 967.5237 | 16205.5983 | 1 | 0.99406 | 8.0019 | 9.0208 | 1.8059 | 0.14194 | 7.832 | 8.8304 | 8.2015 | 9.1542 | 0.36953 | 0.32387 | 0.071532 | 0.070703 | 207.8125 | 178.3493 | 0 | 0.0010529 | FALSE | 0.88705 | TRUE | [M-H]- | 117.3459 | 7.9298 | 8.0822 | 0.15242 | 7.9395 | 8.0719 | 0.13236 | 0 | 0.98434 | 0.95468 | 8.5055 | 0 | 27 | 8 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4441.4575 | No Control Sample | 0 | off | off | 4441.4575 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9548 | 5.5958 | 2.2269 | 77.9651 | -3.2577 | -0.25399 | 0 | 0 | 0 | 0 | 0 | 4.0029 | FALSE | 0.0625 | 0.059944
2733 | 2734 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FHpSA | 0.241854590347088 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.58 | 8.86 | Baseline | Baseline | 7584.1921 | 74093.2583 | 0.10236 | 1677.1698 | 16194.4595 | 0.10356 | 4.522 | 4.5752 | 1677.1698 | 16205.5983 | 1 | 0.99406 | 8.5836 | 9.0208 | 0.041783 | 0.003585 | 8.3139 | 8.8304 | 8.8405 | 9.1542 | 0.52655 | 0.32387 | 0.07106 | 0.070703 | 95.4828 | 178.3493 | 0 | 0.0010529 | FALSE | 0.95153 | TRUE | [M-H]- | 96.7418 | 8.5119 | 8.6592 | 0.14729 | 8.5191 | 8.6487 | 0.12953 | 0 | 1.0936 | 1.1553 | 3.9737 | 0 | 29 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 7584.1921 | No Control Sample | 0 | off | off | 7584.1921 | 0 | off | 0 | off | 0 | off | 0 | off | 447.952 | 5.8725 | 2.6306 | 77.9648 | -7.4749 | -0.58278 | 0 | 0 | 0 | 0 | 0 | 2.4452 | FALSE | 0.0625 | 0.10236
2735 | 2736 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | FOSA | 0.279485832882297 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.88 | 8.86 | Baseline | Baseline | 4711.7687 | 74093.2583 | 0.063592 | 1001.5623 | 16194.4595 | 0.061846 | 4.7044 | 4.5752 | 1002.2326 | 16205.5983 | 0.98102 | 0.99406 | 9.0266 | 9.0208 | 1.6508 | 0.14659 | 8.8833 | 8.8304 | 9.146 | 9.1542 | 0.26277 | 0.32387 | 0.074754 | 0.070703 | 124.5162 | 178.3493 | 0.0011861 | 0.0010529 | FALSE | 1.0006 | TRUE | [M-H]- | 111.7943 | 8.9473 | 9.0995 | 0.15214 | 8.9542 | 9.089 | 0.13482 | 4.521 | 0.91786 | 0.77352 | 6.668 | 0 | 14 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4711.7687 | No Control Sample | 0 | off | off | 4711.7687 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9479 | 3.4368 | 1.7114 | 77.9647 | -8.7396 | -0.68139 | 0 | 0 | 0 | 0 | 0 | 19.3718 | FALSE | 0.0625 | 0.063592
2754 | 2755 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 4:2 FTS | 0.232644248112027 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Total Width | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 5.94 | 5.92 | Manual | Baseline | 2336.6332 | 19781.0328 | 0.11812 | 404.3352 | 3458.0572 | 0.11693 | 5.779 | 5.7203 | 410.3557 | 3462.7032 | 0.95363 | 0.99362 | 5.9548 | 5.9557 | 0.24978 | 0.014837 | 5.8445 | 5.8115 | 6.0793 | 6.138 | 0.23484 | 0.32652 | 0.089131 | 0.089276 | 132.4572 | 179.8164 | 0.0102 | 0.0017933 | TRUE | 0.99986 | TRUE | [M-H]- | 93.0577 | 5.8566 | 6.0508 | 0.19424 | 5.8629 | 6.0445 | 0.18161 | 17.5618 | 1.0126 | 1.0268 | 5.011 | 0 | 29 | 13 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Fail | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2336.6332 | No Control Sample | 0 | off | off | 2336.6332 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9765 | 6.8775 | 2.2488 | 306.9682 | 1.8781 | 0.57652 | 0 | 0 | 0 | 0 | 0 | 20.03 | FALSE | 0.0625 | 0.11812
2756 | 2757 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 6:2 FTS | 0.261043269833642 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 7.92 | 7.92 | Baseline | Baseline | 2554.1077 | 16126.2798 | 0.15838 | 670.8902 | 4068.3014 | 0.16491 | 3.807 | 3.9639 | 670.8902 | 4071.1785 | 1 | 0.99562 | 7.9341 | 7.939 | 0.17802 | 0.014099 | 7.7666 | 7.8012 | 8.1255 | 8.0997 | 0.35885 | 0.29848 | 0.058696 | 0.059375 | 194.7718 | 119.9191 | 0 | 0.0014563 | FALSE | 0.99938 | TRUE | [M-H]- | 104.4173 | 7.8766 | 8.0064 | 0.12987 | 7.8829 | 7.9965 | 0.11359 | 0 | 1.29 | 1.581 | 4.7615 | 0 | 27 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2554.1077 | No Control Sample | 0 | off | off | 2554.1077 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9707 | 6.5585 | 2.8003 | 406.9626 | 2.7483 | 1.1184 | 0 | 0 | 0 | 0 | 0 | 1702.2346 | FALSE | 0.0625 | 0.15838
2758 | 2759 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 8:2 FTS | 0.240111872666709 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 8.88 | 8.89 | Baseline | Baseline | 2287.6971 | 19418.3038 | 0.11781 | 399.2621 | 3732.2063 | 0.10698 | 5.7298 | 5.2029 | 399.2621 | 3732.2076 | 1 | 0.9966 | 8.8969 | 8.89 | 0.18994 | 0.016867 | 8.6656 | 8.7111 | 9.1096 | 9.0796 | 0.44398 | 0.36852 | 0.091579 | 0.081924 | 161.4685 | 113.273 | 0 | 7.4074e-07 | FALSE | 1.0008 | TRUE | [M-H]- | 96.0447 | 8.805 | 8.9828 | 0.17783 | 8.8132 | 8.9704 | 0.15718 | 0 | 0.86632 | 0.66395 | 5.9478 | 0 | 23 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2287.6971 | No Control Sample | 0 | off | off | 2287.6971 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9643 | 5.1854 | 2.7325 | 506.9612 | 10.8981 | 5.5249 | 0 | 0 | 0 | 0 | 0 | 14.5886 | FALSE | 0.0625 | 0.11781
2760 | 2761 | PFAS CS4 0.25ng/ml | 20 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 21) | CS4 | 0 | 0 | Standard | 0 | 0 | 0 | 5 | 1 | 10:2 FTS | 0.192432216854493 | Quantifiers | 10/27/2023 00:23:00 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 0.25 | 4 | 0.0625 | 0 | 9.5 | 8.89 | Baseline | Baseline | 1908.6316 | 19418.3038 | 0.09829 | 467.0787 | 3732.2063 | 0.12515 | 4.0863 | 5.2029 | 467.0787 | 3732.2076 | 1 | 0.9966 | 9.5049 | 8.89 | 0.05158 | 0.0049001 | 9.3278 | 8.7111 | 9.6695 | 9.0796 | 0.34167 | 0.36852 | 0.063167 | 0.081924 | 1814.4745 | 113.273 | 0 | 7.4074e-07 | FALSE | 1.0692 | TRUE | [M-H]- | 76.9729 | 9.4368 | 9.5748 | 0.13797 | 9.447 | 9.5665 | 0.11948 | 0 | 1.1767 | 1.4648 | 3.826 | 0 | 23 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1908.6316 | No Control Sample | 0 | off | off | 1908.6316 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9578 | 4.2739 | 2.6796 | 606.9492 | 2.1796 | 1.3229 | 0 | 0 | 0 | 0 | 0 | 5.9874 | FALSE | 0.0625 | 0.09829
2821 | 2822 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFBA | 1.11810478578288 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 1.49 | 1.48 | Manual | Manual | 20943.6066 | 104425.7097 | 0.20056 | 3005.7696 | 15794.5769 | 0.1903 | 6.9678 | 6.6115 | 3027.4014 | 15836.7307 | 0.93876 | 0.98763 | 1.4402 | 1.4412 | 3.3414 | 0.049787 | 1.3263 | 1.3214 | 1.6605 | 1.6674 | 0.33424 | 0.34601 | 0.10754 | 0.098862 | 98.033 | 224.1393 | 0.0039426 | 0.0011394 | TRUE | 0.99931 | TRUE | [M-H]- | 111.8105 | 1.3509 | 1.5935 | 0.24256 | 1.3608 | 1.5743 | 0.21351 | -35.4554 | 1.3071 | 1.5751 | 1.6521 | 0 | 68 | 23 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 20943.6066 | No Control Sample | 0 | off | off | 20943.6066 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9806 | 6.4449 | 1.3726 | 168.9898 | 4.7605 | 0.80446 | 0 | 0 | 0 | 0 | 0 | 6.4112 | FALSE | 0.25 | 0.20056
2823 | 2824 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFPeA | 1.03985918393042 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 3.18 | 3.21 | Manual | Manual | 17757.6085 | 76359.2956 | 0.23255 | 1883.7762 | 8748.1337 | 0.21533 | 9.4266 | 8.7286 | 1890.9386 | 8777.4714 | 0.96574 | 0.98503 | 3.18 | 3.1744 | 0.00034212 | 1.0879e-05 | 2.9794 | 2.9534 | 3.4458 | 3.4543 | 0.46634 | 0.50088 | 0.15568 | 0.14373 | 238.0628 | 453.9466 | 0 | 0.00093745 | TRUE | 1.0018 | TRUE | [M-H]- | 103.9859 | 3.0231 | 3.352 | 0.32891 | 3.0481 | 3.3307 | 0.28258 | 0 | 1.0789 | 1.218 | 1.8957 | 0 | 84 | 30 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 17757.6085 | No Control Sample | 0 | off | off | 17757.6085 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9776 | 6.1557 | 1.6188 | 218.9866 | 3.8618 | 0.84567 | 0 | 0 | 0 | 0 | 0 | 2.5482 | FALSE | 0.25 | 0.23255
2825 | 2826 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFHxA | 1.04523762730879 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 18.0379 | 6.13 | 6.12 | Baseline | Baseline | 36769.123 | 151154.1922 | 0.24326 | 6688.8108 | 26228.6684 | 0.25502 | 5.4971 | 5.7629 | 6696.217 | 26243.712 | 0.98818 | 0.99542 | 6.1109 | 6.11 | 0.31178 | 0.019112 | 5.9596 | 5.9199 | 6.289 | 6.2925 | 0.32938 | 0.37262 | 0.084842 | 0.090601 | 223.5791 | 241.1017 | 0.00032467 | 0.0011045 | FALSE | 1.0002 | TRUE | [M-H]- | 104.5238 | 6.016 | 6.2057 | 0.18975 | 6.0293 | 6.1901 | 0.16075 | 6.593 | 0.95001 | 0.85709 | 1.5356 | 0 | 36 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 36769.123 | No Control Sample | 203843.5852 | off | off | 36769.123 | 203843.5852 | off | 1 | off | 1 | off | 0.18038 | off | 312.9746 | 5.7731 | 1.8068 | 268.9833 | 3.4857 | 0.9376 | 0 | 0 | 0 | 0 | 0 | 170.3009 | FALSE | 0.25 | 0.24326
2827 | 2828 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFHpA | 1.19060472652118 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 22.6998 | 7.26 | 7.27 | Baseline | Baseline | 51830.0781 | 173974.5489 | 0.29792 | 12185.0906 | 41628.9786 | 0.29271 | 4.2536 | 4.1792 | 12188.0433 | 41652.5673 | 0.99421 | 0.996 | 7.2607 | 7.2622 | 0.010326 | 0.00074969 | 7.1251 | 7.1373 | 7.4188 | 7.4559 | 0.29367 | 0.31858 | 0.065036 | 0.064776 | 243.524 | 293.862 | 0.00052193 | 0.00019936 | FALSE | 0.99979 | TRUE | [M-H]- | 119.0605 | 7.1905 | 7.3332 | 0.14272 | 7.2004 | 7.3239 | 0.12353 | 21.6565 | 1.0707 | 1.1753 | 1.4839 | 0 | 34 | 10 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 51830.0781 | No Control Sample | 228328.7703 | off | off | 51830.0781 | 228328.7703 | off | 1 | off | 1 | off | 0.227 | off | 362.9715 | 5.2664 | 1.9115 | 318.9807 | 4.9492 | 1.5787 | 0 | 0 | 0 | 0 | 0 | 2.8034 | FALSE | 0.25 | 0.29792
2829 | 2830 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFOA | 1.07765075763146 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 22.8531 | 7.96 | 7.94 | Baseline | Baseline | 43758.0047 | 167429.0273 | 0.26135 | 10283.3784 | 40120.8785 | 0.25631 | 4.2552 | 4.1731 | 10297.6124 | 40137.9415 | 0.98329 | 0.99669 | 7.9503 | 7.9515 | 0.12224 | 0.0097304 | 7.8076 | 7.7859 | 8.1906 | 8.1183 | 0.38302 | 0.33233 | 0.064953 | 0.063093 | 103.5026 | 110.5415 | 0.0015504 | 0.00078355 | FALSE | 0.99984 | TRUE | [M-H]- | 107.7651 | 7.8844 | 8.0282 | 0.14371 | 7.8934 | 8.0182 | 0.12482 | -41.6254 | 1.1226 | 1.2656 | 1.6916 | 0 | 28 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 43758.0047 | No Control Sample | 191475.4056 | off | off | 43758.0047 | 191475.4056 | off | 1 | off | 1 | off | 0.22853 | off | 412.9684 | 4.7625 | 1.9667 | 368.9779 | 4.9902 | 1.8413 | 0 | 0 | 0 | 0 | 0 | 0.67962 | FALSE | 0.25 | 0.26135
2831 | 2832 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFNA | 1.0153888597183 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 20.4775 | 8.46 | 8.47 | Baseline | Baseline | 32588.5937 | 138156.0506 | 0.23588 | 6448.0057 | 28655.2278 | 0.22502 | 5.0541 | 4.8213 | 6455.474 | 28670.2745 | 0.9924 | 0.99703 | 8.4626 | 8.4625 | 0.031092 | 0.0026304 | 8.319 | 8.3272 | 8.5989 | 8.81 | 0.27983 | 0.48273 | 0.080899 | 0.075726 | 187.1477 | 137.1449 | 0.0023521 | 0.00072694 | FALSE | 1 | TRUE | [M-H]- | 101.5389 | 8.3809 | 8.543 | 0.16211 | 8.3925 | 8.5323 | 0.13977 | 54.1978 | 1.1002 | 1.2529 | 1.4713 | 0 | 16 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 32588.5937 | No Control Sample | 159143.3885 | off | off | 32588.5937 | 159143.3885 | off | 1 | off | 1 | off | 0.20478 | off | 462.9652 | 4.3241 | 2.0019 | 418.9753 | 4.9936 | 2.0922 | 0 | 0 | 0 | 0 | 0 | 2.6326 | FALSE | 0.25 | 0.23588
2833 | 2834 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFDA | 1.0080526873872 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.87 | 8.86 | Baseline | Baseline | 31483.9891 | 125813.3659 | 0.25024 | 6098.9868 | 25646.3448 | 0.23781 | 5.1622 | 4.9057 | 6101.3884 | 25647.3969 | 0.99321 | 0.9992 | 8.8724 | 8.873 | 0.026833 | 0.0023801 | 8.741 | 8.6923 | 9.0534 | 9.1191 | 0.31247 | 0.42683 | 0.081446 | 0.076734 | 127.9704 | 134.5688 | 0.00062436 | 8.7448e-05 | FALSE | 0.99993 | TRUE | [M-H]- | 100.8053 | 8.7902 | 8.9608 | 0.17061 | 8.8027 | 8.9464 | 0.1437 | -12.1867 | 1.1397 | 1.3064 | 1.7961 | 0 | 16 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 31483.9891 | No Control Sample | 0 | off | off | 31483.9891 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9622 | 4.3017 | 2.2066 | 468.9732 | 6.7764 | 3.1779 | 0 | 0 | 0 | 0 | 0 | 87.1728 | FALSE | 0.25 | 0.25024
2835 | 2836 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFUdA | 1.15148060642155 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.21 | 9.2 | Baseline | Baseline | 27186.1091 | 106461.9013 | 0.25536 | 6206.779 | 23280.3167 | 0.26661 | 4.3801 | 4.573 | 6206.7814 | 23288.6898 | 0.99808 | 0.99634 | 9.2086 | 9.2081 | 0.015211 | 0.0014009 | 8.9953 | 9.0625 | 9.4636 | 9.3869 | 0.46832 | 0.32435 | 0.067785 | 0.070902 | 186.5401 | 165.76 | 8.5773e-07 | 0.00072944 | FALSE | 1.0001 | TRUE | [M-H]- | 115.1481 | 9.1344 | 9.2837 | 0.14934 | 9.1427 | 9.2716 | 0.12895 | 0.011368 | 0.98458 | 0.9086 | 1.8894 | 0 | 26 | 5 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 27186.1091 | No Control Sample | 0 | off | off | 27186.1091 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9592 | 4.2401 | 2.387 | 518.9684 | 2.1561 | 1.1189 | 0 | 0 | 0 | 0 | 0 | 2.1856 | FALSE | 0.25 | 0.25536
2837 | 2838 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFDoA | 1.09436351713759 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.47 | 9.47 | Baseline | Baseline | 25228.7935 | 89085.5844 | 0.2832 | 6222.9566 | 22769.863 | 0.2733 | 4.0541 | 3.9124 | 6224.5823 | 22778.672 | 0.9946 | 0.99522 | 9.4713 | 9.4721 | 0.014101 | 0.0013354 | 9.3468 | 9.3177 | 9.5984 | 9.6324 | 0.25163 | 0.3147 | 0.063982 | 0.061247 | 280.9421 | 348.1031 | 0.00052489 | 0.00079066 | FALSE | 0.99991 | TRUE | [M-H]- | 109.4364 | 9.4088 | 9.5426 | 0.13385 | 9.4171 | 9.5336 | 0.11648 | 12.9806 | 1.2018 | 1.46 | 1.3231 | 0 | 17 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 25228.7935 | No Control Sample | 0 | off | off | 25228.7935 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9569 | 5.2969 | 3.2467 | 568.9663 | 5.1603 | 2.936 | 0 | 0 | 0 | 0 | 0 | 385.2298 | FALSE | 0.25 | 0.2832
2839 | 2840 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFTrDA | 0.984357323237237 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.68 | 9.47 | Baseline | Baseline | 16541.4247 | 89085.5844 | 0.18568 | 4741.1707 | 22769.863 | 0.20822 | 3.4889 | 3.9124 | 4744.1208 | 22778.672 | 0.99119 | 0.99522 | 9.6895 | 9.4721 | 0.098419 | 0.0095269 | 9.5846 | 9.3177 | 9.7836 | 9.6324 | 0.19907 | 0.3147 | 0.053397 | 0.061247 | 500.3152 | 348.1031 | 0.0011107 | 0.00079066 | FALSE | 1.0229 | TRUE | [M-H]- | 98.4357 | 9.6306 | 9.7499 | 0.11931 | 9.6375 | 9.7406 | 0.10309 | 26.4526 | 0.93481 | 0.80977 | 1.2713 | 0 | 17 | 6 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 16541.4247 | No Control Sample | 0 | off | off | 16541.4247 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9544 | 5.8973 | 3.9096 | 618.9643 | 4.2282 | 2.6171 | 0 | 0 | 0 | 0 | 0 | 2.5569 | FALSE | 0.25 | 0.18568
2841 | 2842 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFTeDA | 1.05610136069981 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.87 | 9.86 | Baseline | Baseline | 8679.537 | 23180.2089 | 0.37444 | 2584.8389 | 6827.4375 | 0.3786 | 3.3579 | 3.3952 | 2588.0008 | 6827.9582 | 0.96713 | 0.99766 | 9.8693 | 9.8691 | 0.0073416 | 0.00072461 | 9.7845 | 9.7665 | 9.9612 | 9.9954 | 0.17672 | 0.22888 | 0.052331 | 0.053259 | 512.6825 | 999.7728 | 0.0025093 | 0.00017153 | FALSE | 1 | TRUE | [M-H]- | 105.6101 | 9.8166 | 9.9302 | 0.11363 | 9.8233 | 9.9208 | 0.097422 | 36.7043 | 1.1712 | 1.334 | 1.6158 | 0 | 19 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8679.537 | No Control Sample | 0 | off | off | 8679.537 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9494 | 3.0226 | 2.1549 | 668.9597 | 2.6254 | 1.7563 | 0 | 0 | 0 | 0 | 0 | 352.7953 | FALSE | 0.25 | 0.37444
2848 | 2849 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFBS | 1.07026372481916 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 3.5 | 3.5 | Manual | Manual | 12171.3601 | 42964.9552 | 0.28329 | 1264.2427 | 3901.0069 | 0.32408 | 9.6274 | 11.0138 | 1267.2963 | 3910.7836 | 0.98561 | 0.98862 | 3.8141 | 3.8148 | 8.9746 | 0.31411 | 3.5705 | 3.552 | 4.0602 | 4.1212 | 0.48972 | 0.56919 | 0.13838 | 0.17215 | 255.307 | 909.2773 | 0.0045563 | 0 | TRUE | 0.99983 | TRUE | [M-H]- | 107.0264 | 3.6465 | 4.0212 | 0.37469 | 3.6684 | 3.978 | 0.30952 | -11.7624 | 1.2163 | 1.3426 | 11.1712 | 0 | 77 | 23 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 12171.3601 | No Control Sample | 0 | off | off | 12171.3601 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9448 | 5.8892 | 1.7605 | 79.9565 | -9.0183 | -0.72108 | 0 | 0 | 0 | 0 | 0 | 5.2627 | FALSE | 0.25 | 0.28329
2850 | 2851 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFPeS | 1.12054332873824 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 6.38 | 3.5 | Baseline | Manual | 12830.9011 | 42964.9552 | 0.29864 | 2466.3819 | 3901.0069 | 0.63224 | 5.2023 | 11.0138 | 2466.424 | 3910.7836 | 0.99355 | 0.98862 | 6.362 | 3.8148 | 0.2817 | 0.017973 | 6.2118 | 3.552 | 6.5262 | 4.1212 | 0.31437 | 0.56919 | 0.079747 | 0.17215 | 252.8114 | 909.2773 | 3.7035e-05 | 0 | FALSE | 1.6677 | TRUE | [M-H]- | 112.0543 | 6.2812 | 6.4582 | 0.17705 | 6.2922 | 6.446 | 0.15376 | 0.29056 | 1.1732 | 1.3869 | 14.1465 | 0 | 34 | 11 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 12830.9011 | No Control Sample | 0 | off | off | 12830.9011 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9415 | 4.8332 | 1.6865 | 79.9569 | -3.1675 | -0.25326 | 0 | 0 | 0 | 0 | 0 | 0.37777 | FALSE | 0.25 | 0.29864
2852 | 2853 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | L-PFHxS | 0.825028127389708 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.811 | 4 | 0.20275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 9352.5668 | 60687.3918 | 0.15411 | 2247.1411 | 15039.4367 | 0.14942 | 4.162 | 4.0352 | 2288.0702 | 15055.7754 | 0.73599 | 0.98928 | 7.3372 | 7.3376 | 0.038172 | 0.0028018 | 7.2589 | 7.1951 | 7.4378 | 7.4739 | 0.1789 | 0.27885 | 0.065442 | 0.06252 | 86.7542 | 176.0772 | 0.016582 | 0.0016937 | FALSE | 0.99994 | TRUE | [M-H]- | 101.7297 | 7.2746 | 7.4046 | 0.13 | 7.2813 | 7.3978 | 0.11646 | -208.2877 | 1.0639 | 1.142 | 14.6926 | 0 | 22 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9352.5668 | No Control Sample | 0 | off | off | 9352.5668 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9387 | 5.1847 | 2.0684 | 79.9569 | -3.8879 | -0.31086 | 0 | 0 | 0 | 0 | 0 | 33.1489 | FALSE | 0.20275 | 0.15411
2854 | 2855 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | Br-PFHxS | 0.16918511356814 | Qualifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.189 | 4 | 0.04725 | 0 | 7.16 | 7.35 | Manual | Baseline | 2605.7216 | 60687.3918 | 0.042937 | 505.8921 | 15039.4367 | 0.033638 | 5.1507 | 4.0352 | 2288.0702 | 15055.7754 | 0.20535 | 0.98928 | 7.1547 | 7.3376 | 0.073421 | 0.005257 | 7.062 | 7.1951 | 7.2604 | 7.4739 | 0.19844 | 0.27885 | 0.083043 | 0.06252 | 19.5307 | 176.0772 | 0.0091368 | 0.0016937 | TRUE | 0.97507 | TRUE | [M-H]- | 89.5159 | 7.0772 | 7.2676 | 0.19038 | 7.0842 | 7.255 | 0.17078 | -23.2928 | 1.085 | 1.1153 | 0.27861 | 0 | 23 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2605.7216 | No Control Sample | 0 | off | off | 9352.5668 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9387 | 5.1847 | 2.0684 | 79.9603 | 39.1108 | 3.1272 | 0 | 0 | 0 | 0 | 0 | 33.1489 | FALSE | 0.04725 | 0.042937
2856 | 2857 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFHpS | 1.15318053079286 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy, Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 7.98 | 7.35 | Baseline | Baseline | 10651.0631 | 60687.3918 | 0.17551 | 2476.8966 | 15039.4367 | 0.16469 | 4.3002 | 4.0352 | 2480.4309 | 15055.7754 | 0.99075 | 0.98928 | 7.9734 | 7.3376 | 0.082894 | 0.0066149 | 7.8629 | 7.1951 | 8.0985 | 7.4739 | 0.2356 | 0.27885 | 0.066931 | 0.06252 | 122.2895 | 176.0772 | 0.0024188 | 0.0016937 | FALSE | 1.0866 | TRUE | [M-H]- | 115.3181 | 7.9065 | 8.0526 | 0.14609 | 7.9164 | 8.044 | 0.12757 | -25.4296 | 1.0715 | 1.1923 | 19.5025 | 0 | 18 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10651.0631 | No Control Sample | 0 | off | off | 10651.0631 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9354 | 4.4453 | 1.9956 | 79.957 | -2.7357 | -0.21874 | 0 | 0 | 0 | 0 | 0 | 0.47665 | FALSE | 0.25 | 0.17551
2858 | 2859 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | L-PFOS | 0.783209670604349 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 0.788 | 4 | 0.197 | 0 | 8.46 | 8.45 | Manual | Baseline | 8096.0499 | 39768.555 | 0.20358 | 1780.7242 | 8277.5646 | 0.21513 | 4.5465 | 4.8044 | 1798.8026 | 8277.606 | 0.66976 | 0.99656 | 8.4589 | 8.4614 | 0.012545 | 0.0010613 | 8.3818 | 8.2965 | 8.5498 | 8.6762 | 0.16794 | 0.37972 | 0.072498 | 0.075865 | 54.9558 | 117.2321 | 0 | 1.1017e-05 | TRUE | 0.99971 | TRUE | [M-H]- | 99.3921 | 8.3787 | 8.5319 | 0.15315 | 8.3787 | 8.5244 | 0.14568 | 0 | 0.955 | 0.8168 | 14.9343 | 0 | 10 | 5 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8096.0499 | No Control Sample | 0 | off | off | 8096.0499 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9323 | 4.2137 | 2.1023 | 79.9567 | -7.0159 | -0.56097 | 0 | 0 | 0 | 0 | 0 | 0.95982 | FALSE | 0.197 | 0.20358
2860 | 2861 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | Br-PFOS | 0.153232899170131 | Qualifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.211 | 4 | 0.05275 | 0 | 8.3 | 8.45 | Manual | Baseline | 3391.011 | 39768.555 | 0.085269 | 592.3613 | 8277.5646 | 0.071562 | 5.7246 | 4.8044 | 1798.8026 | 8277.606 | 0.28053 | 0.99656 | 8.3105 | 8.4614 | 0.12676 | 0.010521 | 8.1952 | 8.2965 | 8.3818 | 8.6762 | 0.1866 | 0.37972 | 0.095206 | 0.075865 | 18.2811 | 117.2321 | 0.0041408 | 1.1017e-05 | TRUE | 0.98217 | TRUE | [M-H]- | 72.6222 | 8.2234 | 8.3974 | 0.17403 | 8.2322 | 8.3974 | 0.16519 | 13.145 | 1.0471 | 1.2246 | 0.41885 | 0 | 13 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3391.011 | No Control Sample | 0 | off | off | 8096.0499 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9323 | 4.2137 | 2.1023 | 79.9572 | -1.8269 | -0.14607 | 0 | 0 | 0 | 0 | 0 | 0.95982 | FALSE | 0.05275 | 0.085269
2862 | 2863 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFNS | 1.08022606575503 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.85 | 8.45 | Baseline | Baseline | 9521.7022 | 39768.555 | 0.23943 | 1920.302 | 8277.5646 | 0.23199 | 4.9584 | 4.8044 | 1921.0452 | 8277.606 | 0.98697 | 0.99656 | 8.8583 | 8.4614 | 0.093581 | 0.008282 | 8.7071 | 8.2965 | 9.02 | 8.6762 | 0.31297 | 0.37972 | 0.077132 | 0.075865 | 106.2709 | 117.2321 | 0.00069788 | 1.1017e-05 | FALSE | 1.0469 | TRUE | [M-H]- | 108.0226 | 8.778 | 8.944 | 0.16604 | 8.7895 | 8.9307 | 0.14115 | -4.282 | 1.1782 | 1.3961 | 19.7448 | 0 | 16 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9521.7022 | No Control Sample | 0 | off | off | 9521.7022 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9287 | 3.1139 | 1.7093 | 79.9569 | -3.0645 | -0.24503 | 0 | 0 | 0 | 0 | 0 | 1.5708 | FALSE | 0.25 | 0.23943
2864 | 2865 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFDS | 1.1682934897151 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.19 | 8.45 | Baseline | Baseline | 10580.7389 | 39768.555 | 0.26606 | 2298.3137 | 8277.5646 | 0.27766 | 4.6037 | 4.8044 | 2298.3216 | 8277.606 | 0.99328 | 0.99656 | 9.1859 | 8.4614 | 0.04425 | 0.0040666 | 9.0587 | 8.2965 | 9.3675 | 8.6762 | 0.30883 | 0.37972 | 0.071729 | 0.075865 | 185.1058 | 117.2321 | 1.9705e-06 | 1.1017e-05 | FALSE | 1.0856 | TRUE | [M-H]- | 116.8293 | 9.1173 | 9.267 | 0.14962 | 9.1231 | 9.2537 | 0.13065 | -0.014665 | 0.96901 | 0.82801 | 16.757 | 0 | 17 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10580.7389 | No Control Sample | 0 | off | off | 10580.7389 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9255 | 2.8141 | 1.6854 | 79.9567 | -2.7656 | -0.22113 | 0 | 0 | 0 | 0 | 0 | 1.365 | FALSE | 0.25 | 0.26606
2868 | 2869 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FBSA | 0.860091434677942 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 4.6 | 8.86 | Manual | Baseline | 15646.9013 | 68993.9029 | 0.22679 | 1556.0333 | 14571.1676 | 0.10679 | 10.0556 | 4.735 | 1571.8306 | 14584.9804 | 0.95093 | 0.99323 | 4.938 | 9.0032 | 7.3476 | 0.33799 | 4.7457 | 8.8147 | 5.1415 | 9.1811 | 0.39579 | 0.36635 | 0.16184 | 0.073548 | 126.4236 | 116.4723 | 0 | 0.0017543 | TRUE | 0.54847 | TRUE | [M-H]- | 86.0091 | 4.7515 | 5.1322 | 0.38077 | 4.7795 | 5.1187 | 0.33921 | 0 | 0.972 | 1.0212 | 5.8935 | 0 | 72 | 30 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 15646.9013 | No Control Sample | 0 | off | off | 15646.9013 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9607 | 5.9054 | 1.7596 | 77.9652 | -1.888 | -0.1472 | 0 | 0 | 0 | 0 | 0 | 3.5651 | FALSE | 0.25 | 0.22679
2870 | 2871 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FPeSA | 1.16097330338916 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 7 | 8.86 | Baseline | Baseline | 32798.372 | 68993.9029 | 0.47538 | 7076.8707 | 14571.1676 | 0.48568 | 4.6346 | 4.735 | 7081.5667 | 14584.9804 | 0.99185 | 0.99323 | 7.0852 | 9.0032 | 1.2172 | 0.085202 | 6.9304 | 8.8147 | 7.2494 | 9.1811 | 0.31907 | 0.36635 | 0.070467 | 0.073548 | 227.06 | 116.4723 | 0.0013151 | 0.0017543 | FALSE | 0.78697 | TRUE | [M-H]- | 116.0973 | 7.0066 | 7.1638 | 0.15718 | 7.0194 | 7.1532 | 0.13382 | 29.169 | 0.96401 | 0.94671 | 3.6988 | 0 | 34 | 9 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 32798.372 | No Control Sample | 0 | off | off | 32798.372 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9575 | 4.8701 | 1.6946 | 77.9652 | -1.5195 | -0.11847 | 0 | 0 | 0 | 0 | 0 | 0.68263 | FALSE | 0.25 | 0.47538
2872 | 2873 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FHxSA | 1.21680290557326 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 7.86 | 8.86 | Baseline | Baseline | 17154.1688 | 68993.9029 | 0.24863 | 3837.5437 | 14571.1676 | 0.26337 | 4.4701 | 4.735 | 3838.4206 | 14584.9804 | 0.99319 | 0.99323 | 7.9765 | 9.0032 | 1.4821 | 0.11649 | 7.8475 | 8.8147 | 8.1604 | 9.1811 | 0.31295 | 0.36635 | 0.068284 | 0.073548 | 92.6604 | 116.4723 | 0.00056133 | 0.0017543 | FALSE | 0.88596 | TRUE | [M-H]- | 121.6803 | 7.908 | 8.0606 | 0.15255 | 7.9171 | 8.0468 | 0.12972 | 6.8833 | 1.1759 | 1.3258 | 9.3083 | 0 | 23 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 17154.1688 | No Control Sample | 0 | off | off | 17154.1688 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9547 | 5.2469 | 2.088 | 77.9651 | -3.539 | -0.27592 | 0 | 0 | 0 | 0 | 0 | 0.70303 | FALSE | 0.25 | 0.24863
2874 | 2875 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FHpSA | 1.18191362679386 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.58 | 8.86 | Baseline | Baseline | 34512.2083 | 68993.9029 | 0.50022 | 7541.8836 | 14571.1676 | 0.51759 | 4.5761 | 4.735 | 7548.4589 | 14584.9804 | 0.99124 | 0.99323 | 8.5579 | 9.0032 | 0.25784 | 0.022123 | 8.3395 | 8.8147 | 8.7044 | 9.1811 | 0.36495 | 0.36635 | 0.070426 | 0.073548 | 126.5359 | 116.4723 | 0.001384 | 0.0017543 | FALSE | 0.95054 | TRUE | [M-H]- | 118.1914 | 8.4839 | 8.6345 | 0.15058 | 8.4965 | 8.6253 | 0.12874 | 28.6019 | 0.91611 | 0.85037 | 3.9232 | 0 | 21 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 34512.2083 | No Control Sample | 0 | off | off | 34512.2083 | 0 | off | 0 | off | 0 | off | 0 | off | 447.951 | 3.6609 | 1.6399 | 77.965 | -5.1106 | -0.39845 | 0 | 0 | 0 | 0 | 0 | 1.0378 | FALSE | 0.25 | 0.50022
2876 | 2877 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FOSA | 1.08692144685352 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.88 | 8.86 | Baseline | Baseline | 17062.9588 | 68993.9029 | 0.24731 | 3549.757 | 14571.1676 | 0.24362 | 4.8068 | 4.735 | 3549.757 | 14584.9804 | 0.99305 | 0.99323 | 8.9995 | 9.0032 | 1.3461 | 0.11953 | 8.7436 | 8.8147 | 9.293 | 9.1811 | 0.5494 | 0.36635 | 0.075077 | 0.073548 | 152.6616 | 116.4723 | 0 | 0.0017543 | FALSE | 0.99959 | TRUE | [M-H]- | 108.6921 | 8.9178 | 9.0787 | 0.16091 | 8.9306 | 9.0691 | 0.13849 | 0 | 1.0047 | 1.0577 | 7.8352 | 0 | 28 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 17062.9588 | No Control Sample | 0 | off | off | 17062.9588 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9485 | 4.5443 | 2.2628 | 77.965 | -5.1152 | -0.39881 | 0 | 0 | 0 | 0 | 0 | 1.2917 | FALSE | 0.25 | 0.24731
2895 | 2896 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 4:2 FTS | 1.02650866882859 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 5.94 | 5.92 | Baseline | Baseline | 9453.0682 | 18136.81 | 0.52121 | 1660.9217 | 3144.5809 | 0.52819 | 5.6915 | 5.7676 | 1685.6681 | 3184.972 | 0.95622 | 0.95905 | 5.9228 | 5.9167 | 0.28909 | 0.017172 | 5.8091 | 5.8033 | 6.0807 | 6.0833 | 0.27155 | 0.28 | 0.090031 | 0.092042 | 209.0763 | 177.863 | 0.023837 | 0.017738 | FALSE | 1.001 | TRUE | [M-H]- | 102.6509 | 5.8236 | 6.0149 | 0.19139 | 5.8315 | 6.0037 | 0.17213 | -145.7963 | 1.0026 | 0.96782 | 5.0493 | 0 | 32 | 12 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9453.0682 | No Control Sample | 0 | off | off | 9453.0682 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9763 | 6.1411 | 2.008 | 306.9689 | 4.261 | 1.308 | 0 | 0 | 0 | 0 | 0 | 8.6834 | FALSE | 0.25 | 0.52121
2897 | 2898 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 6:2 FTS | 1.0121209239373 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 7.92 | 7.92 | Baseline | Baseline | 9946.1265 | 16196.7941 | 0.61408 | 2339.0679 | 3894.0877 | 0.60067 | 4.2522 | 4.1593 | 2339.0679 | 3896.4717 | 1 | 0.99548 | 7.9127 | 7.9125 | 0.092297 | 0.00731 | 7.7103 | 7.774 | 8.148 | 8.0249 | 0.43768 | 0.25082 | 0.064999 | 0.063517 | 85.019 | 94.1855 | 0 | 0.0010984 | FALSE | 1 | TRUE | [M-H]- | 101.2121 | 7.8427 | 7.9867 | 0.144 | 7.8523 | 7.9747 | 0.12235 | 0 | 1.1474 | 1.3011 | 5.133 | 0 | 32 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9946.1265 | No Control Sample | 0 | off | off | 9946.1265 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9697 | 4.3084 | 1.8396 | 406.9638 | 5.7679 | 2.3473 | 0 | 0 | 0 | 0 | 0 | 6.0898 | FALSE | 0.25 | 0.61408
2899 | 2900 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 8:2 FTS | 1.0539712062972 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.88 | 8.89 | Baseline | Baseline | 9975.8927 | 19290.7638 | 0.51713 | 1993.8086 | 3765.1598 | 0.52954 | 5.0034 | 5.1235 | 1993.8122 | 3765.1598 | 0.99764 | 1 | 8.8694 | 8.8694 | 0.11883 | 0.010552 | 8.708 | 8.6269 | 9.0786 | 9.1261 | 0.3707 | 0.4992 | 0.078022 | 0.079483 | 130.1793 | 151.1931 | 3.9549e-06 | 0 | FALSE | 1 | TRUE | [M-H]- | 105.3971 | 8.7893 | 8.954 | 0.16469 | 8.7971 | 8.9445 | 0.14743 | 0.021272 | 1.0072 | 0.99148 | 5.0044 | 0 | 19 | 5 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9975.8927 | No Control Sample | 0 | off | off | 9975.8927 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9639 | 4.5875 | 2.4174 | 506.9575 | 3.6402 | 1.8454 | 0 | 0 | 0 | 0 | 0 | 5.7212 | FALSE | 0.25 | 0.51713
2901 | 2902 | PFAS CS5 1ng/ml | 21 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 22) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 10:2 FTS | 0.908661176637003 | Quantifiers | 10/27/2023 00:39:39 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.5 | 8.89 | Baseline | Baseline | 8953.3268 | 19290.7638 | 0.46413 | 2148.408 | 3765.1598 | 0.5706 | 4.1674 | 5.1235 | 2162.8813 | 3765.1598 | 0.974 | 1 | 9.4905 | 8.8694 | 0.099612 | 0.0094631 | 9.3688 | 8.6269 | 9.5721 | 9.1261 | 0.20332 | 0.4992 | 0.065864 | 0.079483 | 135.1575 | 151.1931 | 0.011606 | 0 | FALSE | 1.07 | TRUE | [M-H]- | 90.8661 | 9.4206 | 9.5539 | 0.13323 | 9.4281 | 9.546 | 0.11791 | 122.6431 | 1.0385 | 1.0811 | 4.1047 | 0 | 14 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Fail | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8953.3268 | No Control Sample | 0 | off | off | 8953.3268 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9583 | 5.003 | 3.1367 | 606.949 | 1.8805 | 1.1414 | 0 | 0 | 0 | 0 | 0 | 2.9566 | FALSE | 0.25 | 0.46413
2962 | 2963 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFBA | 1.1186161155684 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 1.49 | 1.48 | Manual | Baseline | 21221.8608 | 105764.7288 | 0.20065 | 2899.0593 | 15245.7651 | 0.19016 | 7.3203 | 6.9373 | 2911.6992 | 15314.0207 | 0.9533 | 0.97894 | 1.4705 | 1.4674 | 1.3118 | 0.019546 | 1.3423 | 1.3254 | 1.6999 | 1.67 | 0.35751 | 0.34457 | 0.10373 | 0.098575 | 116.092 | 187.1907 | 0.0029687 | 0.010458 | TRUE | 1.0021 | TRUE | [M-H]- | 111.8616 | 1.3808 | 1.6204 | 0.23959 | 1.3894 | 1.6057 | 0.21625 | -24.0735 | 1.2564 | 1.4924 | 1.6193 | 0 | 73 | 22 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 21221.8608 | No Control Sample | 0 | off | off | 21221.8608 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9805 | 5.8744 | 1.2511 | 168.9895 | 2.9165 | 0.49286 | 0 | 0 | 0 | 0 | 0 | 7.2169 | FALSE | 0.25 | 0.20065
2964 | 2965 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFPeA | 1.03046150238385 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 3.18 | 3.21 | Manual | Manual | 17755.3865 | 77046.0412 | 0.23045 | 1879.6038 | 8379.9287 | 0.2243 | 9.4463 | 9.1941 | 1887.8755 | 8400.1326 | 0.96274 | 0.98843 | 3.2017 | 3.1954 | 0.68276 | 0.021712 | 2.9985 | 2.9838 | 3.4466 | 3.4816 | 0.44807 | 0.49775 | 0.13652 | 0.14615 | 254.4921 | 624.0205 | 0.0012682 | 0.0016706 | TRUE | 1.002 | TRUE | [M-H]- | 103.0462 | 3.0438 | 3.3819 | 0.33816 | 3.0644 | 3.3578 | 0.2934 | 5.3197 | 1.0218 | 1.0257 | 1.941 | 0 | 80 | 26 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 17755.3865 | No Control Sample | 0 | off | off | 17755.3865 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9776 | 5.9877 | 1.5746 | 218.9867 | 3.9612 | 0.86746 | 0 | 0 | 0 | 0 | 0 | 3.0529 | FALSE | 0.25 | 0.23045
2966 | 2967 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFHxA | 1.07978883707149 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 18.1401 | 6.13 | 6.12 | Baseline | Baseline | 36977.4038 | 147146.3666 | 0.2513 | 6379.0203 | 27093.886 | 0.23544 | 5.7967 | 5.431 | 6391.8496 | 27107.5896 | 0.98405 | 0.99498 | 6.1397 | 6.1371 | 0.15785 | 0.0096763 | 5.968 | 5.9706 | 6.3073 | 6.3758 | 0.33935 | 0.40518 | 0.089348 | 0.083424 | 209.8405 | 252.5804 | 0.0032196 | 0.0011804 | FALSE | 1.0004 | TRUE | [M-H]- | 107.9789 | 6.0435 | 6.2297 | 0.18625 | 6.0578 | 6.2187 | 0.16091 | 60.5206 | 0.93876 | 0.89543 | 1.5022 | 0 | 37 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 36977.4038 | No Control Sample | 203843.5852 | off | off | 36977.4038 | 203843.5852 | off | 1 | off | 1 | off | 0.1814 | off | 312.9746 | 5.6035 | 1.7538 | 268.9831 | 2.4574 | 0.661 | 0 | 0 | 0 | 0 | 0 | 176.1361 | FALSE | 0.25 | 0.2513
2968 | 2969 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFHpA | 1.25172988835081 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 22.769 | 7.26 | 7.27 | Baseline | Baseline | 51988.2155 | 165983.8166 | 0.31321 | 12400.2488 | 38732.7217 | 0.32015 | 4.1925 | 4.2854 | 12409.7607 | 38761.5134 | 0.99148 | 0.99531 | 7.2961 | 7.2969 | 0.49782 | 0.036142 | 7.1521 | 7.1357 | 7.4607 | 7.4461 | 0.30858 | 0.31035 | 0.06545 | 0.066578 | 254.7374 | 207.1364 | 0.0016856 | 0.0013388 | FALSE | 0.99989 | TRUE | [M-H]- | 125.173 | 7.228 | 7.3668 | 0.13876 | 7.2381 | 7.358 | 0.11983 | 67.7365 | 1.1225 | 1.3165 | 1.5372 | 0 | 36 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 51988.2155 | No Control Sample | 228328.7703 | off | off | 51988.2155 | 228328.7703 | off | 1 | off | 1 | off | 0.22769 | off | 362.9715 | 5.2995 | 1.9235 | 318.9806 | 4.8289 | 1.5403 | 0 | 0 | 0 | 0 | 0 | 0.71564 | FALSE | 0.25 | 0.31321
2970 | 2971 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFOA | 1.08742796807881 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 22.7354 | 7.96 | 7.94 | Baseline | Baseline | 43532.7838 | 165069.6474 | 0.26372 | 10059.8139 | 39625.0208 | 0.25388 | 4.3274 | 4.1658 | 10084.9388 | 39630.0739 | 0.97556 | 0.99849 | 7.9876 | 7.9863 | 0.34735 | 0.027649 | 7.8592 | 7.8384 | 8.1417 | 8.1808 | 0.28248 | 0.34238 | 0.066672 | 0.064094 | 96.7975 | 150.2247 | 0.0021118 | 0.00028385 | FALSE | 1.0002 | TRUE | [M-H]- | 108.7428 | 7.9221 | 8.0691 | 0.14704 | 7.9301 | 8.0582 | 0.12808 | -75.204 | 1.2105 | 1.4315 | 1.7211 | 0 | 21 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 43532.7838 | No Control Sample | 191475.4056 | off | off | 43532.7838 | 191475.4056 | off | 1 | off | 1 | off | 0.22735 | off | 412.9687 | 5.4508 | 2.251 | 368.9778 | 4.6165 | 1.7034 | 0 | 0 | 0 | 0 | 0 | 0.62912 | FALSE | 0.25 | 0.26372
2972 | 2973 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFNA | 1.1060928954457 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 23.6498 | 8.46 | 8.47 | Baseline | Baseline | 37637.0572 | 146474.067 | 0.25695 | 7784.5441 | 30761.6917 | 0.25306 | 4.8348 | 4.7616 | 7793.1482 | 30762.2762 | 0.99352 | 0.99921 | 8.5038 | 8.5016 | 0.5172 | 0.043755 | 8.3544 | 8.3101 | 8.6302 | 8.7274 | 0.27587 | 0.41737 | 0.075138 | 0.074841 | 219.0296 | 147.6343 | 0.0019128 | 4.0955e-05 | FALSE | 1.0003 | TRUE | [M-H]- | 110.6093 | 8.421 | 8.5805 | 0.15952 | 8.4333 | 8.5726 | 0.13931 | 53.9774 | 0.87835 | 0.77608 | 1.2772 | 0 | 16 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 37637.0572 | No Control Sample | 159143.3885 | off | off | 37637.0572 | 159143.3885 | off | 1 | off | 1 | off | 0.2365 | off | 462.9661 | 6.1008 | 2.8245 | 418.9752 | 4.7072 | 1.9722 | 0 | 0 | 0 | 0 | 0 | 0.91517 | FALSE | 0.25 | 0.25695
2974 | 2975 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFDA | 1.0742038908895 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.87 | 8.86 | Baseline | Baseline | 33022.0048 | 123833.1656 | 0.26667 | 6547.5015 | 25768.8249 | 0.25409 | 5.0435 | 4.8055 | 6551.0767 | 25768.8979 | 0.99459 | 0.99968 | 8.9165 | 8.9143 | 0.52397 | 0.046476 | 8.7205 | 8.7101 | 9.0713 | 9.2163 | 0.35082 | 0.5062 | 0.078499 | 0.074581 | 173.4569 | 138.6244 | 0.00090707 | 6.9817e-06 | FALSE | 1.0002 | TRUE | [M-H]- | 107.4204 | 8.834 | 9.0015 | 0.16741 | 8.8456 | 8.9896 | 0.14402 | 16.9292 | 0.9062 | 0.78128 | 1.7274 | 0 | 18 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 33022.0048 | No Control Sample | 0 | off | off | 33022.0048 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9627 | 5.1855 | 2.6599 | 468.9715 | 3.199 | 1.5002 | 0 | 0 | 0 | 0 | 0 | 93.3182 | FALSE | 0.25 | 0.26667
2976 | 2977 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFUdA | 1.10275403130484 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.21 | 9.2 | Baseline | Baseline | 25426.4856 | 103970.8211 | 0.24455 | 5569.7277 | 22543.8964 | 0.24706 | 4.5651 | 4.6119 | 5570.4065 | 22552.013 | 0.9966 | 0.99687 | 9.2498 | 9.2519 | 0.43226 | 0.039811 | 9.0912 | 9.1186 | 9.3967 | 9.4023 | 0.30552 | 0.28373 | 0.071311 | 0.072766 | 228.0274 | 116.3788 | 0.00023195 | 0.00066802 | FALSE | 0.99977 | TRUE | [M-H]- | 110.2754 | 9.1746 | 9.3274 | 0.15281 | 9.1852 | 9.3201 | 0.13482 | 4.2286 | 1.0256 | 1.112 | 1.9452 | 0 | 17 | 5 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 25426.4856 | No Control Sample | 0 | off | off | 25426.4856 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9599 | 5.391 | 3.0349 | 518.9689 | 3.0916 | 1.6044 | 0 | 0 | 0 | 0 | 0 | 1.6275 | FALSE | 0.25 | 0.24455
2978 | 2979 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFDoA | 0.905023351423665 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.47 | 9.47 | Baseline | Baseline | 19963.3012 | 85240.3318 | 0.2342 | 4770.0779 | 22653.9403 | 0.21056 | 4.1851 | 3.7627 | 4774.8206 | 22668.1783 | 0.99085 | 0.99519 | 9.5178 | 9.5176 | 0.50438 | 0.047764 | 9.3976 | 9.3695 | 9.6412 | 9.6313 | 0.24368 | 0.26183 | 0.065232 | 0.057423 | 310.1026 | 359.039 | 0.0010349 | 0.0011104 | FALSE | 1 | TRUE | [M-H]- | 90.5023 | 9.4505 | 9.5872 | 0.13672 | 9.4588 | 9.5781 | 0.11936 | 20.2582 | 1.1284 | 1.2825 | 1.8067 | 0 | 17 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 19963.3012 | No Control Sample | 0 | off | off | 19963.3012 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9559 | 3.7219 | 2.2814 | 568.9667 | 5.8473 | 3.3269 | 0 | 0 | 0 | 0 | 0 | 376.9797 | FALSE | 0.25 | 0.2342
2980 | 2981 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFTrDA | 1.14941453246348 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.68 | 9.47 | Baseline | Baseline | 18481.3852 | 85240.3318 | 0.21682 | 5178.5337 | 22653.9403 | 0.22859 | 3.5688 | 3.7627 | 5184.0331 | 22668.1783 | 0.98939 | 0.99519 | 9.7357 | 9.5176 | 0.57528 | 0.055687 | 9.5839 | 9.3695 | 9.8341 | 9.6313 | 0.25022 | 0.26183 | 0.056366 | 0.057423 | 486.8431 | 359.039 | 0.0017539 | 0.0011104 | FALSE | 1.0229 | TRUE | [M-H]- | 114.9415 | 9.6778 | 9.7936 | 0.11585 | 9.6852 | 9.7863 | 0.10104 | 36.2991 | 1.0991 | 1.2317 | 1.0752 | 0 | 22 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 18481.3852 | No Control Sample | 0 | off | off | 18481.3852 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9532 | 4.1877 | 2.7762 | 618.9626 | 1.4039 | 0.86893 | 0 | 0 | 0 | 0 | 0 | 4.0909 | FALSE | 0.25 | 0.21682
2982 | 2983 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFTeDA | 0.799649702870833 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy, Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.87 | 9.86 | Baseline | Baseline | 7137.8474 | 25176.4138 | 0.28351 | 2167.9566 | 7653.9984 | 0.28324 | 3.2924 | 3.2893 | 2168.1203 | 7657.7316 | 0.95765 | 0.99658 | 9.9128 | 9.9115 | 0.43325 | 0.042762 | 9.8047 | 9.8068 | 10.0097 | 10.0106 | 0.205 | 0.20375 | 0.050952 | 0.049875 | 307.4236 | 934.4216 | 0.00013515 | 0.00096164 | FALSE | 1.0001 | TRUE | [M-H]- | 79.965 | 9.861 | 9.9704 | 0.10934 | 9.8674 | 9.963 | 0.095654 | 1.4292 | 1.0035 | 0.98628 | 1.6489 | 0 | 23 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 7137.8474 | No Control Sample | 0 | off | off | 7137.8474 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9502 | 4.0563 | 2.8919 | 668.961 | 4.6925 | 3.1391 | 0 | 0 | 0 | 0 | 0 | 413.6793 | FALSE | 0.25 | 0.28351
2989 | 2990 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFBS | 1.12564662259211 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 3.5 | 3.5 | Manual | Manual | 12169.1591 | 40843.6509 | 0.29794 | 1087.3316 | 3930.6635 | 0.27663 | 11.1918 | 10.391 | 1091.5576 | 3942.1406 | 0.97639 | 0.98708 | 3.8355 | 3.8308 | 9.5858 | 0.3355 | 3.6211 | 3.5906 | 4.0795 | 4.1525 | 0.45839 | 0.56189 | 0.16922 | 0.15438 | 156.4201 | 839.5533 | 0.0008619 | 0.00065847 | TRUE | 1.0012 | TRUE | [M-H]- | 112.5647 | 3.6695 | 4.0591 | 0.38968 | 3.6837 | 4.0141 | 0.33036 | -2.0445 | 1.1031 | 1.0347 | 11.2771 | 0 | 72 | 28 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 12169.1591 | No Control Sample | 0 | off | off | 12169.1591 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9446 | 5.5135 | 1.6482 | 79.9579 | 8.3387 | 0.66674 | 0 | 0 | 0 | 0 | 0 | 4.9267 | FALSE | 0.25 | 0.29794
2991 | 2992 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFPeS | 1.23417848124466 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 6.38 | 3.5 | Baseline | Manual | 13434.3502 | 40843.6509 | 0.32892 | 2698.1455 | 3930.6635 | 0.68644 | 4.9791 | 10.391 | 2699.1731 | 3942.1406 | 0.99017 | 0.98708 | 6.3955 | 3.8308 | 0.24367 | 0.015546 | 6.2481 | 3.5906 | 6.5531 | 4.1525 | 0.305 | 0.56189 | 0.076959 | 0.15438 | 208.1976 | 839.5533 | 0.00080039 | 0.00065847 | FALSE | 1.6695 | TRUE | [M-H]- | 123.4178 | 6.3086 | 6.4803 | 0.17167 | 6.3204 | 6.4653 | 0.14488 | 7.0806 | 1.0187 | 0.99853 | 13.4734 | 0 | 33 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 13434.3502 | No Control Sample | 0 | off | off | 13434.3502 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9416 | 5.054 | 1.7636 | 79.9569 | -4.2678 | -0.34124 | 0 | 0 | 0 | 0 | 0 | 1.0499 | FALSE | 0.25 | 0.32892
2993 | 2994 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | L-PFHxS | 0.89490634334673 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.811 | 4 | 0.20275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 10114.5633 | 60507.0471 | 0.16716 | 2598.5865 | 15118.2513 | 0.17188 | 3.8923 | 4.0023 | 2618.0935 | 15146.4876 | 0.76795 | 0.98785 | 7.3743 | 7.3715 | 0.46707 | 0.034283 | 7.283 | 7.2589 | 7.4796 | 7.4813 | 0.19658 | 0.2224 | 0.059498 | 0.062414 | 89.3514 | 171.0043 | 0.014266 | 0.0035084 | FALSE | 1.0004 | TRUE | [M-H]- | 110.346 | 7.3084 | 7.444 | 0.13558 | 7.3149 | 7.4319 | 0.11706 | -188.5808 | 0.98526 | 0.87823 | 14.8356 | 0 | 24 | 9 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10114.5633 | No Control Sample | 0 | off | off | 10114.5633 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9386 | 5.0728 | 2.0237 | 79.9569 | -4.2559 | -0.34029 | 0 | 0 | 0 | 0 | 0 | 27.3616 | FALSE | 0.20275 | 0.16716
2995 | 2996 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | Br-PFHxS | 0.173987461165628 | Qualifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.189 | 4 | 0.04725 | 0 | 7.16 | 7.35 | Manual | Baseline | 2671.7222 | 60507.0471 | 0.044156 | 498.2478 | 15118.2513 | 0.032957 | 5.3622 | 4.0023 | 2618.0935 | 15146.4876 | 0.20311 | 0.98785 | 7.1783 | 7.3715 | 0.25498 | 0.018256 | 7.0886 | 7.2589 | 7.2837 | 7.4813 | 0.1951 | 0.2224 | 0.084132 | 0.062414 | 17.1321 | 171.0043 | 0.0095536 | 0.0035084 | TRUE | 0.97379 | TRUE | [M-H]- | 92.0569 | 7.0991 | 7.2916 | 0.19247 | 7.1062 | 7.2711 | 0.1649 | 24.3983 | 1.1571 | 1.1669 | 0.26415 | 0 | 23 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2671.7222 | No Control Sample | 0 | off | off | 10114.5633 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9386 | 5.0728 | 2.0237 | 79.9547 | -31.2507 | -2.4987 | 0 | 0 | 0 | 0 | 0 | 27.3616 | FALSE | 0.04725 | 0.044156
2997 | 2998 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFHpS | 1.18203805617355 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 7.98 | 7.35 | Baseline | Baseline | 10885.1546 | 60507.0471 | 0.1799 | 2384.7746 | 15118.2513 | 0.15774 | 4.5644 | 4.0023 | 2385.3071 | 15146.4876 | 0.99173 | 0.98785 | 8.0142 | 7.3715 | 0.42914 | 0.034246 | 7.8891 | 7.2589 | 8.1929 | 7.4813 | 0.3038 | 0.2224 | 0.071593 | 0.062414 | 127.5005 | 171.0043 | 0.00052593 | 0.0035084 | FALSE | 1.0872 | TRUE | [M-H]- | 118.2038 | 7.9424 | 8.0951 | 0.15273 | 7.9526 | 8.0842 | 0.13157 | 4.1285 | 1.0369 | 1.0741 | 18.6976 | 0 | 22 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10885.1546 | No Control Sample | 0 | off | off | 10885.1546 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9352 | 3.8961 | 1.7491 | 79.9568 | -4.6234 | -0.36967 | 0 | 0 | 0 | 0 | 0 | 1.3748 | FALSE | 0.25 | 0.1799
2999 | 3000 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | L-PFOS | 0.744312548046531 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 0.788 | 4 | 0.197 | 0 | 8.46 | 8.45 | Manual | Baseline | 7890.1712 | 40782.6799 | 0.19347 | 1666.2101 | 9168.3002 | 0.18174 | 4.7354 | 4.4482 | 1683.1259 | 9175.3908 | 0.66414 | 0.99362 | 8.5002 | 8.5016 | 0.47487 | 0.040174 | 8.418 | 8.3665 | 8.596 | 8.6748 | 0.17806 | 0.30838 | 0.073841 | 0.069351 | 49.9496 | 167.0249 | 0 | 0.00019773 | TRUE | 0.99984 | TRUE | [M-H]- | 94.4559 | 8.415 | 8.5775 | 0.16244 | 8.415 | 8.5665 | 0.15149 | 0 | 0.96329 | 0.79666 | 15.9015 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 7890.1712 | No Control Sample | 0 | off | off | 7890.1712 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9326 | 4.8489 | 2.4192 | 79.9569 | -4.5508 | -0.36387 | 0 | 0 | 0 | 0 | 0 | 1.377 | FALSE | 0.197 | 0.19347
3001 | 3002 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | Br-PFOS | 0.151029059562657 | Qualifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.211 | 4 | 0.05275 | 0 | 8.3 | 8.45 | Manual | Baseline | 3427.4698 | 40782.6799 | 0.084042 | 541.1143 | 9168.3002 | 0.05902 | 6.3341 | 4.4482 | 1683.1259 | 9175.3908 | 0.28896 | 0.99362 | 8.3482 | 8.5016 | 0.58033 | 0.048168 | 8.2509 | 8.3665 | 8.4247 | 8.6748 | 0.17382 | 0.30838 | 0.10806 | 0.069351 | 16.2215 | 167.0249 | 0 | 0.00019773 | TRUE | 0.98196 | TRUE | [M-H]- | 71.5778 | 8.2616 | 8.4364 | 0.17484 | 8.273 | 8.4364 | 0.16335 | 0 | 1.1138 | 1.438 | 0.4344 | 0 | 12 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3427.4698 | No Control Sample | 0 | off | off | 7890.1712 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9326 | 4.8489 | 2.4192 | 79.9594 | 26.8189 | 2.1444 | 0 | 0 | 0 | 0 | 0 | 1.377 | FALSE | 0.05275 | 0.084042
3003 | 3004 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFNS | 1.20910296250796 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy, Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.85 | 8.45 | Baseline | Baseline | 10929.4718 | 40782.6799 | 0.26799 | 2266.3192 | 9168.3002 | 0.24719 | 4.8226 | 4.4482 | 2266.3192 | 9175.3908 | 0.99642 | 0.99362 | 8.9017 | 8.5016 | 0.58414 | 0.051696 | 8.651 | 8.3665 | 9.1731 | 8.6748 | 0.52215 | 0.30838 | 0.074655 | 0.069351 | 116.2659 | 167.0249 | 0 | 0.00019773 | FALSE | 1.0471 | TRUE | [M-H]- | 120.9103 | 8.8266 | 8.986 | 0.15932 | 8.8334 | 8.9725 | 0.13907 | 0 | 0.9562 | 0.81787 | 17.1249 | 0 | 27 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Fail | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10929.4718 | No Control Sample | 0 | off | off | 10929.4718 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9291 | 3.8291 | 2.1019 | 79.9568 | -4.3456 | -0.34747 | 0 | 0 | 0 | 0 | 0 | 1.2792 | FALSE | 0.25 | 0.26799
3005 | 3006 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFDS | 1.07799788947798 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.19 | 8.45 | Baseline | Baseline | 10011.9322 | 40782.6799 | 0.24549 | 2186.6301 | 9168.3002 | 0.2385 | 4.5787 | 4.4482 | 2189.3097 | 9175.3908 | 0.97298 | 0.99362 | 9.2324 | 8.5016 | 0.46119 | 0.042384 | 9.0955 | 8.3665 | 9.401 | 8.6748 | 0.30552 | 0.30838 | 0.0717 | 0.069351 | 110.7553 | 167.0249 | 0.0027017 | 0.00019773 | FALSE | 1.086 | TRUE | [M-H]- | 107.7998 | 9.1581 | 9.3121 | 0.15402 | 9.1661 | 9.3024 | 0.13631 | 19.3362 | 1.0235 | 1.0281 | 16.8223 | 0 | 17 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10011.9322 | No Control Sample | 0 | off | off | 10011.9322 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9261 | 3.7314 | 2.2348 | 79.9569 | -0.58307 | -0.046621 | 0 | 0 | 0 | 0 | 0 | 1.5633 | FALSE | 0.25 | 0.24549
3009 | 3010 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FBSA | 0.907711385588335 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 4.6 | 8.86 | Manual | Baseline | 16666.877 | 69635.9399 | 0.23934 | 1597.3824 | 15379.4698 | 0.10386 | 10.4339 | 4.5279 | 1610.1363 | 15382.3287 | 0.96372 | 0.99767 | 4.9464 | 9.0428 | 7.5305 | 0.3464 | 4.7476 | 8.8762 | 5.1796 | 9.239 | 0.43208 | 0.36273 | 0.17518 | 0.068539 | 156.8469 | 135.3332 | 0.0050652 | 0.000374 | TRUE | 0.547 | TRUE | [M-H]- | 90.7711 | 4.7624 | 5.1531 | 0.39066 | 4.7878 | 5.1173 | 0.32949 | -18.7262 | 1.1611 | 1.3064 | 5.5651 | 0 | 78 | 33 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 16666.877 | No Control Sample | 0 | off | off | 16666.877 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9605 | 5.0415 | 1.5022 | 77.9656 | 3.6555 | 0.285 | 0 | 0 | 0 | 0 | 0 | 1.8841 | FALSE | 0.25 | 0.23934
3011 | 3012 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FPeSA | 1.12495714803103 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 7 | 8.86 | Baseline | Baseline | 32076.6321 | 69635.9399 | 0.46063 | 7022.8789 | 15379.4698 | 0.45664 | 4.5674 | 4.5279 | 7047.6857 | 15382.3287 | 0.98053 | 0.99767 | 7.1158 | 9.0428 | 1.6537 | 0.11576 | 6.9806 | 8.8762 | 7.2383 | 9.239 | 0.25765 | 0.36273 | 0.071118 | 0.068539 | 185.0115 | 135.3332 | 0.0066946 | 0.000374 | FALSE | 0.7869 | TRUE | [M-H]- | 112.4957 | 7.0375 | 7.1987 | 0.16121 | 7.0492 | 7.1872 | 0.13802 | 182.4779 | 1.0403 | 1.0984 | 3.8103 | 0 | 28 | 9 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 32076.6321 | No Control Sample | 0 | off | off | 32076.6321 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9574 | 4.6339 | 1.6124 | 77.9648 | -6.6032 | -0.51482 | 0 | 0 | 0 | 0 | 0 | 1.1615 | FALSE | 0.25 | 0.46063
3013 | 3014 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FHxSA | 1.25515062723484 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 7.86 | 8.86 | Baseline | Baseline | 17859.4474 | 69635.9399 | 0.25647 | 3849.5396 | 15379.4698 | 0.2503 | 4.6394 | 4.5279 | 3849.5894 | 15382.3287 | 0.99536 | 0.99767 | 8.0175 | 9.0428 | 2.0033 | 0.15746 | 7.7938 | 8.8762 | 8.207 | 9.239 | 0.41322 | 0.36273 | 0.072574 | 0.068539 | 109.4613 | 135.3332 | 2.3041e-05 | 0.000374 | FALSE | 0.88661 | TRUE | [M-H]- | 125.5151 | 7.9484 | 8.0992 | 0.15087 | 7.9554 | 8.089 | 0.13362 | 0.21465 | 1.1595 | 1.304 | 8.8176 | 0 | 30 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 17859.4474 | No Control Sample | 0 | off | off | 17859.4474 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9542 | 4.0563 | 1.6142 | 77.9651 | -4.383 | -0.34172 | 0 | 0 | 0 | 0 | 0 | 0.67212 | FALSE | 0.25 | 0.25647
3015 | 3016 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FHpSA | 1.13014638073522 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.58 | 8.86 | Baseline | Baseline | 33307.684 | 69635.9399 | 0.47831 | 7436.2258 | 15379.4698 | 0.48352 | 4.4791 | 4.5279 | 7439.3947 | 15382.3287 | 0.99357 | 0.99767 | 8.6009 | 9.0428 | 0.24368 | 0.020908 | 8.3944 | 8.8762 | 8.7606 | 9.239 | 0.36618 | 0.36273 | 0.069076 | 0.068539 | 117.3729 | 135.3332 | 0.00074039 | 0.000374 | FALSE | 0.95113 | TRUE | [M-H]- | 113.0146 | 8.5253 | 8.6793 | 0.15404 | 8.5353 | 8.6657 | 0.13039 | 15.0354 | 1.0412 | 1.0388 | 4.1964 | 0 | 21 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 33307.684 | No Control Sample | 0 | off | off | 33307.684 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9516 | 4.9418 | 2.2137 | 77.965 | -4.6204 | -0.36023 | 0 | 0 | 0 | 0 | 0 | 1.2818 | FALSE | 0.25 | 0.47831
3017 | 3018 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FOSA | 1.25047577585735 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.88 | 8.86 | Baseline | Baseline | 19813.1805 | 69635.9399 | 0.28453 | 4322.8765 | 15379.4698 | 0.28108 | 4.5833 | 4.5279 | 4322.9631 | 15382.3287 | 0.98793 | 0.99767 | 9.0415 | 9.0428 | 1.8186 | 0.1615 | 8.8663 | 8.8762 | 9.2316 | 9.239 | 0.36535 | 0.36273 | 0.070963 | 0.068539 | 126.8923 | 135.3332 | 4.1606e-05 | 0.000374 | FALSE | 0.99985 | TRUE | [M-H]- | 125.0476 | 8.9621 | 9.1205 | 0.15836 | 8.9753 | 9.1088 | 0.13348 | 0.49229 | 1.0661 | 1.185 | 7.0765 | 0 | 19 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 19813.1805 | No Control Sample | 0 | off | off | 19813.1805 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9484 | 4.4161 | 2.199 | 77.965 | -5.1613 | -0.40241 | 0 | 0 | 0 | 0 | 0 | 0.44028 | FALSE | 0.25 | 0.28453
3036 | 3037 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 4:2 FTS | 1.0589008214774 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 5.94 | 5.92 | Baseline | Baseline | 9510.0111 | 17687.9081 | 0.53766 | 1592.2203 | 2957.4365 | 0.53838 | 5.9728 | 5.9808 | 1599.6258 | 2960.2085 | 0.9797 | 0.99292 | 5.9468 | 5.9434 | 0.11491 | 0.0068255 | 5.8092 | 5.7685 | 6.1076 | 6.1385 | 0.29847 | 0.37 | 0.099557 | 0.093432 | 167.9621 | 210.9829 | 0.0088541 | 0.0019706 | FALSE | 1.0006 | TRUE | [M-H]- | 105.8901 | 5.839 | 6.0367 | 0.19764 | 5.8524 | 6.0288 | 0.1764 | -47.2334 | 0.86845 | 0.75643 | 4.9317 | 0 | 35 | 14 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9510.0111 | No Control Sample | 0 | off | off | 9510.0111 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9765 | 6.6858 | 2.1861 | 306.969 | 4.6235 | 1.4193 | 0 | 0 | 0 | 0 | 0 | 8.2417 | FALSE | 0.25 | 0.53766
3038 | 3039 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 6:2 FTS | 1.0801660894157 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 7.92 | 7.92 | Baseline | Baseline | 10049.4433 | 15334.1234 | 0.65536 | 2337.1172 | 3529.3768 | 0.66219 | 4.2999 | 4.3447 | 2337.1172 | 3529.9286 | 1 | 0.99816 | 7.9508 | 7.9504 | 0.38914 | 0.03082 | 7.7528 | 7.8148 | 8.1787 | 8.0892 | 0.42587 | 0.2745 | 0.065176 | 0.066629 | 77.573 | 114.1893 | 0 | 0.00033345 | FALSE | 1.0001 | TRUE | [M-H]- | 108.0166 | 7.8814 | 8.0316 | 0.15017 | 7.8919 | 8.0174 | 0.12553 | 0 | 1.0688 | 1.1003 | 5.1438 | 0 | 31 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10049.4433 | No Control Sample | 0 | off | off | 10049.4433 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9701 | 5.2168 | 2.2274 | 406.9633 | 4.3555 | 1.7725 | 0 | 0 | 0 | 0 | 0 | 5.5018 | FALSE | 0.25 | 0.65536
3040 | 3041 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 8:2 FTS | 1.13943114261931 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.88 | 8.89 | Baseline | Baseline | 10360.2062 | 18531.3353 | 0.55906 | 2127.3717 | 3701.3275 | 0.57476 | 4.87 | 5.0067 | 2127.3717 | 3703.3015 | 1 | 0.9933 | 8.911 | 8.9111 | 0.34898 | 0.03099 | 8.6875 | 8.774 | 9.1357 | 9.0481 | 0.4482 | 0.27408 | 0.076291 | 0.078189 | 234.6599 | 166.8358 | 0 | 0.0009269 | FALSE | 0.99999 | TRUE | [M-H]- | 113.9431 | 8.8328 | 8.9969 | 0.16416 | 8.8424 | 8.9877 | 0.14528 | 0 | 1.0399 | 1.0961 | 5.1267 | 0 | 23 | 5 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10360.2062 | No Control Sample | 0 | off | off | 10360.2062 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9642 | 5.1457 | 2.7116 | 506.9578 | 4.1413 | 2.0995 | 0 | 0 | 0 | 0 | 0 | 5.7212 | FALSE | 0.25 | 0.55906
3042 | 3043 | PFAS CS5 1ng/ml | 22 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 23) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 10:2 FTS | 0.894309245021685 | Quantifiers | 10/27/2023 00:56:19 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.5 | 8.89 | Baseline | Baseline | 8465.01 | 18531.3353 | 0.45679 | 2078.5845 | 3701.3275 | 0.56158 | 4.0725 | 5.0067 | 2078.5845 | 3703.3015 | 0.99644 | 0.9933 | 9.5345 | 8.9111 | 0.36286 | 0.034472 | 9.3505 | 8.774 | 9.7092 | 9.0481 | 0.35875 | 0.27408 | 0.063782 | 0.078189 | 200.9964 | 166.8358 | 0 | 0.0009269 | FALSE | 1.07 | TRUE | [M-H]- | 89.4309 | 9.4702 | 9.6025 | 0.13222 | 9.4766 | 9.5937 | 0.11707 | 0 | 1.1307 | 1.2489 | 4.193 | 0 | 25 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8465.01 | No Control Sample | 0 | off | off | 8465.01 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9575 | 3.8199 | 2.3949 | 606.9549 | 11.4533 | 6.9515 | 0 | 0 | 0 | 0 | 0 | 1.7203 | FALSE | 0.25 | 0.45679
3103 | 3104 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFBA | 1.10774678698103 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 1.49 | 1.48 | Manual | Baseline | 20568.4737 | 103514.2252 | 0.1987 | 2924.6757 | 14641.7069 | 0.19975 | 7.0327 | 7.0698 | 2941.6152 | 14698.1088 | 0.93943 | 0.98489 | 1.4671 | 1.4697 | 1.5354 | 0.022878 | 1.3585 | 1.3451 | 1.6518 | 1.7044 | 0.2933 | 0.35923 | 0.11493 | 0.11735 | 129.3849 | 208.5006 | 0.0050603 | 0.0005833 | TRUE | 0.99828 | TRUE | [M-H]- | 110.7747 | 1.3794 | 1.6221 | 0.24261 | 1.3862 | 1.5992 | 0.21297 | 50.459 | 1.332 | 1.5254 | 1.5296 | 0 | 60 | 25 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 20568.4737 | No Control Sample | 0 | off | off | 20568.4737 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9804 | 5.5238 | 1.1765 | 168.9899 | 5.3871 | 0.91036 | 0 | 0 | 0 | 0 | 0 | 11.7225 | FALSE | 0.25 | 0.1987
3105 | 3106 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFPeA | 1.02234185309624 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 3.18 | 3.21 | Manual | Baseline | 17291.1756 | 75627.6054 | 0.22864 | 1846.9665 | 7878.4165 | 0.23443 | 9.3619 | 9.5993 | 1853.5469 | 7955.1305 | 0.97026 | 0.968 | 3.2144 | 3.212 | 1.0822 | 0.034416 | 3.0108 | 3.025 | 3.5043 | 3.4596 | 0.49351 | 0.43463 | 0.13064 | 0.15452 | 264.27 | 694.7257 | 0.0032438 | 0.0029532 | TRUE | 1.0007 | TRUE | [M-H]- | 102.2342 | 3.0702 | 3.3945 | 0.32431 | 3.0904 | 3.3812 | 0.29076 | -12.1397 | 1.1682 | 1.4525 | 1.9475 | 0 | 86 | 25 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 17291.1756 | No Control Sample | 0 | off | off | 17291.1756 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9775 | 5.6531 | 1.4866 | 218.9873 | 6.8009 | 1.4893 | 0 | 0 | 0 | 0 | 0 | 0.53211 | FALSE | 0.25 | 0.22864
3107 | 3108 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFHxA | 1.05866683105669 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 18.1581 | 6.13 | 6.12 | Baseline | Baseline | 37014.1244 | 150231.1993 | 0.24638 | 6727.3577 | 27807.9395 | 0.24192 | 5.502 | 5.4025 | 6746.2698 | 27819.1749 | 0.98243 | 0.99609 | 6.1456 | 6.1457 | 0.2548 | 0.015619 | 5.9839 | 5.9443 | 6.2864 | 6.365 | 0.30247 | 0.42072 | 0.083207 | 0.084438 | 182.2608 | 234.2185 | 0.0053061 | 0.00083262 | FALSE | 0.99999 | TRUE | [M-H]- | 105.8667 | 6.0502 | 6.2405 | 0.19028 | 6.0645 | 6.2272 | 0.16269 | 118.0162 | 1.0392 | 1.1068 | 1.5347 | 0 | 33 | 10 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 37014.1244 | No Control Sample | 203843.5852 | off | off | 37014.1244 | 203843.5852 | off | 1 | off | 1 | off | 0.18158 | off | 312.9747 | 6.0625 | 1.8974 | 268.984 | 5.8049 | 1.5614 | 0 | 0 | 0 | 0 | 0 | 176.4342 | FALSE | 0.25 | 0.24638
3109 | 3110 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFHpA | 1.22663084289245 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 22.5149 | 7.26 | 7.27 | Baseline | Baseline | 51407.9219 | 167489.5153 | 0.30693 | 12149.1622 | 40692.8914 | 0.29856 | 4.2314 | 4.1159 | 12152.3319 | 40723.1604 | 0.99413 | 0.99492 | 7.2922 | 7.2905 | 0.443 | 0.032162 | 7.1494 | 7.1615 | 7.4924 | 7.4347 | 0.34295 | 0.27323 | 0.064521 | 0.064562 | 284.0541 | 228.2392 | 0.00039081 | 0.0014483 | FALSE | 1.0002 | TRUE | [M-H]- | 122.6631 | 7.2239 | 7.3629 | 0.13899 | 7.2331 | 7.3542 | 0.1211 | 13.8447 | 1.0983 | 1.2396 | 1.4543 | 0 | 40 | 10 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 51407.9219 | No Control Sample | 228328.7703 | off | off | 51407.9219 | 228328.7703 | off | 1 | off | 1 | off | 0.22515 | off | 362.9716 | 5.5482 | 2.0138 | 318.9806 | 4.7934 | 1.529 | 0 | 0 | 0 | 0 | 0 | 1.6536 | FALSE | 0.25 | 0.30693
3111 | 3112 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFOA | 1.19771799285881 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 24.6939 | 7.96 | 7.94 | Baseline | Baseline | 47282.7159 | 162779.2797 | 0.29047 | 11011.0523 | 38781.4573 | 0.28393 | 4.2941 | 4.1973 | 11022.202 | 38792.2707 | 0.98733 | 0.9972 | 7.978 | 7.9787 | 0.22588 | 0.01798 | 7.8624 | 7.8416 | 8.1443 | 8.1835 | 0.28193 | 0.34185 | 0.06633 | 0.064354 | 104.0738 | 125.6832 | 0.00093409 | 0.00066749 | FALSE | 0.99991 | TRUE | [M-H]- | 119.7718 | 7.9135 | 8.0571 | 0.14364 | 7.9214 | 8.0455 | 0.12409 | -36.4812 | 1.1865 | 1.3566 | 1.6517 | 0 | 21 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 47282.7159 | No Control Sample | 191475.4056 | off | off | 47282.7159 | 191475.4056 | off | 1 | off | 1 | off | 0.24694 | off | 412.9686 | 5.3735 | 2.2191 | 368.9778 | 4.6581 | 1.7187 | 0 | 0 | 0 | 0 | 0 | 1.9955 | FALSE | 0.25 | 0.29047
3113 | 3114 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFNA | 1.04527055502359 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 21.3408 | 8.46 | 8.47 | Baseline | Baseline | 33962.3934 | 139864.1004 | 0.24282 | 6821.0479 | 30553.0154 | 0.22325 | 4.9791 | 4.5778 | 6826.0474 | 30567.4253 | 0.99007 | 0.99628 | 8.4861 | 8.4847 | 0.30875 | 0.026121 | 8.3381 | 8.2806 | 8.6661 | 8.6393 | 0.32793 | 0.35863 | 0.077792 | 0.071024 | 174.7398 | 143.6726 | 0.0016965 | 0.00081138 | FALSE | 1.0002 | TRUE | [M-H]- | 104.5271 | 8.4038 | 8.5642 | 0.16041 | 8.4145 | 8.5554 | 0.14092 | 35.2873 | 1.0976 | 1.2594 | 1.467 | 0 | 19 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 33962.3934 | No Control Sample | 159143.3885 | off | off | 33962.3934 | 159143.3885 | off | 1 | off | 1 | off | 0.21341 | off | 462.9657 | 5.4379 | 2.5176 | 418.9755 | 5.4635 | 2.289 | 0 | 0 | 0 | 0 | 0 | 4.2714 | FALSE | 0.25 | 0.24282
3115 | 3116 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFDA | 1.08725011766645 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.87 | 8.86 | Baseline | Baseline | 35233.519 | 130540.9645 | 0.2699 | 7329.1242 | 27544.3302 | 0.26608 | 4.8073 | 4.7393 | 7337.547 | 27545.2748 | 0.98925 | 0.99923 | 8.8914 | 8.8931 | 0.24143 | 0.021415 | 8.7416 | 8.693 | 9.0539 | 9.1593 | 0.3123 | 0.46637 | 0.075149 | 0.073642 | 177.3268 | 123.3897 | 0.0017876 | 7.8298e-05 | FALSE | 0.99981 | TRUE | [M-H]- | 108.725 | 8.8148 | 8.9785 | 0.16372 | 8.8259 | 8.9669 | 0.14104 | 41.9518 | 1.1432 | 1.3321 | 1.6447 | 0 | 16 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 35233.519 | No Control Sample | 0 | off | off | 35233.519 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9626 | 5.0893 | 2.6106 | 468.9717 | 3.6228 | 1.699 | 0 | 0 | 0 | 0 | 0 | 82.2524 | FALSE | 0.25 | 0.2699
3117 | 3118 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFUdA | 1.27668650650649 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.21 | 9.2 | Baseline | Baseline | 26043.678 | 91986.0225 | 0.28313 | 5724.1609 | 19497.874 | 0.29358 | 4.5498 | 4.7177 | 5732.8122 | 19504.9949 | 0.9893 | 0.99614 | 9.2243 | 9.2253 | 0.15569 | 0.014339 | 9.0928 | 9.082 | 9.3803 | 9.4361 | 0.28747 | 0.35407 | 0.071153 | 0.074807 | 163.0088 | 171.5069 | 0.0030511 | 0.00092911 | FALSE | 0.99989 | TRUE | [M-H]- | 127.6687 | 9.1536 | 9.3034 | 0.1498 | 9.1607 | 9.2905 | 0.12986 | 60.7552 | 0.97767 | 0.86682 | 1.7351 | 0 | 16 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 26043.678 | No Control Sample | 0 | off | off | 26043.678 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9597 | 5.0298 | 2.8316 | 518.9703 | 5.6947 | 2.9554 | 0 | 0 | 0 | 0 | 0 | 0.30959 | FALSE | 0.25 | 0.28313
3119 | 3120 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFDoA | 0.944702351334928 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.47 | 9.47 | Baseline | Baseline | 20571.5336 | 84148.0839 | 0.24447 | 5119.3856 | 20804.7222 | 0.24607 | 4.0184 | 4.0447 | 5119.4385 | 20826.8247 | 0.995 | 0.99389 | 9.4912 | 9.4933 | 0.22373 | 0.021187 | 9.311 | 9.353 | 9.6564 | 9.6028 | 0.34545 | 0.2497 | 0.062216 | 0.063588 | 236.7074 | 284.9743 | 1.8234e-05 | 0.0019648 | FALSE | 0.99978 | TRUE | [M-H]- | 94.4702 | 9.4247 | 9.5589 | 0.13419 | 9.4337 | 9.5505 | 0.1168 | 0.27022 | 0.92682 | 0.84187 | 1.5582 | 0 | 23 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 20571.5336 | No Control Sample | 0 | off | off | 20571.5336 | 0 | off | 0 | off | 0 | off | 0 | off | 612.957 | 5.4719 | 3.354 | 568.9665 | 5.3636 | 3.0517 | 0 | 0 | 0 | 0 | 0 | 376.1142 | FALSE | 0.25 | 0.24447
3121 | 3122 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFTrDA | 0.924839064047499 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.68 | 9.47 | Baseline | Baseline | 14679.9001 | 84148.0839 | 0.17445 | 4172.1135 | 20804.7222 | 0.20054 | 3.5186 | 4.0447 | 4173.9496 | 20826.8247 | 0.99053 | 0.99389 | 9.7097 | 9.4933 | 0.30639 | 0.029659 | 9.5842 | 9.353 | 9.8144 | 9.6028 | 0.23013 | 0.2497 | 0.054811 | 0.063588 | 458.0182 | 284.9743 | 0.0008203 | 0.0019648 | FALSE | 1.0228 | TRUE | [M-H]- | 92.4839 | 9.6481 | 9.7678 | 0.11975 | 9.6579 | 9.7593 | 0.10145 | 14.8713 | 1.0129 | 1.0588 | 1.2904 | 0 | 20 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 14679.9001 | No Control Sample | 0 | off | off | 14679.9001 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9532 | 4.0696 | 2.6979 | 618.9642 | 4.0055 | 2.4793 | 0 | 0 | 0 | 0 | 0 | 7.4149 | FALSE | 0.25 | 0.17445
3123 | 3124 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFTeDA | 1.13006966221139 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.87 | 9.86 | Baseline | Baseline | 8040.1854 | 20067.2211 | 0.40066 | 2251.5059 | 5709.0658 | 0.39437 | 3.571 | 3.515 | 2251.6043 | 5709.0658 | 0.97256 | 0.99861 | 9.8863 | 9.8852 | 0.16508 | 0.016293 | 9.775 | 9.711 | 9.9864 | 10.0379 | 0.2114 | 0.32692 | 0.055398 | 0.054741 | 423.1779 | 1086.4678 | 8.114e-05 | 0 | FALSE | 1.0001 | TRUE | [M-H]- | 113.007 | 9.8301 | 9.9486 | 0.11852 | 9.8372 | 9.9413 | 0.1041 | 0.86418 | 1.0661 | 1.1479 | 1.4448 | 0 | 23 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8040.1854 | No Control Sample | 0 | off | off | 8040.1854 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9513 | 5.6367 | 4.0187 | 668.9621 | 6.2634 | 4.1899 | 0 | 0 | 0 | 0 | 0 | 399.6864 | FALSE | 0.25 | 0.40066
3130 | 3131 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFBS | 1.08478671525704 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 3.5 | 3.5 | Manual | Manual | 12129.0304 | 42242.3222 | 0.28713 | 1115.1747 | 3899.2323 | 0.286 | 10.8763 | 10.8335 | 1120.2592 | 3911.2819 | 0.98015 | 0.98683 | 3.8687 | 3.8585 | 10.5342 | 0.3687 | 3.634 | 3.6248 | 4.1469 | 4.1701 | 0.51294 | 0.54529 | 0.16603 | 0.1691 | 165.5247 | 816.416 | 0.0029679 | 0 | TRUE | 1.0026 | TRUE | [M-H]- | 108.4787 | 3.6793 | 4.0776 | 0.39831 | 3.7122 | 4.0367 | 0.32449 | -6.4525 | 1.034 | 1.0321 | 11.0939 | 0 | 80 | 27 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 12129.0304 | No Control Sample | 0 | off | off | 12129.0304 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9447 | 5.5994 | 1.6739 | 79.9569 | -3.2555 | -0.2603 | 0 | 0 | 0 | 0 | 0 | 1.1316 | FALSE | 0.25 | 0.28713
3132 | 3133 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFPeS | 1.15901481114237 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 6.38 | 3.5 | Baseline | Manual | 13048.2092 | 42242.3222 | 0.30889 | 2745.0346 | 3899.2323 | 0.70399 | 4.7534 | 10.8335 | 2746.0035 | 3911.2819 | 0.99537 | 0.98683 | 6.3964 | 3.8585 | 0.25699 | 0.016396 | 6.247 | 3.6248 | 6.5704 | 4.1701 | 0.32345 | 0.54529 | 0.073694 | 0.1691 | 218.9204 | 816.416 | 0.00078574 | 0 | FALSE | 1.6577 | TRUE | [M-H]- | 115.9015 | 6.3197 | 6.4832 | 0.16354 | 6.3293 | 6.4702 | 0.14082 | 6.6683 | 1.1391 | 1.2681 | 13.9876 | 0 | 35 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 13048.2092 | No Control Sample | 0 | off | off | 13048.2092 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9414 | 4.5755 | 1.5966 | 79.9568 | -4.3801 | -0.35022 | 0 | 0 | 0 | 0 | 0 | 1.4102 | FALSE | 0.25 | 0.30889
3134 | 3135 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | L-PFHxS | 0.861430056343338 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.811 | 4 | 0.20275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 9946.9808 | 61816.9631 | 0.16091 | 2535.3931 | 14707.4849 | 0.17239 | 3.9232 | 4.2031 | 2557.6361 | 14711.9321 | 0.79469 | 0.99513 | 7.3676 | 7.3674 | 0.37558 | 0.027568 | 7.2803 | 7.2382 | 7.4768 | 7.5965 | 0.19655 | 0.35825 | 0.060529 | 0.064865 | 92.4686 | 195.7059 | 0.014478 | 0.00017617 | FALSE | 1 | TRUE | [M-H]- | 106.2183 | 7.3001 | 7.4357 | 0.13555 | 7.3105 | 7.4285 | 0.11804 | -186.7563 | 1.0297 | 1.1283 | 14.8176 | 0 | 24 | 9 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9946.9808 | No Control Sample | 0 | off | off | 9946.9808 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9384 | 4.6002 | 1.8352 | 79.9573 | 1.4675 | 0.11734 | 0 | 0 | 0 | 0 | 0 | 25.3869 | FALSE | 0.20275 | 0.16091
3136 | 3137 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | Br-PFHxS | 0.137821937695659 | Qualifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.189 | 4 | 0.04725 | 0 | 7.16 | 7.35 | Manual | Baseline | 2162.1878 | 61816.9631 | 0.034977 | 405.4021 | 14707.4849 | 0.027564 | 5.3334 | 4.2031 | 2557.6361 | 14711.9321 | 0.17274 | 0.99513 | 7.1886 | 7.3674 | 0.39977 | 0.028624 | 7.0878 | 7.2382 | 7.2774 | 7.5965 | 0.18958 | 0.35825 | 0.084029 | 0.064865 | 14.7855 | 195.7059 | 0.0098469 | 0.00017617 | TRUE | 0.97574 | TRUE | [M-H]- | 72.9217 | 7.1015 | 7.2803 | 0.17877 | 7.1069 | 7.2695 | 0.16258 | 21.0564 | 0.9213 | 0.77471 | 0.21737 | 0 | 22 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2162.1878 | No Control Sample | 0 | off | off | 9946.9808 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9384 | 4.6002 | 1.8352 | 79.9541 | -39.3328 | -3.1449 | 0 | 0 | 0 | 0 | 0 | 25.3869 | FALSE | 0.04725 | 0.034977
3138 | 3139 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFHpS | 1.22287372536594 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 7.98 | 7.35 | Baseline | Baseline | 11504.9958 | 61816.9631 | 0.18611 | 2564.6587 | 14707.4849 | 0.17438 | 4.486 | 4.2031 | 2564.6612 | 14711.9321 | 0.99502 | 0.99513 | 7.999 | 7.3674 | 0.23852 | 0.019034 | 7.8514 | 7.2382 | 8.1796 | 7.5965 | 0.32822 | 0.35825 | 0.070461 | 0.064865 | 129.3123 | 195.7059 | 2.0873e-06 | 0.00017617 | FALSE | 1.0857 | TRUE | [M-H]- | 122.2874 | 7.9365 | 8.082 | 0.14551 | 7.9445 | 8.0715 | 0.12695 | 0.01631 | 1.1069 | 1.1997 | 17.7163 | 0 | 24 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 11504.9958 | No Control Sample | 0 | off | off | 11504.9958 | 0 | off | 0 | off | 0 | off | 0 | off | 448.935 | 3.6059 | 1.6188 | 79.9574 | 2.4217 | 0.19364 | 0 | 0 | 0 | 0 | 0 | 1.0606 | FALSE | 0.25 | 0.18611
3140 | 3141 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | L-PFOS | 0.900033912682261 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 0.788 | 4 | 0.197 | 0 | 8.46 | 8.45 | Manual | Baseline | 8908.3164 | 38078.6403 | 0.23395 | 1862.4384 | 7713.0912 | 0.24146 | 4.7831 | 4.9369 | 1882.054 | 7714.5359 | 0.68465 | 0.99487 | 8.4844 | 8.482 | 0.28878 | 0.02443 | 8.4001 | 8.3502 | 8.5754 | 8.7142 | 0.17529 | 0.364 | 0.075536 | 0.07689 | 54.9821 | 128.7496 | 0.010156 | 5.7172e-05 | TRUE | 1.0003 | TRUE | [M-H]- | 114.2175 | 8.3985 | 8.5554 | 0.15688 | 8.3995 | 8.5476 | 0.14809 | -107.9058 | 0.94791 | 0.81271 | 13.8074 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8908.3164 | No Control Sample | 0 | off | off | 8908.3164 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9324 | 4.2937 | 2.1422 | 79.9566 | -8.5587 | -0.68433 | 0 | 0 | 0 | 0 | 0 | 3.7627 | FALSE | 0.197 | 0.23395
3142 | 3143 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | Br-PFOS | 0.161718636710603 | Qualifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.211 | 4 | 0.05275 | 0 | 8.3 | 8.45 | Manual | Baseline | 3426.7219 | 38078.6403 | 0.089991 | 657.1069 | 7713.0912 | 0.085194 | 5.2149 | 4.9369 | 1882.054 | 7714.5359 | 0.26412 | 0.99487 | 8.3337 | 8.482 | 0.40594 | 0.033693 | 8.2366 | 8.3502 | 8.4018 | 8.7142 | 0.16513 | 0.364 | 0.084112 | 0.07689 | 19.3988 | 128.7496 | 0.0092381 | 5.7172e-05 | TRUE | 0.98251 | TRUE | [M-H]- | 76.6439 | 8.2529 | 8.4172 | 0.1643 | 8.2628 | 8.4172 | 0.15446 | 36.7613 | 1.1595 | 1.5317 | 0.38467 | 0 | 12 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3426.7219 | No Control Sample | 0 | off | off | 8908.3164 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9324 | 4.2937 | 2.1422 | 79.9555 | -23.0355 | -1.8419 | 0 | 0 | 0 | 0 | 0 | 3.7627 | FALSE | 0.05275 | 0.089991
3144 | 3145 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFNS | 1.14878994941108 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.85 | 8.45 | Baseline | Baseline | 9695.7674 | 38078.6403 | 0.25462 | 1930.1454 | 7713.0912 | 0.25024 | 5.0233 | 4.9369 | 1930.1478 | 7714.5359 | 0.99382 | 0.99487 | 8.8758 | 8.482 | 0.29137 | 0.025786 | 8.6886 | 8.3502 | 9.0963 | 8.7142 | 0.4077 | 0.364 | 0.078249 | 0.07689 | 109.7701 | 128.7496 | 2.758e-06 | 5.7172e-05 | FALSE | 1.0464 | TRUE | [M-H]- | 114.879 | 8.7963 | 8.9678 | 0.17151 | 8.8085 | 8.9528 | 0.14427 | 0.013057 | 1.166 | 1.351 | 19.0015 | 0 | 21 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9695.7674 | No Control Sample | 0 | off | off | 9695.7674 | 0 | off | 0 | off | 0 | off | 0 | off | 548.929 | 3.5874 | 1.9692 | 79.957 | -1.7167 | -0.13727 | 0 | 0 | 0 | 0 | 0 | 1.7959 | FALSE | 0.25 | 0.25462
3146 | 3147 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFDS | 1.16832869501221 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.19 | 8.45 | Baseline | Baseline | 10131.429 | 38078.6403 | 0.26607 | 2121.8021 | 7713.0912 | 0.27509 | 4.7749 | 4.9369 | 2123.5866 | 7714.5359 | 0.98595 | 0.99487 | 9.2044 | 8.482 | 0.15702 | 0.01443 | 9.0972 | 8.3502 | 9.3674 | 8.7142 | 0.27022 | 0.364 | 0.075412 | 0.07689 | 201.5408 | 128.7496 | 0.00038863 | 5.7172e-05 | FALSE | 1.0852 | TRUE | [M-H]- | 116.8329 | 9.1317 | 9.2827 | 0.15102 | 9.1399 | 9.272 | 0.13211 | -3.0516 | 1.2016 | 1.4154 | 15.9354 | 0 | 15 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10131.429 | No Control Sample | 0 | off | off | 10131.429 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9258 | 3.3209 | 1.989 | 79.957 | 1.0269 | 0.082111 | 0 | 0 | 0 | 0 | 0 | 1.2443 | FALSE | 0.25 | 0.26607
3150 | 3151 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FBSA | 0.835615156238452 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 4.6 | 8.86 | Manual | Baseline | 16001.248 | 72623.0597 | 0.22033 | 1513.4805 | 15352.4881 | 0.098582 | 10.5725 | 4.7304 | 1521.1243 | 15352.4881 | 0.9729 | 0.99839 | 4.9679 | 9.0219 | 7.9976 | 0.36789 | 4.7381 | 8.7533 | 5.2088 | 9.3002 | 0.47072 | 0.54685 | 0.17258 | 0.073801 | 139.8055 | 113.8663 | 0 | 1.9447e-09 | TRUE | 0.55065 | TRUE | [M-H]- | 83.5615 | 4.7731 | 5.165 | 0.3919 | 4.7953 | 5.1451 | 0.34983 | 0 | 0.96735 | 0.93914 | 5.8161 | 0 | 85 | 32 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 16001.248 | No Control Sample | 0 | off | off | 16001.248 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9605 | 5.1749 | 1.5419 | 77.9649 | -4.9227 | -0.3838 | 0 | 0 | 0 | 0 | 0 | 0.1962 | FALSE | 0.25 | 0.22033
3152 | 3153 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FPeSA | 1.07191855690552 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 7 | 8.86 | Baseline | Baseline | 31875.4021 | 72623.0597 | 0.43892 | 7006.8884 | 15352.4881 | 0.4564 | 4.5492 | 4.7304 | 7033.1237 | 15352.4881 | 0.9795 | 0.99839 | 7.114 | 9.0219 | 1.628 | 0.11396 | 6.9677 | 8.7533 | 7.2261 | 9.3002 | 0.25843 | 0.54685 | 0.070594 | 0.073801 | 229.5018 | 113.8663 | 0.0063893 | 1.9447e-09 | FALSE | 0.78852 | TRUE | [M-H]- | 107.1919 | 7.0364 | 7.1949 | 0.15851 | 7.0468 | 7.1841 | 0.13734 | 173.2329 | 1.0429 | 1.0935 | 3.8557 | 0 | 28 | 9 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 31875.4021 | No Control Sample | 0 | off | off | 31875.4021 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9576 | 5.2383 | 1.8227 | 77.9651 | -2.3924 | -0.18653 | 0 | 0 | 0 | 0 | 0 | 1.1054 | FALSE | 0.25 | 0.43892
3154 | 3155 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FHxSA | 1.10938590893279 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 7.86 | 8.86 | Baseline | Baseline | 16462.5048 | 72623.0597 | 0.22668 | 3640.6412 | 15352.4881 | 0.23714 | 4.5219 | 4.7304 | 3645.6498 | 15352.4881 | 0.99468 | 0.99839 | 8.0046 | 9.0219 | 1.8402 | 0.14464 | 7.8905 | 8.7533 | 8.2097 | 9.3002 | 0.31925 | 0.54685 | 0.070806 | 0.073801 | 110.9058 | 113.8663 | 0.0020531 | 1.9447e-09 | FALSE | 0.88725 | TRUE | [M-H]- | 110.9386 | 7.932 | 8.0811 | 0.14908 | 7.943 | 8.0718 | 0.12878 | -23.4136 | 1.1024 | 1.2753 | 9.6986 | 0 | 23 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 16462.5048 | No Control Sample | 0 | off | off | 16462.5048 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9546 | 5.0056 | 1.992 | 77.9651 | -4.259 | -0.33206 | 0 | 0 | 0 | 0 | 0 | 1.4038 | FALSE | 0.25 | 0.22668
3156 | 3157 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FHpSA | 1.04916394596345 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.58 | 8.86 | Baseline | Baseline | 32247.3622 | 72623.0597 | 0.44404 | 7352.5631 | 15352.4881 | 0.47892 | 4.3859 | 4.7304 | 7352.6633 | 15352.4881 | 0.99639 | 0.99839 | 8.5807 | 9.0219 | 0.0084608 | 0.00072594 | 8.3407 | 8.7533 | 8.7835 | 9.3002 | 0.44278 | 0.54685 | 0.068276 | 0.073801 | 185.6084 | 113.8663 | 2.4115e-05 | 1.9447e-09 | FALSE | 0.9511 | TRUE | [M-H]- | 104.9164 | 8.5099 | 8.6567 | 0.14674 | 8.5187 | 8.6462 | 0.12757 | 0.40043 | 0.98915 | 0.94972 | 4.2151 | 0 | 25 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 32247.3622 | No Control Sample | 0 | off | off | 32247.3622 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9515 | 4.6455 | 2.0809 | 77.9648 | -7.5556 | -0.58907 | 0 | 0 | 0 | 0 | 0 | 1.0231 | FALSE | 0.25 | 0.44404
3158 | 3159 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FOSA | 1.06810254858959 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.88 | 8.86 | Baseline | Baseline | 17649.5227 | 72623.0597 | 0.24303 | 3855.2487 | 15352.4881 | 0.25112 | 4.5781 | 4.7304 | 3859.3327 | 15352.4881 | 0.98275 | 0.99839 | 9.0203 | 9.0219 | 1.5801 | 0.14031 | 8.889 | 8.7533 | 9.1337 | 9.3002 | 0.24473 | 0.54685 | 0.071952 | 0.073801 | 140.0383 | 113.8663 | 0.0018246 | 1.9447e-09 | FALSE | 0.99983 | TRUE | [M-H]- | 106.8103 | 8.9407 | 9.0939 | 0.15321 | 8.9534 | 9.0873 | 0.13391 | 28.7431 | 0.97797 | 1.0379 | 7.4832 | 0 | 13 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 17649.5227 | No Control Sample | 0 | off | off | 17649.5227 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9486 | 4.8758 | 2.4279 | 77.9651 | -3.6388 | -0.2837 | 0 | 0 | 0 | 0 | 0 | 1.1451 | FALSE | 0.25 | 0.24303
3177 | 3178 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 4:2 FTS | 1.11631925090528 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 5.94 | 5.92 | Baseline | Baseline | 10304.3259 | 18179.4988 | 0.56681 | 1959.6562 | 3282.7344 | 0.59696 | 5.2582 | 5.5379 | 1960.6805 | 3300.0668 | 0.99475 | 0.97833 | 5.9555 | 5.9536 | 0.26087 | 0.015496 | 5.8168 | 5.8026 | 6.1161 | 6.0812 | 0.29928 | 0.27857 | 0.075872 | 0.079622 | 155.4197 | 171.6512 | 0.0011404 | 5.2196e-06 | FALSE | 1.0003 | TRUE | [M-H]- | 111.6319 | 5.8552 | 6.0555 | 0.2003 | 5.8735 | 6.0399 | 0.16635 | 7.4674 | 1.0361 | 1.1232 | 4.4882 | 0 | 35 | 11 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10304.3259 | No Control Sample | 0 | off | off | 10304.3259 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9765 | 6.7004 | 2.1908 | 306.9688 | 3.9342 | 1.2077 | 0 | 0 | 0 | 0 | 0 | 9.7578 | FALSE | 0.25 | 0.56681
3179 | 3180 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 6:2 FTS | 0.97914216218427 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 7.92 | 7.92 | Baseline | Baseline | 9912.4618 | 16685.6551 | 0.59407 | 2250.5353 | 4018.0061 | 0.56011 | 4.4045 | 4.1527 | 2250.5356 | 4018.0061 | 0.99962 | 0.99871 | 7.9424 | 7.9416 | 0.28245 | 0.02237 | 7.7723 | 7.7344 | 8.1178 | 8.2033 | 0.3455 | 0.4689 | 0.067769 | 0.063003 | 91.3106 | 95.5967 | 2.6225e-07 | 4.5469e-14 | FALSE | 1.0001 | TRUE | [M-H]- | 97.9142 | 7.8706 | 8.0179 | 0.14721 | 7.8806 | 8.0051 | 0.12446 | 0.0017083 | 1.0187 | 0.99822 | 5.1228 | 0 | 26 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9912.4618 | No Control Sample | 0 | off | off | 9912.4618 | 0 | off | 0 | off | 0 | off | 0 | off | 426.97 | 4.9884 | 2.1299 | 406.9616 | 0.17642 | 0.071797 | 0 | 0 | 0 | 0 | 0 | 5.5018 | FALSE | 0.25 | 0.59407
3181 | 3182 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 8:2 FTS | 1.1559855716252 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.88 | 8.89 | Baseline | Baseline | 10045.089 | 17710.3758 | 0.56719 | 1932.5782 | 3430.9978 | 0.56327 | 5.1978 | 5.1619 | 1932.5782 | 3432.1268 | 1 | 0.99491 | 8.8869 | 8.8885 | 0.077883 | 0.006916 | 8.654 | 8.7359 | 9.1572 | 9.0494 | 0.50325 | 0.31348 | 0.080654 | 0.081219 | 115.3167 | 131.684 | 0 | 0.00070858 | FALSE | 0.99982 | TRUE | [M-H]- | 115.5986 | 8.8084 | 8.9815 | 0.17303 | 8.8149 | 8.969 | 0.15401 | 0 | 1.0293 | 0.98621 | 5.0365 | 0 | 26 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Fail | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10045.089 | No Control Sample | 0 | off | off | 10045.089 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9644 | 5.4166 | 2.8543 | 506.9572 | 2.8833 | 1.4617 | 0 | 0 | 0 | 0 | 0 | 8.2314 | FALSE | 0.25 | 0.56719
3183 | 3184 | PFAS CS5 1ng/ml | 23 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 24) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 10:2 FTS | 0.812569033782957 | Quantifiers | 10/27/2023 01:12:53 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.5 | 8.89 | Baseline | Baseline | 7350.5711 | 17710.3758 | 0.41504 | 1884.4696 | 3430.9978 | 0.54925 | 3.9006 | 5.1619 | 1893.045 | 3432.1268 | 0.97975 | 0.99491 | 9.5029 | 8.8885 | 0.030613 | 0.0029082 | 9.4023 | 8.7359 | 9.601 | 9.0494 | 0.19877 | 0.31348 | 0.06102 | 0.081219 | 260.9706 | 131.684 | 0.0089388 | 0.00070858 | FALSE | 1.0691 | TRUE | [M-H]- | 81.2569 | 9.4404 | 9.5717 | 0.13135 | 9.4487 | 9.5655 | 0.11676 | 84.7471 | 1.1002 | 1.2738 | 5.0149 | 0 | 14 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Fail | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 7350.5711 | No Control Sample | 0 | off | off | 7350.5711 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9575 | 3.7596 | 2.3571 | 606.9491 | 2.052 | 1.2455 | 0 | 0 | 0 | 0 | 0 | 2.7696 | FALSE | 0.25 | 0.41504
3244 | 3245 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFBA | 1.08564831904169 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 1.49 | 1.48 | Manual | Baseline | 20653.1314 | 106055.9927 | 0.19474 | 3014.0046 | 15662.0008 | 0.19244 | 6.8524 | 6.7715 | 3044.4962 | 15709.0786 | 0.9253 | 0.98412 | 1.4434 | 1.4389 | 3.1271 | 0.046594 | 1.333 | 1.3138 | 1.633 | 1.6682 | 0.29993 | 0.35443 | 0.10649 | 0.10179 | 108.7296 | 246.4451 | 0 | 0.0060809 | TRUE | 1.0031 | TRUE | [M-H]- | 108.5648 | 1.3493 | 1.5993 | 0.25 | 1.3608 | 1.5784 | 0.21757 | 0 | 1.4741 | 1.9669 | 1.5968 | 0 | 61 | 23 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 20653.1314 | No Control Sample | 0 | off | off | 20653.1314 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9806 | 6.625 | 1.411 | 168.9891 | 0.67001 | 0.11322 | 0 | 0 | 0 | 0 | 0 | 5.3469 | FALSE | 0.25 | 0.19474
3246 | 3247 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFPeA | 1.04377176558151 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 3.18 | 3.21 | Manual | Manual | 17749.4101 | 76037.9407 | 0.23343 | 1957.9599 | 8689.1213 | 0.22533 | 9.0653 | 8.7509 | 1967.4259 | 8715.4882 | 0.96245 | 0.98644 | 3.1699 | 3.1614 | 0.31732 | 0.010091 | 2.9768 | 2.9596 | 3.4564 | 3.4562 | 0.47953 | 0.49664 | 0.13376 | 0.138 | 247.3421 | 578.4874 | 0.0035524 | 0.001925 | TRUE | 1.0027 | TRUE | [M-H]- | 104.3772 | 3.0325 | 3.3574 | 0.32494 | 3.0441 | 3.3241 | 0.28 | -14.5047 | 1.2408 | 1.3475 | 1.9454 | 0 | 85 | 25 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 17749.4101 | No Control Sample | 0 | off | off | 17749.4101 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9775 | 5.6275 | 1.4799 | 218.9865 | 3.4187 | 0.74865 | 0 | 0 | 0 | 0 | 0 | 1.1588 | FALSE | 0.25 | 0.23343
3248 | 3249 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFHxA | 1.05570728540275 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 18.5529 | 6.13 | 6.12 | Baseline | Baseline | 37818.9183 | 153927.9712 | 0.24569 | 7112.6512 | 28006.0363 | 0.25397 | 5.3171 | 5.4962 | 7125.4223 | 28016.7832 | 0.98422 | 0.99618 | 6.1028 | 6.1015 | 0.44427 | 0.027234 | 5.9245 | 5.9102 | 6.2448 | 6.3082 | 0.3203 | 0.39798 | 0.081195 | 0.083023 | 213.437 | 230.7318 | 0.0028647 | 0.00080617 | FALSE | 1.0002 | TRUE | [M-H]- | 105.5707 | 6.0109 | 6.1931 | 0.1822 | 6.023 | 6.182 | 0.15897 | 63.6131 | 0.97914 | 0.9649 | 1.4817 | 0 | 36 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 37818.9183 | No Control Sample | 203843.5852 | off | off | 37818.9183 | 203843.5852 | off | 1 | off | 1 | off | 0.18553 | off | 312.9747 | 6.0015 | 1.8783 | 268.9837 | 4.7286 | 1.2719 | 0 | 0 | 0 | 0 | 0 | 157.5172 | FALSE | 0.25 | 0.24569
3250 | 3251 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFHpA | 1.20326515220529 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 21.672 | 7.26 | 7.27 | Baseline | Baseline | 49483.3636 | 164349.8575 | 0.30109 | 11889.8564 | 38442.4243 | 0.30929 | 4.1618 | 4.2752 | 11906.8225 | 38462.2338 | 0.9879 | 0.99588 | 7.2572 | 7.2587 | 0.038782 | 0.0028156 | 7.1113 | 7.1139 | 7.3806 | 7.4177 | 0.26927 | 0.30373 | 0.064043 | 0.066335 | 244.522 | 259.1212 | 0.0026879 | 0.00076586 | FALSE | 0.99979 | TRUE | [M-H]- | 120.3265 | 7.1888 | 7.3295 | 0.14076 | 7.1978 | 7.3201 | 0.12233 | 118.6896 | 1.1014 | 1.2269 | 1.5135 | 0 | 31 | 10 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 49483.3636 | No Control Sample | 228328.7703 | off | off | 49483.3636 | 228328.7703 | off | 1 | off | 1 | off | 0.21672 | off | 362.9718 | 5.8735 | 2.1319 | 318.9801 | 3.2333 | 1.0313 | 0 | 0 | 0 | 0 | 0 | 0.8458 | FALSE | 0.25 | 0.30109
3252 | 3253 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFOA | 1.09910925183314 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 23.1853 | 7.96 | 7.94 | Baseline | Baseline | 44394.2151 | 166546.9946 | 0.26656 | 11127.7477 | 38981.8207 | 0.28546 | 3.9895 | 4.2724 | 11168.5339 | 38990.5048 | 0.97239 | 0.99687 | 7.9505 | 7.9517 | 0.11955 | 0.0095165 | 7.8218 | 7.7894 | 8.096 | 8.1668 | 0.2742 | 0.3774 | 0.06067 | 0.064493 | 110.0712 | 103.4718 | 0.0011746 | 0.00049854 | FALSE | 0.99985 | TRUE | [M-H]- | 109.9109 | 7.8838 | 8.0267 | 0.14291 | 7.8935 | 8.0138 | 0.12031 | 47.669 | 1.0821 | 1.137 | 1.6124 | 0 | 21 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 44394.2151 | No Control Sample | 191475.4056 | off | off | 44394.2151 | 191475.4056 | off | 1 | off | 1 | off | 0.23185 | off | 412.9687 | 5.6249 | 2.3229 | 368.9777 | 4.3215 | 1.5945 | 0 | 0 | 0 | 0 | 0 | 0.96057 | FALSE | 0.25 | 0.26656
3254 | 3255 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFNA | 0.992312968031163 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 20.279 | 8.46 | 8.47 | Baseline | Baseline | 32272.6089 | 139998.0841 | 0.23052 | 6405.1749 | 28961.2115 | 0.22116 | 5.0385 | 4.834 | 6416.3377 | 28961.306 | 0.98308 | 0.9983 | 8.4671 | 8.4656 | 0.083633 | 0.0070754 | 8.3373 | 8.2957 | 8.6153 | 8.655 | 0.27802 | 0.35933 | 0.079774 | 0.076082 | 159.1442 | 162.3642 | 0.00045335 | 6.4906e-06 | FALSE | 1.0002 | TRUE | [M-H]- | 99.2313 | 8.3877 | 8.5466 | 0.1589 | 8.3966 | 8.5381 | 0.14157 | 10.4446 | 0.90294 | 0.78924 | 1.5295 | 0 | 16 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 32272.6089 | No Control Sample | 159143.3885 | off | off | 32272.6089 | 159143.3885 | off | 1 | off | 1 | off | 0.20279 | off | 462.9652 | 4.2828 | 1.9828 | 418.9742 | 2.431 | 1.0185 | 0 | 0 | 0 | 0 | 0 | 2.8837 | FALSE | 0.25 | 0.23052
3256 | 3257 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFDA | 0.984716522148336 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.87 | 8.86 | Baseline | Baseline | 30804.0545 | 126013.4546 | 0.24445 | 6214.3976 | 26148.8777 | 0.23765 | 4.9569 | 4.8191 | 6221.8422 | 26148.8788 | 0.98952 | 0.99967 | 8.8828 | 8.8803 | 0.14486 | 0.012849 | 8.6662 | 8.6741 | 9.0338 | 9.1581 | 0.36753 | 0.48407 | 0.076958 | 0.075338 | 116.2548 | 173.6288 | 0.0019853 | 1.0348e-07 | FALSE | 1.0003 | TRUE | [M-H]- | 98.4717 | 8.8023 | 8.9682 | 0.16587 | 8.8098 | 8.9563 | 0.14658 | 33.5681 | 0.99841 | 0.9388 | 1.8628 | 0 | 19 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 30804.0545 | No Control Sample | 0 | off | off | 30804.0545 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9624 | 4.6107 | 2.3651 | 468.9715 | 3.1987 | 1.5001 | 0 | 0 | 0 | 0 | 0 | 91.9307 | FALSE | 0.25 | 0.24445
3258 | 3259 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFUdA | 1.26518210519568 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.21 | 9.2 | Baseline | Baseline | 26423.1077 | 94174.7869 | 0.28058 | 5575.762 | 20270.9034 | 0.27506 | 4.7389 | 4.6458 | 5575.8007 | 20270.9052 | 0.99737 | 0.99909 | 9.2165 | 9.2183 | 0.07061 | 0.0065032 | 9.0534 | 9.0049 | 9.4292 | 9.4842 | 0.37582 | 0.4793 | 0.074868 | 0.072979 | 239.8925 | 197.3637 | 1.55e-05 | 1.9639e-07 | FALSE | 0.9998 | TRUE | [M-H]- | 126.5182 | 9.1407 | 9.2959 | 0.15516 | 9.1522 | 9.2857 | 0.13347 | 0.22996 | 1.1356 | 1.3464 | 1.6329 | 0 | 21 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 26423.1077 | No Control Sample | 0 | off | off | 26423.1077 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9593 | 4.3403 | 2.4434 | 518.9692 | 3.668 | 1.9036 | 0 | 0 | 0 | 0 | 0 | 2.7575 | FALSE | 0.25 | 0.28058
3260 | 3261 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFDoA | 1.00249104715676 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.47 | 9.47 | Baseline | Baseline | 20957.4467 | 80784.9433 | 0.25942 | 5117.474 | 21058.9117 | 0.24301 | 4.0953 | 3.8361 | 5119.5305 | 21068.3427 | 0.99349 | 0.99612 | 9.4857 | 9.486 | 0.16621 | 0.01574 | 9.3627 | 9.3519 | 9.5968 | 9.6455 | 0.23417 | 0.2936 | 0.064546 | 0.059837 | 247.1688 | 356.1606 | 0.00072186 | 0.00097508 | FALSE | 0.99997 | TRUE | [M-H]- | 100.2491 | 9.419 | 9.5531 | 0.1341 | 9.4293 | 9.5458 | 0.11655 | 15.7755 | 1.1089 | 1.3221 | 1.5896 | 0 | 16 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 20957.4467 | No Control Sample | 0 | off | off | 20957.4467 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9562 | 4.1916 | 2.5692 | 568.9673 | 6.8007 | 3.8693 | 0 | 0 | 0 | 0 | 0 | 382.4645 | FALSE | 0.25 | 0.25942
3262 | 3263 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFTrDA | 0.931652034526276 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.68 | 9.47 | Baseline | Baseline | 14197.0092 | 80784.9433 | 0.17574 | 3562.4469 | 21058.9117 | 0.16917 | 3.9852 | 3.8361 | 3563.0851 | 21068.3427 | 0.99113 | 0.99612 | 9.7042 | 9.486 | 0.2503 | 0.024229 | 9.5541 | 9.3519 | 9.8334 | 9.6455 | 0.2793 | 0.2936 | 0.063227 | 0.059837 | 373.6162 | 356.1606 | 0.00032648 | 0.00097508 | FALSE | 1.023 | TRUE | [M-H]- | 93.1652 | 9.6406 | 9.7676 | 0.12696 | 9.6474 | 9.7596 | 0.11215 | 4.1642 | 0.97101 | 0.91416 | 1.2061 | 0 | 24 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 14197.0092 | No Control Sample | 0 | off | off | 14197.0092 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9529 | 3.7418 | 2.4806 | 618.9634 | 2.7145 | 1.6802 | 0 | 0 | 0 | 0 | 0 | 2.9676 | FALSE | 0.25 | 0.17574
3264 | 3265 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFTeDA | 0.932674824479615 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.87 | 9.86 | Baseline | Baseline | 6840.5718 | 20686.5725 | 0.33068 | 2062.907 | 6434.3965 | 0.32061 | 3.316 | 3.215 | 2065.6721 | 6443.3142 | 0.96055 | 0.99049 | 9.8772 | 9.8766 | 0.073317 | 0.0072364 | 9.8041 | 9.7861 | 9.9623 | 9.98 | 0.15818 | 0.19387 | 0.050944 | 0.049302 | 360.6692 | 602.8734 | 9.9679e-06 | 0.0029356 | FALSE | 1.0001 | TRUE | [M-H]- | 93.2675 | 9.8265 | 9.9397 | 0.11319 | 9.8327 | 9.9316 | 0.098915 | 0.12999 | 1.1789 | 1.3645 | 1.5724 | 0 | 17 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 6840.5718 | No Control Sample | 0 | off | off | 6840.5718 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9498 | 3.4947 | 2.4915 | 668.9622 | 6.4605 | 4.3218 | 0 | 0 | 0 | 0 | 0 | 386.7349 | FALSE | 0.25 | 0.33068
3271 | 3272 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFBS | 1.06275643852894 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 3.5 | 3.5 | Manual | Manual | 11931.262 | 42414.9241 | 0.2813 | 1222.9056 | 4148.3356 | 0.29479 | 9.7565 | 10.2246 | 1228.7526 | 4157.9538 | 0.97454 | 0.98786 | 3.8056 | 3.7994 | 8.7321 | 0.30562 | 3.5588 | 3.5653 | 4.0525 | 4.0969 | 0.4937 | 0.53161 | 0.16499 | 0.15837 | 216.51 | 857.9244 | 0.0010657 | 0.00091118 | TRUE | 1.0016 | TRUE | [M-H]- | 106.2756 | 3.6378 | 4.003 | 0.36517 | 3.6641 | 3.9623 | 0.29825 | 2.6396 | 1.1322 | 1.2088 | 11.382 | 0 | 77 | 28 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 11931.262 | No Control Sample | 0 | off | off | 11931.262 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9446 | 5.5244 | 1.6515 | 79.9571 | -1.3224 | -0.10573 | 0 | 0 | 0 | 0 | 0 | 5.2009 | FALSE | 0.25 | 0.2813
3273 | 3274 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFPeS | 1.13550122810349 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 6.38 | 3.5 | Baseline | Manual | 12835.7261 | 42414.9241 | 0.30262 | 2529.8984 | 4148.3356 | 0.60986 | 5.0736 | 10.2246 | 2540.2887 | 4157.9538 | 0.97385 | 0.98786 | 6.3573 | 3.7994 | 0.35577 | 0.022698 | 6.209 | 3.5653 | 6.4759 | 4.0969 | 0.26683 | 0.53161 | 0.078123 | 0.15837 | 235.7605 | 857.9244 | 0.0066667 | 0.00091118 | FALSE | 1.6732 | TRUE | [M-H]- | 113.5501 | 6.2773 | 6.4369 | 0.1596 | 6.2888 | 6.4282 | 0.13936 | 63.2079 | 1.0474 | 1.1546 | 14.2569 | 0 | 29 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 12835.7261 | No Control Sample | 0 | off | off | 12835.7261 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9415 | 4.8716 | 1.6999 | 79.9568 | -5.49 | -0.43897 | 0 | 0 | 0 | 0 | 0 | 0.63655 | FALSE | 0.25 | 0.30262
3275 | 3276 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | L-PFHxS | 0.866613786708527 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.811 | 4 | 0.20275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 10010.9047 | 61842.0875 | 0.16188 | 2368.0136 | 15259.6697 | 0.15518 | 4.2276 | 4.0526 | 2395.8056 | 15287.7275 | 0.75039 | 0.98535 | 7.3348 | 7.3336 | 0.071499 | 0.005248 | 7.2458 | 7.2191 | 7.434 | 7.4528 | 0.18815 | 0.23372 | 0.067926 | 0.063064 | 87.8176 | 203.4825 | 0.0063002 | 0.00312 | FALSE | 1.0002 | TRUE | [M-H]- | 106.8574 | 7.2674 | 7.4053 | 0.13789 | 7.2749 | 7.3963 | 0.12145 | -79.2929 | 1.0745 | 1.1417 | 15.3474 | 0 | 23 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10010.9047 | No Control Sample | 0 | off | off | 10010.9047 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9388 | 5.3784 | 2.1457 | 79.957 | -2.8235 | -0.22576 | 0 | 0 | 0 | 0 | 0 | 28.9036 | FALSE | 0.20275 | 0.16188
3277 | 3278 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | Br-PFHxS | 0.177936689383399 | Qualifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 0.189 | 4 | 0.04725 | 0 | 7.16 | 7.35 | Manual | Baseline | 2792.6534 | 61842.0875 | 0.045158 | 503.0496 | 15259.6697 | 0.032966 | 5.5514 | 4.0526 | 2395.8056 | 15287.7275 | 0.20933 | 0.98535 | 7.1518 | 7.3336 | 0.11399 | 0.0081614 | 7.0314 | 7.2191 | 7.2417 | 7.4528 | 0.21022 | 0.23372 | 0.087658 | 0.063064 | 18.6555 | 203.4825 | 0.013773 | 0.00312 | TRUE | 0.97522 | TRUE | [M-H]- | 94.1464 | 7.0619 | 7.2458 | 0.18389 | 7.0674 | 7.2234 | 0.15595 | 32.9585 | 1.0324 | 0.86662 | 0.27896 | 0 | 24 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2792.6534 | No Control Sample | 0 | off | off | 10010.9047 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9388 | 5.3784 | 2.1457 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 28.9036 | FALSE | 0.04725 | 0.045158
3279 | 3280 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFHpS | 1.14387355274834 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Total Width | 0 | 0 | 1 | 4 | 0.25 | 0 | 7.98 | 7.35 | Baseline | Baseline | 10766.1232 | 61842.0875 | 0.17409 | 2650.8037 | 15259.6697 | 0.17371 | 4.0615 | 4.0526 | 2651.5163 | 15287.7275 | 0.99587 | 0.98535 | 7.9744 | 7.3336 | 0.069736 | 0.0055649 | 7.8643 | 7.2191 | 8.0994 | 7.4528 | 0.2351 | 0.23372 | 0.062988 | 0.063064 | 132.6729 | 203.4825 | 0.00053267 | 0.00312 | FALSE | 1.0874 | TRUE | [M-H]- | 114.3874 | 7.9115 | 8.0502 | 0.13865 | 7.9197 | 8.0404 | 0.12071 | 6.006 | 1.0758 | 1.1435 | 19.2035 | 0 | 18 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10766.1232 | No Control Sample | 0 | off | off | 10766.1232 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9352 | 3.9122 | 1.7563 | 79.9572 | 0.59701 | 0.047735 | 0 | 0 | 0 | 0 | 0 | 0.71388 | FALSE | 0.25 | 0.17409
3281 | 3282 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | L-PFOS | 0.743985951338721 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 0.788 | 4 | 0.197 | 0 | 8.46 | 8.45 | Manual | Baseline | 7374.5727 | 38134.3896 | 0.19338 | 1500.6914 | 7968.9266 | 0.18832 | 4.9141 | 4.7854 | 1515.9269 | 7973.004 | 0.64712 | 0.99257 | 8.4632 | 8.4659 | 0.038185 | 0.0032305 | 8.3848 | 8.3129 | 8.5642 | 8.6416 | 0.17946 | 0.3287 | 0.079396 | 0.075424 | 42.5125 | 151.1981 | 0 | 0.00048315 | TRUE | 0.99968 | TRUE | [M-H]- | 94.4145 | 8.3789 | 8.5425 | 0.16354 | 8.3789 | 8.5318 | 0.15281 | 0 | 1.0207 | 0.90732 | 15.8583 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 7374.5727 | No Control Sample | 0 | off | off | 7374.5727 | 0 | off | 0 | off | 0 | off | 0 | off | 498.932 | 3.6579 | 1.825 | 79.9574 | 1.3278 | 0.10617 | 0 | 0 | 0 | 0 | 0 | 0.7478 | FALSE | 0.197 | 0.19338
3283 | 3284 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | Br-PFOS | 0.159677548243149 | Qualifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.211 | 4 | 0.05275 | 0 | 8.3 | 8.45 | Manual | Baseline | 3388.4261 | 38134.3896 | 0.088855 | 642.492 | 7968.9266 | 0.080625 | 5.2739 | 4.7854 | 1515.9269 | 7973.004 | 0.29764 | 0.99257 | 8.3133 | 8.4659 | 0.1605 | 0.013322 | 8.2028 | 8.3129 | 8.3865 | 8.6416 | 0.18369 | 0.3287 | 0.083821 | 0.075424 | 18.2009 | 151.1981 | 0.01125 | 0.00048315 | TRUE | 0.98198 | TRUE | [M-H]- | 75.6766 | 8.2299 | 8.3976 | 0.16771 | 8.2404 | 8.3976 | 0.15722 | 39.348 | 1.0884 | 1.3623 | 0.45947 | 0 | 13 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3388.4261 | No Control Sample | 0 | off | off | 7374.5727 | 0 | off | 0 | off | 0 | off | 0 | off | 498.932 | 3.6579 | 1.825 | 79.9563 | -12.871 | -1.0291 | 0 | 0 | 0 | 0 | 0 | 0.7478 | FALSE | 0.05275 | 0.088855
3285 | 3286 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFNS | 1.11971681217205 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.85 | 8.45 | Baseline | Baseline | 9464.2265 | 38134.3896 | 0.24818 | 1992.8192 | 7968.9266 | 0.25007 | 4.7492 | 4.7854 | 1992.8205 | 7973.004 | 0.99505 | 0.99257 | 8.8675 | 8.4659 | 0.1977 | 0.017497 | 8.6879 | 8.3129 | 9.0569 | 8.6416 | 0.36897 | 0.3287 | 0.073918 | 0.075424 | 119.2991 | 151.1981 | 1.3879e-06 | 0.00048315 | FALSE | 1.0474 | TRUE | [M-H]- | 111.9717 | 8.7901 | 8.9509 | 0.16077 | 8.8006 | 8.9407 | 0.14009 | 0.0074961 | 1.0208 | 1.0505 | 19.5923 | 0 | 19 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9464.2265 | No Control Sample | 0 | off | off | 9464.2265 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9291 | 3.75 | 2.0585 | 79.9571 | 0.23863 | 0.01908 | 0 | 0 | 0 | 0 | 0 | 1.6948 | FALSE | 0.25 | 0.24818
3287 | 3288 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | PFDS | 1.04162339538948 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.19 | 8.45 | Baseline | Baseline | 9045.8994 | 38134.3896 | 0.23721 | 1926.1137 | 7968.9266 | 0.2417 | 4.6965 | 4.7854 | 1926.1137 | 7973.004 | 0.9886 | 0.99257 | 9.2003 | 8.4659 | 0.11155 | 0.010252 | 9.0204 | 8.3129 | 9.4665 | 8.6416 | 0.4461 | 0.3287 | 0.073503 | 0.075424 | 133.7527 | 151.1981 | 1.2694e-10 | 0.00048315 | FALSE | 1.0867 | TRUE | [M-H]- | 104.1623 | 9.1285 | 9.2868 | 0.15831 | 9.1379 | 9.2725 | 0.13456 | -5.4808e-07 | 1.2195 | 1.4252 | 17.6455 | 0 | 25 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9045.8994 | No Control Sample | 0 | off | off | 9045.8994 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9257 | 3.1362 | 1.8783 | 79.9574 | 6.8071 | 0.54428 | 0 | 0 | 0 | 0 | 0 | 1.4212 | FALSE | 0.25 | 0.23721
3291 | 3292 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FBSA | 0.845820311586346 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 4.6 | 8.86 | Manual | Baseline | 16135.2266 | 72347.5705 | 0.22302 | 1521.0592 | 15505.9296 | 0.098095 | 10.6079 | 4.6658 | 1536.6627 | 15507.5278 | 0.95923 | 0.99738 | 4.9196 | 9.0109 | 6.9474 | 0.31958 | 4.6854 | 8.8143 | 5.1523 | 9.219 | 0.46694 | 0.40468 | 0.17259 | 0.072975 | 126.4967 | 146.7424 | 0.010153 | 0.00020321 | TRUE | 0.54596 | TRUE | [M-H]- | 84.582 | 4.705 | 5.1342 | 0.42921 | 4.758 | 5.0968 | 0.33878 | 33.0747 | 0.98137 | 1.0451 | 5.7925 | 0 | 85 | 32 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 16135.2266 | No Control Sample | 0 | off | off | 16135.2266 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9608 | 5.9716 | 1.7793 | 77.9651 | -2.9352 | -0.22885 | 0 | 0 | 0 | 0 | 0 | 1.5892 | FALSE | 0.25 | 0.22302
3293 | 3294 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FPeSA | 1.12496919240442 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 7 | 8.86 | Baseline | Baseline | 33326.0562 | 72347.5705 | 0.46064 | 7090.8934 | 15505.9296 | 0.4573 | 4.6998 | 4.6658 | 7092.5288 | 15507.5278 | 0.99409 | 0.99738 | 7.0786 | 9.0109 | 1.1224 | 0.078565 | 6.9471 | 8.8143 | 7.2535 | 9.219 | 0.3064 | 0.40468 | 0.071958 | 0.072975 | 239.2841 | 146.7424 | 0.00045492 | 0.00020321 | FALSE | 0.78556 | TRUE | [M-H]- | 112.4969 | 6.9997 | 7.1607 | 0.16097 | 7.0107 | 7.1507 | 0.13994 | 10.5281 | 0.95655 | 0.91391 | 3.693 | 0 | 33 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 33326.0562 | No Control Sample | 0 | off | off | 33326.0562 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9576 | 5.244 | 1.8247 | 77.965 | -4.0879 | -0.31871 | 0 | 0 | 0 | 0 | 0 | 0.57732 | FALSE | 0.25 | 0.46064
3295 | 3296 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FHxSA | 1.19974665039574 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 7.86 | 8.86 | Baseline | Baseline | 17735.8589 | 72347.5705 | 0.24515 | 4013.1311 | 15505.9296 | 0.25881 | 4.4195 | 4.6658 | 4014.6572 | 15507.5278 | 0.99384 | 0.99738 | 7.9756 | 9.0109 | 1.4702 | 0.11555 | 7.8226 | 8.8143 | 8.1453 | 9.219 | 0.32268 | 0.40468 | 0.068112 | 0.072975 | 100.2257 | 146.7424 | 0.00080656 | 0.00020321 | FALSE | 0.8851 | TRUE | [M-H]- | 119.9747 | 7.907 | 8.0589 | 0.15193 | 7.9167 | 8.0471 | 0.13047 | 10.031 | 1.1292 | 1.2666 | 9.0866 | 0 | 24 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 17735.8589 | No Control Sample | 0 | off | off | 17735.8589 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9544 | 4.6271 | 1.8414 | 77.9651 | -3.5475 | -0.27659 | 0 | 0 | 0 | 0 | 0 | 0.43669 | FALSE | 0.25 | 0.24515
3297 | 3298 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FHpSA | 1.06581852251699 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.58 | 8.86 | Baseline | Baseline | 32634.992 | 72347.5705 | 0.45109 | 7394.2303 | 15505.9296 | 0.47686 | 4.4136 | 4.6658 | 7395.6931 | 15507.5278 | 0.9928 | 0.99738 | 8.564 | 9.0109 | 0.18624 | 0.015979 | 8.3772 | 8.8143 | 8.7044 | 9.219 | 0.32718 | 0.40468 | 0.069036 | 0.072975 | 197.5035 | 146.7424 | 0.00032165 | 0.00020321 | FALSE | 0.95041 | TRUE | [M-H]- | 106.5819 | 8.4907 | 8.6395 | 0.14888 | 8.5008 | 8.6293 | 0.1285 | 7.2692 | 0.98511 | 0.96482 | 4.272 | 0 | 19 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 32634.992 | No Control Sample | 0 | off | off | 32634.992 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9513 | 4.1455 | 1.857 | 77.9649 | -6.2301 | -0.48573 | 0 | 0 | 0 | 0 | 0 | 1.3447 | FALSE | 0.25 | 0.45109
3299 | 3300 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | FOSA | 1.14235514717869 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.88 | 8.86 | Baseline | Baseline | 18804.88 | 72347.5705 | 0.25992 | 3923.1835 | 15505.9296 | 0.25301 | 4.7933 | 4.6658 | 3926.9508 | 15507.5278 | 0.98807 | 0.99738 | 9.0123 | 9.0109 | 1.49 | 0.13231 | 8.8671 | 8.8143 | 9.1525 | 9.219 | 0.28542 | 0.40468 | 0.075129 | 0.072975 | 112.0852 | 146.7424 | 0.0018051 | 0.00020321 | FALSE | 1.0002 | TRUE | [M-H]- | 114.2355 | 8.9314 | 9.0892 | 0.15779 | 8.9389 | 9.0791 | 0.14027 | 24.8116 | 0.91357 | 0.77702 | 7.2055 | 0 | 15 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 18804.88 | No Control Sample | 0 | off | off | 18804.88 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9483 | 4.2082 | 2.0955 | 77.9651 | -3.4337 | -0.26771 | 0 | 0 | 0 | 0 | 0 | 2.6748 | FALSE | 0.25 | 0.25992
3318 | 3319 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 4:2 FTS | 1.05838266630879 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 5.94 | 5.92 | Baseline | Baseline | 10067.8864 | 18734.682 | 0.53739 | 1787.9962 | 3133.8035 | 0.57055 | 5.6308 | 5.9783 | 1797.0771 | 3142.0564 | 0.9807 | 0.98768 | 5.909 | 5.9076 | 0.52188 | 0.031 | 5.7465 | 5.7739 | 6.0501 | 6.1186 | 0.30358 | 0.34472 | 0.087532 | 0.094865 | 149.0628 | 148.1264 | 0.0093774 | 0.0041291 | FALSE | 1.0002 | TRUE | [M-H]- | 105.8383 | 5.8071 | 6.0027 | 0.19554 | 5.8235 | 5.9901 | 0.16658 | 55.2293 | 0.9567 | 0.94096 | 4.6541 | 0 | 38 | 13 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 10067.8864 | No Control Sample | 0 | off | off | 10067.8864 | 0 | off | 0 | off | 0 | off | 0 | off | 326.976 | 5.1337 | 1.6786 | 306.968 | 1.1621 | 0.35674 | 0 | 0 | 0 | 0 | 0 | 6.5605 | FALSE | 0.25 | 0.53739
3320 | 3321 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 6:2 FTS | 1.09613306830366 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 7.92 | 7.92 | Baseline | Baseline | 9932.404 | 14934.7715 | 0.66505 | 2262.6992 | 3647.9459 | 0.62027 | 4.3896 | 4.094 | 2262.6992 | 3647.9472 | 0.99791 | 0.99872 | 7.9118 | 7.911 | 0.10375 | 0.0082172 | 7.7139 | 7.7485 | 8.1647 | 8.0913 | 0.45077 | 0.3428 | 0.068574 | 0.062606 | 88.6741 | 93.4268 | 0 | 7.5832e-07 | FALSE | 1.0001 | TRUE | [M-H]- | 109.6133 | 7.8414 | 7.9853 | 0.14396 | 7.8526 | 7.9763 | 0.12367 | 0 | 1.0994 | 1.282 | 5.0584 | 0 | 33 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9932.404 | No Control Sample | 0 | off | off | 9932.404 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9699 | 4.6334 | 1.9783 | 406.9626 | 2.6998 | 1.0987 | 0 | 0 | 0 | 0 | 0 | 5.5018 | FALSE | 0.25 | 0.66505
3322 | 3323 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 8:2 FTS | 1.05673076653639 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 1 | 4 | 0.25 | 0 | 8.88 | 8.89 | Baseline | Baseline | 9864.1063 | 19024.7866 | 0.51849 | 2012.7124 | 3822.2556 | 0.52658 | 4.9009 | 4.9774 | 2012.7124 | 3823.1154 | 0.99762 | 0.99589 | 8.8752 | 8.8783 | 0.053588 | 0.0047586 | 8.6706 | 8.697 | 9.0586 | 9.0295 | 0.38803 | 0.33248 | 0.077052 | 0.076702 | 165.1546 | 104.1624 | 3.0366e-08 | 0.00040804 | FALSE | 0.99966 | TRUE | [M-H]- | 105.6731 | 8.7968 | 8.9627 | 0.16586 | 8.8091 | 8.9509 | 0.14186 | 0.00015751 | 1.1238 | 1.3052 | 5.0315 | 0 | 20 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9864.1063 | No Control Sample | 0 | off | off | 9864.1063 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9644 | 5.3775 | 2.8337 | 506.9581 | 4.7273 | 2.3965 | 0 | 0 | 0 | 0 | 0 | 5.7212 | FALSE | 0.25 | 0.51849
3324 | 3325 | PFAS CS5 1ng/ml | 24 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 25) | CS5 | 0 | 0 | Standard | 0 | 0 | 0 | 6 | 1 | 10:2 FTS | 0.705469702485115 | Quantifiers | 10/27/2023 01:29:29 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 1 | 4 | 0.25 | 0 | 9.5 | 8.89 | Baseline | Baseline | 6855.3747 | 19024.7866 | 0.36034 | 1669.2641 | 3822.2556 | 0.43672 | 4.1068 | 4.9774 | 1669.2641 | 3823.1154 | 0.99198 | 0.99589 | 9.5032 | 8.8783 | 0.033275 | 0.0031611 | 9.3147 | 8.697 | 9.6711 | 9.0295 | 0.35642 | 0.33248 | 0.065269 | 0.076702 | 212.2954 | 104.1624 | 0 | 0.00040804 | FALSE | 1.0704 | TRUE | [M-H]- | 70.547 | 9.438 | 9.5704 | 0.13239 | 9.446 | 9.5649 | 0.11887 | 0 | 1.0894 | 1.2515 | 4.8614 | 0 | 24 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 6855.3747 | No Control Sample | 0 | off | off | 6855.3747 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9576 | 3.8714 | 2.4272 | 606.9547 | 11.2369 | 6.8202 | 0 | 0 | 0 | 0 | 0 | 21.2142 | FALSE | 0.25 | 0.36034
3385 | 3386 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFBA | 3.8852303267632 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 1.49 | 1.48 | Manual | Baseline | 73337.443 | 105231.9124 | 0.69691 | 11446.4944 | 14888.1688 | 0.76883 | 6.407 | 7.0682 | 11540.6677 | 14930.3812 | 0.96558 | 0.98395 | 1.4761 | 1.479 | 0.93047 | 0.013864 | 1.3721 | 1.3292 | 1.6614 | 1.6942 | 0.28929 | 0.36498 | 0.090513 | 0.10859 | 200.5963 | 240.6353 | 0.0050675 | 0.007823 | TRUE | 0.99808 | TRUE | [M-H]- | 97.1308 | 1.3898 | 1.6249 | 0.23503 | 1.3979 | 1.6053 | 0.2074 | -200.5064 | 1.4697 | 1.8834 | 1.6949 | 0 | 59 | 20 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 73337.443 | No Control Sample | 0 | off | off | 73337.443 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9806 | 6.5312 | 1.391 | 168.9893 | 1.5811 | 0.26718 | 0 | 0 | 0 | 0 | 0 | 2.3576 | FALSE | 1 | 0.69691
3387 | 3388 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFPeA | 3.89894106258222 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 3.18 | 3.21 | Baseline | Baseline | 65316.6442 | 74908.1656 | 0.87196 | 7125.5005 | 8086.3601 | 0.88118 | 9.1666 | 9.2635 | 7226.74 | 8166.9641 | 0.9479 | 0.96579 | 3.2252 | 3.2246 | 1.4201 | 0.045159 | 3.0449 | 3.0316 | 3.4667 | 3.4619 | 0.42178 | 0.43028 | 0.14001 | 0.14357 | 494.8385 | 578.996 | 0.0040707 | 0.0039148 | FALSE | 1.0002 | TRUE | [M-H]- | 97.4735 | 3.0849 | 3.4045 | 0.31963 | 3.1028 | 3.3775 | 0.27467 | 68.7693 | 0.99672 | 0.92867 | 1.9901 | 0 | 72 | 27 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 65316.6442 | No Control Sample | 0 | off | off | 65316.6442 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9776 | 5.9409 | 1.5623 | 218.9868 | 4.5374 | 0.99362 | 0 | 0 | 0 | 0 | 0 | 1.5919 | FALSE | 1 | 0.87196
3389 | 3390 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFHxA | 4.1041688616124 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 69.3536 | 6.13 | 6.12 | Baseline | Baseline | 141372.7825 | 148010.5067 | 0.95515 | 26611.1294 | 27206.5089 | 0.97812 | 5.3125 | 5.4403 | 26702.4647 | 27240.9353 | 0.9812 | 0.99216 | 6.1571 | 6.1538 | 0.44212 | 0.027102 | 5.9925 | 5.9698 | 6.2968 | 6.3277 | 0.30432 | 0.35792 | 0.083288 | 0.086303 | 292.141 | 299.6733 | 0.0044445 | 0.0023717 | FALSE | 1.0005 | TRUE | [M-H]- | 102.6042 | 6.067 | 6.2489 | 0.18199 | 6.0788 | 6.2351 | 0.15624 | 388.65 | 0.97155 | 0.9107 | 1.4819 | 0 | 33 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 141372.7825 | No Control Sample | 203843.5852 | off | off | 141372.7825 | 203843.5852 | off | 0.25 | off | 4 | off | 2.7741 | off | 312.9743 | 4.783 | 1.497 | 268.9829 | 1.8092 | 0.48663 | 0 | 0 | 0 | 0 | 0 | 40.7795 | FALSE | 1 | 0.95515
3391 | 3392 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFHpA | 4.53208344234139 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 80.0486 | 7.26 | 7.27 | Baseline | Baseline | 182773.9766 | 161171.379 | 1.134 | 45351.4032 | 38548.2026 | 1.1765 | 4.0302 | 4.181 | 45402.1346 | 38573.8936 | 0.99177 | 0.99474 | 7.2961 | 7.2976 | 0.49682 | 0.036069 | 7.1502 | 7.1528 | 7.4409 | 7.435 | 0.29073 | 0.28223 | 0.061445 | 0.063561 | 215.4286 | 210.9314 | 0.0018647 | 0.0012199 | FALSE | 0.99979 | TRUE | [M-H]- | 113.3021 | 7.2283 | 7.3683 | 0.14001 | 7.2378 | 7.3585 | 0.12068 | 290.8716 | 1.0219 | 1.0475 | 1.5982 | 0 | 34 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Pass | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 182773.9766 | No Control Sample | 228328.7703 | off | off | 182773.9766 | 228328.7703 | off | 0.25 | off | 4 | off | 3.2019 | off | 362.9709 | 3.5795 | 1.2992 | 318.9801 | 3.2801 | 1.0463 | 0 | 0 | 0 | 0 | 0 | 0.36736 | FALSE | 1 | 1.134
3393 | 3394 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFOA | 4.38796867147886 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 90.0124 | 7.96 | 7.94 | Baseline | Baseline | 172351.6547 | 161958.3043 | 1.0642 | 40788.3905 | 38531.6153 | 1.0586 | 4.2255 | 4.2033 | 40814.1065 | 38533.0894 | 0.99251 | 0.99945 | 7.9788 | 7.9764 | 0.23583 | 0.018772 | 7.8487 | 7.8027 | 8.1426 | 8.2295 | 0.29388 | 0.42677 | 0.064908 | 0.063969 | 105.7864 | 138.5744 | 7.8973e-05 | 8.6582e-05 | FALSE | 1.0003 | TRUE | [M-H]- | 109.6992 | 7.9137 | 8.0566 | 0.14289 | 7.9217 | 8.0452 | 0.12351 | -10.9608 | 1.1955 | 1.3847 | 1.4251 | 0 | 22 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 172351.6547 | No Control Sample | 191475.4056 | off | off | 172351.6547 | 191475.4056 | off | 0.25 | off | 4 | off | 3.6005 | off | 412.968 | 3.9171 | 1.6176 | 368.9767 | 1.7005 | 0.62743 | 0 | 0 | 0 | 0 | 0 | 0.47738 | FALSE | 1 | 1.0642
3395 | 3396 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFNA | 3.77709112113977 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 76.4208 | 8.46 | 8.47 | Baseline | Baseline | 121618.675 | 138605.1891 | 0.87745 | 25062.7945 | 29637.3344 | 0.84565 | 4.8526 | 4.6767 | 25063.8232 | 29638.8233 | 0.99769 | 0.99924 | 8.4875 | 8.485 | 0.32526 | 0.027517 | 8.3187 | 8.279 | 8.7192 | 8.7079 | 0.40042 | 0.42897 | 0.075962 | 0.07259 | 142.8774 | 170.6229 | 9.2898e-05 | 0.00010364 | FALSE | 1.0003 | TRUE | [M-H]- | 94.4273 | 8.4049 | 8.5627 | 0.15778 | 8.4148 | 8.5539 | 0.13907 | 5.8147 | 0.87479 | 0.73167 | 1.4527 | 0 | 23 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 121618.675 | No Control Sample | 159143.3885 | off | off | 121618.675 | 159143.3885 | off | 0.25 | off | 4 | off | 3.0568 | off | 462.9652 | 4.1497 | 1.9212 | 418.9745 | 3.1579 | 1.3231 | 0 | 0 | 0 | 0 | 0 | 0.64753 | FALSE | 1 | 0.87745
3397 | 3398 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFDA | 3.63331633524382 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.87 | 8.86 | Baseline | Baseline | 107369.4261 | 119041.3053 | 0.90195 | 22227.2173 | 24885.5445 | 0.89318 | 4.8305 | 4.7836 | 22229.7591 | 24896.7048 | 0.99703 | 0.9967 | 8.9017 | 8.8963 | 0.35732 | 0.031694 | 8.7388 | 8.7288 | 9.1446 | 9.0747 | 0.40587 | 0.34592 | 0.074792 | 0.074514 | 141.0168 | 160.2245 | 8.9982e-05 | 0.00084719 | FALSE | 1.0006 | TRUE | [M-H]- | 90.8329 | 8.821 | 8.9845 | 0.16343 | 8.8289 | 8.9732 | 0.14429 | -4.9279 | 0.99215 | 0.93698 | 1.3857 | 0 | 21 | 5 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 107369.4261 | No Control Sample | 0 | off | off | 107369.4261 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9621 | 4.0829 | 2.0943 | 468.9718 | 3.8816 | 1.8204 | 0 | 0 | 0 | 0 | 0 | 22.1254 | FALSE | 1 | 0.90195
3399 | 3400 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFUdA | 4.19402508412904 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.21 | 9.2 | Baseline | Baseline | 90554.1697 | 97360.1469 | 0.93009 | 18914.7949 | 21493.8619 | 0.88001 | 4.7875 | 4.5297 | 18929.8479 | 21510.3579 | 0.99441 | 0.99484 | 9.235 | 9.2331 | 0.271 | 0.024959 | 9.0863 | 9.0945 | 9.3917 | 9.3811 | 0.30535 | 0.28652 | 0.075709 | 0.071019 | 220.6433 | 212.3955 | 0.0015294 | 0.0015811 | FALSE | 1.0002 | TRUE | [M-H]- | 104.8506 | 9.1596 | 9.3159 | 0.15627 | 9.1688 | 9.3027 | 0.13393 | 94.7374 | 0.93719 | 0.80407 | 1.3616 | 0 | 17 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 90554.1697 | No Control Sample | 0 | off | off | 90554.1697 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9587 | 3.3787 | 1.9021 | 518.9686 | 2.4217 | 1.2568 | 0 | 0 | 0 | 0 | 0 | 0.57425 | FALSE | 1 | 0.93009
3401 | 3402 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFDoA | 4.01704238339449 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.47 | 9.47 | Baseline | Baseline | 76542.5529 | 73632.4221 | 1.0395 | 19243.9754 | 18072.8724 | 1.0648 | 3.9775 | 4.0742 | 19243.9804 | 18079.0353 | 0.99783 | 0.99605 | 9.5045 | 9.5032 | 0.36472 | 0.034539 | 9.3409 | 9.3301 | 9.6796 | 9.6706 | 0.33865 | 0.3405 | 0.061641 | 0.063216 | 331.9422 | 205.4846 | 5.2346e-07 | 0.00068322 | FALSE | 1.0001 | TRUE | [M-H]- | 100.4261 | 9.4375 | 9.5726 | 0.1351 | 9.446 | 9.5627 | 0.11673 | 0.029746 | 0.95339 | 0.87169 | 1.8031 | 0 | 23 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 76542.5529 | No Control Sample | 0 | off | off | 76542.5529 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9556 | 3.2527 | 1.9938 | 568.9651 | 2.9005 | 1.6503 | 0 | 0 | 0 | 0 | 0 | 104.2372 | FALSE | 1 | 1.0395
3403 | 3404 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFTrDA | 3.95353457335163 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.68 | 9.47 | Baseline | Baseline | 54912.0087 | 73632.4221 | 0.74576 | 14739.5681 | 18072.8724 | 0.81556 | 3.7255 | 4.0742 | 14744.2567 | 18079.0353 | 0.99531 | 0.99605 | 9.7196 | 9.5032 | 0.40949 | 0.039639 | 9.5666 | 9.3301 | 9.842 | 9.6706 | 0.27537 | 0.3405 | 0.059199 | 0.063216 | 622.5302 | 205.4846 | 0.00057272 | 0.00068322 | FALSE | 1.0228 | TRUE | [M-H]- | 98.8384 | 9.6609 | 9.7802 | 0.11937 | 9.6693 | 9.7723 | 0.103 | 30.6561 | 1.0703 | 1.1762 | 1.1328 | 0 | 24 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 54912.0087 | No Control Sample | 0 | off | off | 54912.0087 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9523 | 2.7442 | 1.8193 | 618.9624 | 1.1014 | 0.68171 | 0 | 0 | 0 | 0 | 0 | 0.83089 | FALSE | 1 | 0.74576
3405 | 3406 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFTeDA | 4.01068127052494 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.87 | 9.86 | Baseline | Baseline | 24169.0132 | 16996.7998 | 1.422 | 6705.1596 | 4727.4757 | 1.4183 | 3.6045 | 3.5953 | 6705.5898 | 4730.916 | 0.98359 | 0.98914 | 9.8999 | 9.8992 | 0.30304 | 0.02991 | 9.7724 | 9.8047 | 10.0227 | 10.0029 | 0.25027 | 0.19813 | 0.057997 | 0.056982 | 547.0412 | 522.3587 | 0.00013212 | 0.0014616 | FALSE | 1.0001 | TRUE | [M-H]- | 100.267 | 9.8462 | 9.9575 | 0.11128 | 9.8532 | 9.9503 | 0.097058 | 3.5396 | 1.1848 | 1.4293 | 1.6581 | 0 | 28 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 24169.0132 | No Control Sample | 0 | off | off | 24169.0132 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9498 | 3.6011 | 2.5674 | 668.9609 | 4.4767 | 2.9947 | 0 | 0 | 0 | 0 | 0 | 89.2883 | FALSE | 1 | 1.422
3412 | 3413 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFBS | 4.00738540994558 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 3.5 | 3.5 | Manual | Manual | 43790.718 | 41284.4982 | 1.0607 | 4290.8045 | 3868.2058 | 1.1092 | 10.2057 | 10.6728 | 4306.5291 | 3881.9294 | 0.98199 | 0.9865 | 3.8724 | 3.8696 | 10.6402 | 0.37241 | 3.6361 | 3.6342 | 4.1391 | 4.1456 | 0.50303 | 0.51134 | 0.15599 | 0.14789 | 302.7363 | 900.2814 | 0.0010391 | 0 | TRUE | 1.0007 | TRUE | [M-H]- | 100.1846 | 3.7042 | 4.0768 | 0.37252 | 3.7218 | 4.0311 | 0.30929 | -8.8636 | 1.0963 | 1.0306 | 11.9821 | 0 | 79 | 26 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 43790.718 | No Control Sample | 0 | off | off | 43790.718 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9443 | 4.2363 | 1.2664 | 79.9572 | 0.15646 | 0.01251 | 0 | 0 | 0 | 0 | 0 | 0.29163 | FALSE | 1 | 1.0607
3414 | 3415 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFPeS | 4.38821015593255 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 6.38 | 3.5 | Baseline | Manual | 48282.3687 | 41284.4982 | 1.1695 | 9807.8788 | 3868.2058 | 2.5355 | 4.9228 | 10.6728 | 9825.3681 | 3881.9294 | 0.98671 | 0.9865 | 6.4086 | 3.8696 | 0.44844 | 0.02861 | 6.247 | 3.6342 | 6.5418 | 4.1456 | 0.29485 | 0.51134 | 0.075433 | 0.14789 | 289.7892 | 900.2814 | 0.0031994 | 0 | FALSE | 1.6562 | TRUE | [M-H]- | 109.7053 | 6.3253 | 6.4913 | 0.166 | 6.3377 | 6.4811 | 0.14338 | 106.4252 | 0.97497 | 0.97225 | 13.4992 | 0 | 32 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 48282.3687 | No Control Sample | 0 | off | off | 48282.3687 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9407 | 2.5658 | 0.8953 | 79.9567 | -6.2453 | -0.49935 | 0 | 0 | 0 | 0 | 0 | 1.0623 | FALSE | 1 | 1.1695
3416 | 3417 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | L-PFHxS | 3.1104907713819 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 3.244 | 4 | 0.811 | 0 | 7.34 | 7.35 | Baseline | Baseline | 36513.0836 | 62842.8687 | 0.58102 | 8887.8408 | 15564.7049 | 0.57103 | 4.1082 | 4.0375 | 8989.7153 | 15591.4515 | 0.77504 | 0.98629 | 7.3707 | 7.3704 | 0.41863 | 0.030728 | 7.2812 | 7.2572 | 7.5048 | 7.4873 | 0.22357 | 0.23012 | 0.064167 | 0.062907 | 107.1105 | 206.7584 | 0.014935 | 0.0027152 | FALSE | 1 | TRUE | [M-H]- | 95.8844 | 7.3084 | 7.4407 | 0.13224 | 7.3167 | 7.4319 | 0.11523 | -593.7507 | 0.99155 | 0.9721 | 14.3319 | 0 | 27 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 36513.0836 | No Control Sample | 0 | off | off | 36513.0836 | 0 | off | 0 | off | 0 | off | 0 | off | 398.938 | 3.4902 | 1.3924 | 79.9568 | -4.3774 | -0.35001 | 0 | 0 | 0 | 0 | 0 | 4.126 | FALSE | 0.811 | 0.58102
3418 | 3419 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | Br-PFHxS | 0.554100070533205 | Qualifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.756 | 4 | 0.189 | 0 | 7.16 | 7.35 | Manual | Baseline | 8837.1372 | 62842.8687 | 0.14062 | 1574.8162 | 15564.7049 | 0.10118 | 5.6115 | 4.0375 | 8989.7153 | 15591.4515 | 0.18765 | 0.98629 | 7.1845 | 7.3704 | 0.34186 | 0.024477 | 7.0797 | 7.2572 | 7.2802 | 7.4873 | 0.20049 | 0.23012 | 0.087718 | 0.062907 | 18.9787 | 206.7584 | 0.012392 | 0.0027152 | TRUE | 0.97478 | TRUE | [M-H]- | 73.2937 | 7.1058 | 7.2812 | 0.17545 | 7.1115 | 7.2726 | 0.16115 | 97.3408 | 1.1759 | 1.3377 | 0.24203 | 0 | 23 | 12 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8837.1372 | No Control Sample | 0 | off | off | 36513.0836 | 0 | off | 0 | off | 0 | off | 0 | off | 398.938 | 3.4902 | 1.3924 | 79.9649 | 96.8242 | 7.7418 | 0 | 0 | 0 | 0 | 0 | 4.126 | FALSE | 0.189 | 0.14062
3420 | 3421 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFHpS | 4.22823372308768 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 7.98 | 7.35 | Baseline | Baseline | 40440.0953 | 62842.8687 | 0.64351 | 9218.3793 | 15564.7049 | 0.59226 | 4.3869 | 4.0375 | 9235.1389 | 15591.4515 | 0.98734 | 0.98629 | 8.0006 | 7.3704 | 0.25767 | 0.020562 | 7.8923 | 7.2572 | 8.1303 | 7.4873 | 0.23798 | 0.23012 | 0.069764 | 0.062907 | 150.5978 | 206.7584 | 0.0031237 | 0.0027152 | FALSE | 1.0855 | TRUE | [M-H]- | 105.7058 | 7.9359 | 8.0791 | 0.14317 | 7.9439 | 8.0685 | 0.12458 | 120.9996 | 1.0945 | 1.1703 | 14.9088 | 0 | 18 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 40440.0953 | No Control Sample | 0 | off | off | 40440.0953 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9344 | 2.2412 | 1.0062 | 79.9567 | -5.6967 | -0.45549 | 0 | 0 | 0 | 0 | 0 | 1.6242 | FALSE | 1 | 0.64351
3422 | 3423 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | L-PFOS | 2.97504871294803 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 3.152 | 4 | 0.788 | 0 | 8.46 | 8.45 | Manual | Baseline | 30138.5229 | 38973.7824 | 0.7733 | 6285.3352 | 8413.8628 | 0.74702 | 4.7951 | 4.6321 | 6335.4041 | 8423.7071 | 0.68145 | 0.99223 | 8.4854 | 8.4868 | 0.30077 | 0.025446 | 8.4041 | 8.2962 | 8.581 | 8.6236 | 0.17696 | 0.32745 | 0.075338 | 0.071525 | 49.2967 | 150.9508 | 0.0050652 | 0.0019677 | TRUE | 0.99985 | TRUE | [M-H]- | 94.3861 | 8.397 | 8.5618 | 0.16488 | 8.397 | 8.5504 | 0.15348 | -179.909 | 0.99666 | 0.85548 | 15.8159 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 30138.5229 | No Control Sample | 0 | off | off | 30138.5229 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9312 | 1.8752 | 0.93557 | 79.9566 | -8.4747 | -0.67761 | 0 | 0 | 0 | 0 | 0 | 0.58945 | FALSE | 0.788 | 0.7733
3424 | 3425 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | Br-PFOS | 0.561449086067763 | Qualifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.844 | 4 | 0.211 | 0 | 8.3 | 8.45 | Manual | Baseline | 12176.439 | 38973.7824 | 0.31243 | 2148.6847 | 8413.8628 | 0.25537 | 5.6669 | 4.6321 | 6335.4041 | 8423.7071 | 0.27572 | 0.99223 | 8.3392 | 8.4868 | 0.47212 | 0.039186 | 8.2212 | 8.2962 | 8.4032 | 8.6236 | 0.18201 | 0.32745 | 0.095173 | 0.071525 | 16.8524 | 150.9508 | 0.019209 | 0.0019677 | TRUE | 0.98261 | TRUE | [M-H]- | 66.5224 | 8.2482 | 8.4157 | 0.16757 | 8.2587 | 8.4157 | 0.1571 | 226.7698 | 0.90297 | 0.90851 | 0.40402 | 0 | 13 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 12176.439 | No Control Sample | 0 | off | off | 30138.5229 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9312 | 1.8752 | 0.93557 | 79.9603 | 36.9578 | 2.955 | 0 | 0 | 0 | 0 | 0 | 0.58945 | FALSE | 0.211 | 0.31243
3426 | 3427 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFNS | 4.07688797919062 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.85 | 8.45 | Baseline | Baseline | 35217.7384 | 38973.7824 | 0.90363 | 7427.6905 | 8413.8628 | 0.88279 | 4.7414 | 4.6321 | 7433.7053 | 8423.7071 | 0.98973 | 0.99223 | 8.8849 | 8.4868 | 0.39477 | 0.034937 | 8.762 | 8.2962 | 9.0708 | 8.6236 | 0.30883 | 0.32745 | 0.074005 | 0.071525 | 160.1323 | 150.9508 | 1.5436e-06 | 0.0019677 | FALSE | 1.0469 | TRUE | [M-H]- | 101.9222 | 8.8079 | 8.9676 | 0.15964 | 8.8167 | 8.9564 | 0.13962 | -0.037126 | 1.0055 | 0.97759 | 16.0972 | 0 | 16 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 35217.7384 | No Control Sample | 0 | off | off | 35217.7384 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9273 | 0.47199 | 0.25909 | 79.9567 | -4.5001 | -0.35982 | 0 | 0 | 0 | 0 | 0 | 2.2892 | FALSE | 1 | 0.90363
3428 | 3429 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFDS | 3.91090106617076 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.19 | 8.45 | Baseline | Baseline | 34711.5187 | 38973.7824 | 0.89064 | 7483.5669 | 8413.8628 | 0.88943 | 4.6384 | 4.6321 | 7486.5708 | 8423.7071 | 0.9927 | 0.99223 | 9.2148 | 8.4868 | 0.27039 | 0.024849 | 9.0717 | 8.2962 | 9.396 | 8.6236 | 0.32432 | 0.32745 | 0.072773 | 0.071525 | 184.6873 | 150.9508 | 0.00092361 | 0.0019677 | FALSE | 1.0858 | TRUE | [M-H]- | 97.7725 | 9.1417 | 9.2953 | 0.15359 | 9.152 | 9.2834 | 0.13147 | 21.3121 | 1.1574 | 1.3445 | 14.4673 | 0 | 18 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 34711.5187 | No Control Sample | 0 | off | off | 34711.5187 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9244 | 0.87319 | 0.52298 | 79.9568 | -1.6477 | -0.13174 | 0 | 0 | 0 | 0 | 0 | 1.6805 | FALSE | 1 | 0.89064
3432 | 3433 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FBSA | 3.15439898116958 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 4 | 4 | 1 | 0 | 4.6 | 8.86 | Manual | Baseline | 59185.2987 | 71158.0786 | 0.83174 | 5313.137 | 15042.1872 | 0.35322 | 11.1394 | 4.7306 | 5367.0774 | 15046.7882 | 0.96257 | 0.99672 | 4.9783 | 9.0271 | 8.2242 | 0.37831 | 4.7523 | 8.8329 | 5.2029 | 9.1936 | 0.45062 | 0.36073 | 0.16912 | 0.073668 | 184.0693 | 158.3176 | 0 | 0.00057489 | TRUE | 0.55148 | TRUE | [M-H]- | 78.86 | 4.7914 | 5.1809 | 0.38947 | 4.8122 | 5.1602 | 0.34808 | 0 | 1.1999 | 1.4594 | 5.8358 | 0 | 81 | 32 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 59185.2987 | No Control Sample | 0 | off | off | 59185.2987 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9606 | 5.608 | 1.671 | 77.9648 | -6.7894 | -0.52934 | 0 | 0 | 0 | 0 | 0 | 0.49683 | FALSE | 1 | 0.83174
3434 | 3435 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FPeSA | 4.04617226624746 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7 | 8.86 | Baseline | Baseline | 117892.9739 | 71158.0786 | 1.6568 | 25203.7468 | 15042.1872 | 1.6755 | 4.6776 | 4.7306 | 25242.6632 | 15046.7882 | 0.99003 | 0.99672 | 7.1228 | 9.0271 | 1.7549 | 0.12284 | 6.9685 | 8.8329 | 7.2632 | 9.1936 | 0.29467 | 0.36073 | 0.072602 | 0.073668 | 236.9965 | 158.3176 | 0.0028617 | 0.00057489 | FALSE | 0.78905 | TRUE | [M-H]- | 101.1543 | 7.0453 | 7.2031 | 0.15784 | 7.0559 | 7.1912 | 0.13536 | 244.7726 | 1.0204 | 1.0274 | 3.8445 | 0 | 32 | 9 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 117892.9739 | No Control Sample | 0 | off | off | 117892.9739 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9569 | 3.1212 | 1.086 | 77.9648 | -6.5977 | -0.51439 | 0 | 0 | 0 | 0 | 0 | 1.1326 | FALSE | 1 | 1.6568
3436 | 3437 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FHxSA | 4.10300685280981 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7.86 | 8.86 | Baseline | Baseline | 59657.5183 | 71158.0786 | 0.83838 | 12867.8698 | 15042.1872 | 0.85545 | 4.6362 | 4.7306 | 12870.38 | 15046.7882 | 0.99735 | 0.99672 | 8.0045 | 9.0271 | 1.839 | 0.14455 | 7.836 | 8.8329 | 8.2398 | 9.1936 | 0.40377 | 0.36073 | 0.072624 | 0.073668 | 94.7726 | 158.3176 | 0.0004671 | 0.00057489 | FALSE | 0.88672 | TRUE | [M-H]- | 102.5752 | 7.9354 | 8.0874 | 0.15201 | 7.9443 | 8.077 | 0.13271 | 14.8862 | 1.2018 | 1.4388 | 8.1028 | 0 | 29 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 59657.5183 | No Control Sample | 0 | off | off | 59657.5183 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9539 | 3.3133 | 1.3185 | 77.9651 | -3.7729 | -0.29416 | 0 | 0 | 0 | 0 | 0 | 1.1406 | FALSE | 1 | 0.83838
3438 | 3439 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FHpSA | 3.89741549197305 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.58 | 8.86 | Baseline | Baseline | 117375.4354 | 71158.0786 | 1.6495 | 26158.7368 | 15042.1872 | 1.739 | 4.487 | 4.7306 | 26211.0419 | 15046.7882 | 0.98664 | 0.99672 | 8.5863 | 9.0271 | 0.072899 | 0.0062548 | 8.4566 | 8.8329 | 8.7405 | 9.1936 | 0.28387 | 0.36073 | 0.070303 | 0.073668 | 123.2192 | 158.3176 | 0.00088322 | 0.00057489 | FALSE | 0.95116 | TRUE | [M-H]- | 97.4354 | 8.5142 | 8.6622 | 0.14794 | 8.521 | 8.6516 | 0.13061 | -81.3902 | 1.084 | 1.1249 | 3.7718 | 0 | 17 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 117375.4354 | No Control Sample | 0 | off | off | 117375.4354 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9504 | 2.2661 | 1.0151 | 77.9649 | -6.919 | -0.53944 | 0 | 0 | 0 | 0 | 0 | 1.6894 | FALSE | 1 | 1.6495
3440 | 3441 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FOSA | 4.23289439454757 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.88 | 8.86 | Baseline | Baseline | 68534.1638 | 71158.0786 | 0.96313 | 14824.0929 | 15042.1872 | 0.9855 | 4.6232 | 4.7306 | 14827.4961 | 15046.7882 | 0.99367 | 0.99672 | 9.0295 | 9.0271 | 1.6836 | 0.1495 | 8.8439 | 8.8329 | 9.2456 | 9.1936 | 0.40175 | 0.36073 | 0.070984 | 0.073668 | 148.8093 | 158.3176 | 0.00047829 | 0.00057489 | FALSE | 1.0003 | TRUE | [M-H]- | 105.8224 | 8.9494 | 9.1054 | 0.15598 | 8.9578 | 9.0962 | 0.13836 | 17.6483 | 0.94807 | 0.87386 | 7.9803 | 0 | 21 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 68534.1638 | No Control Sample | 0 | off | off | 68534.1638 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9468 | 1.1842 | 0.58969 | 77.965 | -5.3925 | -0.42043 | 0 | 0 | 0 | 0 | 0 | 2.0957 | FALSE | 1 | 0.96313
3459 | 3460 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 4:2 FTS | 3.90918290395508 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 5.94 | 5.92 | Baseline | Baseline | 35828.2636 | 18050.5527 | 1.9849 | 6003.8421 | 3227.1166 | 1.8604 | 5.9676 | 5.5934 | 6015.4799 | 3235.0378 | 0.98844 | 0.99026 | 5.9639 | 5.9634 | 0.40277 | 0.023924 | 5.8084 | 5.8279 | 6.1162 | 6.1189 | 0.30785 | 0.29097 | 0.091362 | 0.085449 | 185.3892 | 164.5324 | 0.0028527 | 0.0044259 | FALSE | 1.0001 | TRUE | [M-H]- | 97.7296 | 5.864 | 6.0662 | 0.20228 | 5.878 | 6.0522 | 0.17411 | 55.6353 | 0.96388 | 0.91664 | 4.7936 | 0 | 36 | 12 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 35828.2636 | No Control Sample | 0 | off | off | 35828.2636 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9761 | 5.3924 | 1.7632 | 306.9692 | 5.373 | 1.6493 | 0 | 0 | 0 | 0 | 0 | 5.3291 | FALSE | 1 | 1.9849
3461 | 3462 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 6:2 FTS | 4.1470258892703 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7.92 | 7.92 | Baseline | Baseline | 37438.5999 | 14879.5686 | 2.5161 | 9247.7858 | 3568.5322 | 2.5915 | 4.0484 | 4.1697 | 9248.0181 | 3568.5385 | 0.99944 | 0.99316 | 7.9381 | 7.9393 | 0.22831 | 0.018082 | 7.8133 | 7.8053 | 8.1479 | 8.1062 | 0.33462 | 0.30085 | 0.059576 | 0.062502 | 79.404 | 82.6616 | 4.0455e-05 | 3.0335e-06 | FALSE | 0.99985 | TRUE | [M-H]- | 103.6756 | 7.8693 | 8.0154 | 0.14607 | 7.8799 | 8.0015 | 0.12165 | -1.1181 | 0.99731 | 0.94198 | 4.4476 | 0 | 25 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 37438.5999 | No Control Sample | 0 | off | off | 37438.5999 | 0 | off | 0 | off | 0 | off | 0 | off | 426.97 | 4.8531 | 2.0721 | 406.9633 | 4.4923 | 1.8282 | 0 | 0 | 0 | 0 | 0 | 5.5018 | FALSE | 1 | 2.5161
3463 | 3464 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 8:2 FTS | 4.07745183548025 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.88 | 8.89 | Baseline | Baseline | 33488.0187 | 16738.9038 | 2.0006 | 6743.1426 | 3305.7326 | 2.0398 | 4.9662 | 5.0636 | 6752.9185 | 3308.5383 | 0.99129 | 0.99319 | 8.8939 | 8.897 | 0.15611 | 0.013863 | 8.744 | 8.7522 | 9.0351 | 9.0434 | 0.29115 | 0.29115 | 0.078037 | 0.079 | 142.7305 | 136.2308 | 0.0028774 | 0.0015704 | FALSE | 0.99965 | TRUE | [M-H]- | 101.9363 | 8.8142 | 8.9841 | 0.16988 | 8.8268 | 8.9701 | 0.14327 | 66.6425 | 1.1349 | 1.3008 | 4.0158 | 0 | 15 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 33488.0187 | No Control Sample | 0 | off | off | 33488.0187 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9634 | 3.6605 | 1.929 | 506.9574 | 3.2621 | 1.6537 | 0 | 0 | 0 | 0 | 0 | 5.7212 | FALSE | 1 | 2.0006
3465 | 3466 | PFAS CS6 4ng/ml | 25 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 26) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 10:2 FTS | 2.95958355726113 | Quantifiers | 10/27/2023 01:46:07 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 4 | 4 | 1 | 0 | 9.5 | 8.89 | Baseline | Baseline | 25304.0862 | 16738.9038 | 1.5117 | 5941.7533 | 3305.7326 | 1.7974 | 4.2587 | 5.0636 | 5941.771 | 3308.5383 | 0.99853 | 0.99319 | 9.5178 | 8.897 | 0.18747 | 0.017809 | 9.3802 | 8.7522 | 9.6415 | 9.0434 | 0.26128 | 0.29115 | 0.066456 | 0.079 | 257.4438 | 136.2308 | 5.7865e-06 | 0.0015704 | FALSE | 1.0698 | TRUE | [M-H]- | 73.9896 | 9.4498 | 9.5891 | 0.13937 | 9.4577 | 9.58 | 0.12232 | 0.13159 | 1.1311 | 1.2792 | 5.0075 | 0 | 18 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 25304.0862 | No Control Sample | 0 | off | off | 25304.0862 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9568 | 2.7404 | 1.7181 | 606.9531 | 8.5285 | 5.1763 | 0 | 0 | 0 | 0 | 0 | 0.71191 | FALSE | 1 | 1.5117
3526 | 3527 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFBA | 3.96187627920838 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 1.49 | 1.48 | Baseline | Baseline | 75123.9323 | 105709.9514 | 0.71066 | 10795.5303 | 14496.1052 | 0.74472 | 6.9588 | 7.2923 | 10908.9892 | 14500.1451 | 0.94986 | 0.99685 | 1.4738 | 1.4716 | 1.0905 | 0.016249 | 1.3529 | 1.3186 | 1.657 | 1.7444 | 0.30415 | 0.42578 | 0.10773 | 0.11375 | 211.5676 | 199.9003 | 0.0094696 | 0.00083501 | FALSE | 1.0015 | TRUE | [M-H]- | 99.0469 | 1.3842 | 1.6256 | 0.24147 | 1.3923 | 1.6067 | 0.21443 | 336.1161 | 1.608 | 2.2019 | 1.6453 | 0 | 61 | 23 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 75123.9323 | No Control Sample | 0 | off | off | 75123.9323 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9807 | 6.8463 | 1.4581 | 168.9891 | 0.84337 | 0.14252 | 0 | 0 | 0 | 0 | 0 | 2.1074 | FALSE | 1 | 0.71066
3528 | 3529 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFPeA | 3.6187406956258 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 3.18 | 3.21 | Manual | Baseline | 64070.7103 | 79168.7964 | 0.80929 | 7204.7926 | 8558.0661 | 0.84187 | 8.8928 | 9.2508 | 7217.1246 | 8623.2938 | 0.98722 | 0.97193 | 3.2263 | 3.2231 | 1.4563 | 0.046309 | 3.0223 | 3.0214 | 3.4852 | 3.461 | 0.46285 | 0.4396 | 0.1297 | 0.14098 | 556.3992 | 815.612 | 0.0016189 | 0.0068134 | TRUE | 1.001 | TRUE | [M-H]- | 90.4685 | 3.0767 | 3.4216 | 0.34492 | 3.0973 | 3.3789 | 0.28164 | -25.1996 | 1.1457 | 1.167 | 2.0765 | 0 | 80 | 25 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 64070.7103 | No Control Sample | 0 | off | off | 64070.7103 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9776 | 6.0838 | 1.5999 | 218.9862 | 1.6604 | 0.36361 | 0 | 0 | 0 | 0 | 0 | 4.2957 | FALSE | 1 | 0.80929
3530 | 3531 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFHxA | 3.98203837531122 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 68.5942 | 6.13 | 6.12 | Baseline | Baseline | 139824.8997 | 150879.7779 | 0.92673 | 24796.9474 | 27241.5685 | 0.91026 | 5.6388 | 5.5386 | 24845.5462 | 27248.6139 | 0.98838 | 0.99707 | 6.1573 | 6.1543 | 0.44579 | 0.027327 | 5.9683 | 5.9709 | 6.3074 | 6.4037 | 0.33908 | 0.4328 | 0.087122 | 0.08756 | 287.5726 | 274.8875 | 0.003345 | 0.00058936 | FALSE | 1.0005 | TRUE | [M-H]- | 99.551 | 6.0631 | 6.2482 | 0.18514 | 6.0765 | 6.2365 | 0.16001 | 244.6182 | 0.93751 | 0.87565 | 1.4919 | 0 | 37 | 12 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 139824.8997 | No Control Sample | 203843.5852 | off | off | 139824.8997 | 203843.5852 | off | 0.25 | off | 4 | off | 2.7438 | off | 312.9745 | 5.3441 | 1.6725 | 268.9829 | 1.9099 | 0.51372 | 0 | 0 | 0 | 0 | 0 | 42.8392 | FALSE | 1 | 0.92673
3532 | 3533 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFHpA | 4.36354735674262 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 80.5765 | 7.26 | 7.27 | Baseline | Baseline | 183979.4086 | 168500.4181 | 1.0919 | 42848.9414 | 38945.5617 | 1.1002 | 4.2937 | 4.3266 | 42864.7464 | 38962.9055 | 0.99561 | 0.99678 | 7.2942 | 7.2947 | 0.47083 | 0.034182 | 7.1404 | 7.1525 | 7.4667 | 7.4607 | 0.32632 | 0.30822 | 0.066893 | 0.068159 | 228.2584 | 244.1407 | 0.00066552 | 0.00086867 | FALSE | 0.99993 | TRUE | [M-H]- | 109.0887 | 7.2265 | 7.3669 | 0.14043 | 7.2357 | 7.3569 | 0.12124 | 87.3901 | 1.0033 | 0.9944 | 1.5909 | 0 | 38 | 10 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Pass | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 183979.4086 | No Control Sample | 228328.7703 | off | off | 183979.4086 | 228328.7703 | off | 0.25 | off | 4 | off | 3.2231 | off | 362.9711 | 4.0566 | 1.4724 | 318.9801 | 3.1328 | 0.99931 | 0 | 0 | 0 | 0 | 0 | 1.0183 | FALSE | 1 | 1.0919
3534 | 3535 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFOA | 4.36619166659884 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 88.1049 | 7.96 | 7.94 | Baseline | Baseline | 168699.2463 | 159316.8195 | 1.0589 | 40100.4775 | 36898.4326 | 1.0868 | 4.2069 | 4.3177 | 40156.4525 | 36920.4864 | 0.98837 | 0.99542 | 7.9784 | 7.9782 | 0.23075 | 0.018368 | 7.8351 | 7.8415 | 8.1422 | 8.1337 | 0.30712 | 0.29213 | 0.064788 | 0.067327 | 137.2949 | 136.9191 | 0.00069878 | 0.0012436 | FALSE | 1 | TRUE | [M-H]- | 109.1548 | 7.9155 | 8.0563 | 0.14087 | 7.9235 | 8.0447 | 0.12127 | 91.24 | 1.2237 | 1.4478 | 1.4915 | 0 | 23 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 168699.2463 | No Control Sample | 191475.4056 | off | off | 168699.2463 | 191475.4056 | off | 0.25 | off | 4 | off | 3.5242 | off | 412.9678 | 3.4166 | 1.4109 | 368.9772 | 2.9928 | 1.1043 | 0 | 0 | 0 | 0 | 0 | 0.63333 | FALSE | 1 | 1.0589
3536 | 3537 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFNA | 3.87054218985717 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 83.5618 | 8.46 | 8.47 | Baseline | Baseline | 132983.0127 | 147897.5726 | 0.89916 | 27895.2644 | 30817.5556 | 0.90517 | 4.7672 | 4.7991 | 27917.3459 | 30832.6251 | 0.99401 | 0.9966 | 8.4935 | 8.4917 | 0.39569 | 0.033475 | 8.3358 | 8.2943 | 8.6808 | 8.6523 | 0.34497 | 0.35797 | 0.07418 | 0.075435 | 143.388 | 141.8592 | 0.0015441 | 0.00090471 | FALSE | 1.0002 | TRUE | [M-H]- | 96.7636 | 8.4123 | 8.5723 | 0.16003 | 8.4203 | 8.5606 | 0.14023 | 124.8631 | 0.97373 | 0.89181 | 1.4173 | 0 | 20 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 132983.0127 | No Control Sample | 159143.3885 | off | off | 132983.0127 | 159143.3885 | off | 0.25 | off | 4 | off | 3.3425 | off | 462.9653 | 4.4646 | 2.0669 | 418.9746 | 3.3042 | 1.3844 | 0 | 0 | 0 | 0 | 0 | 0.9446 | FALSE | 1 | 0.89916
3538 | 3539 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFDA | 3.90416378851823 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.87 | 8.86 | Baseline | Baseline | 120197.1674 | 124018.497 | 0.96919 | 25204.1122 | 25509.4328 | 0.98803 | 4.769 | 4.8617 | 25209.2155 | 25523.6314 | 0.99791 | 0.9965 | 8.9045 | 8.9065 | 0.38872 | 0.03448 | 8.7379 | 8.7467 | 9.1238 | 9.0552 | 0.38583 | 0.3085 | 0.073996 | 0.076155 | 192.7889 | 138.3797 | 0.00047259 | 0.0010289 | FALSE | 0.99977 | TRUE | [M-H]- | 97.6041 | 8.826 | 8.9875 | 0.16149 | 8.8381 | 8.9787 | 0.14063 | 30.8714 | 1.0529 | 1.1754 | 1.2194 | 0 | 20 | 5 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 120197.1674 | No Control Sample | 0 | off | off | 120197.1674 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9622 | 4.2604 | 2.1854 | 468.9719 | 3.978 | 1.8656 | 0 | 0 | 0 | 0 | 0 | 24.4122 | FALSE | 1 | 0.96919
3540 | 3541 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFUdA | 3.88243280672166 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.21 | 9.2 | Baseline | Baseline | 87710.1927 | 101870.845 | 0.86099 | 19167.2305 | 21711.3498 | 0.88282 | 4.576 | 4.6921 | 19170.3053 | 21720.3678 | 0.99732 | 0.99574 | 9.2415 | 9.2395 | 0.34198 | 0.031497 | 9.0856 | 9.0748 | 9.4075 | 9.4295 | 0.3219 | 0.35468 | 0.071475 | 0.073744 | 228.9437 | 147.3595 | 0.00031759 | 0.00093379 | FALSE | 1.0002 | TRUE | [M-H]- | 97.0608 | 9.1669 | 9.3204 | 0.15351 | 9.1752 | 9.3118 | 0.13658 | 18.9109 | 1.0184 | 1.0348 | 1.3639 | 0 | 18 | 5 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 87710.1927 | No Control Sample | 0 | off | off | 87710.1927 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9591 | 3.9588 | 2.2286 | 518.9683 | 1.8451 | 0.95755 | 0 | 0 | 0 | 0 | 0 | 0.54761 | FALSE | 1 | 0.86099
3542 | 3543 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFDoA | 4.06678901288235 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.47 | 9.47 | Baseline | Baseline | 81129.3533 | 77090.1571 | 1.0524 | 20740.648 | 19429.9869 | 1.0675 | 3.9116 | 3.9676 | 20748.0895 | 19435.3492 | 0.99528 | 0.99638 | 9.5081 | 9.5103 | 0.40207 | 0.038076 | 9.3746 | 9.3638 | 9.6372 | 9.695 | 0.26265 | 0.3312 | 0.06126 | 0.061986 | 553.5893 | 367.3172 | 0.00070016 | 0.00061141 | FALSE | 0.99977 | TRUE | [M-H]- | 101.6697 | 9.443 | 9.5743 | 0.13128 | 9.4495 | 9.5642 | 0.11468 | 55.2892 | 1.0258 | 0.99608 | 1.6023 | 0 | 18 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 81129.3533 | No Control Sample | 0 | off | off | 81129.3533 | 0 | off | 0 | off | 0 | off | 0 | off | 612.956 | 3.8111 | 2.3361 | 568.9656 | 3.816 | 2.1712 | 0 | 0 | 0 | 0 | 0 | 111.737 | FALSE | 1 | 1.0524
3544 | 3545 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFTrDA | 4.0248546911722 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.68 | 9.47 | Baseline | Baseline | 58527.7514 | 77090.1571 | 0.75921 | 16185.2183 | 19429.9869 | 0.833 | 3.6161 | 3.9676 | 16190.8259 | 19435.3492 | 0.99535 | 0.99638 | 9.7216 | 9.5103 | 0.42951 | 0.041577 | 9.5942 | 9.3638 | 9.8713 | 9.695 | 0.27712 | 0.3312 | 0.055745 | 0.061986 | 482.5399 | 367.3172 | 0.00070937 | 0.00061141 | FALSE | 1.0222 | TRUE | [M-H]- | 100.6214 | 9.6652 | 9.7821 | 0.11691 | 9.671 | 9.7741 | 0.10306 | 41.4311 | 0.94221 | 0.83276 | 1.431 | 0 | 25 | 6 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 58527.7514 | No Control Sample | 0 | off | off | 58527.7514 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9529 | 3.6302 | 2.4066 | 618.9634 | 2.7756 | 1.718 | 0 | 0 | 0 | 0 | 0 | 0.91679 | FALSE | 1 | 0.75921
3546 | 3547 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFTeDA | 4.10714635419518 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 9.87 | 9.86 | Baseline | Baseline | 26785.1912 | 18394.2024 | 1.4562 | 7625.727 | 5433.2042 | 1.4035 | 3.5125 | 3.3855 | 7638.9245 | 5437.3385 | 0.97972 | 0.99374 | 9.9012 | 9.9005 | 0.31632 | 0.031221 | 9.8017 | 9.7837 | 9.9928 | 10.0087 | 0.19117 | 0.22507 | 0.055273 | 0.05287 | 614.2898 | 843.6191 | 0.0032129 | 0.0014795 | FALSE | 1.0001 | TRUE | [M-H]- | 102.6787 | 9.8451 | 9.9585 | 0.11342 | 9.8532 | 9.9515 | 0.098337 | 128.1633 | 0.99044 | 0.99954 | 1.6124 | 0 | 21 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 26785.1912 | No Control Sample | 0 | off | off | 26785.1912 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9495 | 3.0731 | 2.1909 | 668.9604 | 3.7616 | 2.5163 | 0 | 0 | 0 | 0 | 0 | 96.31 | FALSE | 1 | 1.4562
3553 | 3554 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFBS | 3.67118337876489 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 3.5 | 3.5 | Manual | Manual | 43132.9658 | 44388.3804 | 0.97172 | 4165.3866 | 3959.5792 | 1.052 | 10.3551 | 11.2104 | 4175.8998 | 3966.8624 | 0.98952 | 0.99038 | 3.8784 | 3.8779 | 10.8115 | 0.3784 | 3.6402 | 3.6301 | 4.1823 | 4.162 | 0.54209 | 0.53194 | 0.15983 | 0.17053 | 345.3666 | 803.3104 | 0.0016202 | 0.00075895 | TRUE | 1.0001 | TRUE | [M-H]- | 91.7796 | 3.7106 | 4.0848 | 0.37417 | 3.7356 | 4.0545 | 0.31889 | -12.4494 | 1.1048 | 1.2097 | 12.2592 | 0 | 85 | 27 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 43132.9658 | No Control Sample | 0 | off | off | 43132.9658 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9444 | 4.6441 | 1.3883 | 79.9567 | -5.7309 | -0.45823 | 0 | 0 | 0 | 0 | 0 | 1.0937 | FALSE | 1 | 0.97172
3555 | 3556 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFPeS | 4.1087964903024 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 6.38 | 3.5 | Baseline | Manual | 48606.9164 | 44388.3804 | 1.095 | 10042.3444 | 3959.5792 | 2.5362 | 4.8402 | 11.2104 | 10049.026 | 3966.8624 | 0.99274 | 0.99038 | 6.4056 | 3.8779 | 0.40163 | 0.025624 | 6.2387 | 3.6301 | 6.5701 | 4.162 | 0.33132 | 0.53194 | 0.074313 | 0.17053 | 292.6771 | 803.3104 | 0.0013311 | 0.00075895 | FALSE | 1.6518 | TRUE | [M-H]- | 102.7199 | 6.3252 | 6.4869 | 0.16175 | 6.3357 | 6.4771 | 0.14145 | 40.3447 | 1.0532 | 1.1339 | 11.8267 | 0 | 36 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 48606.9164 | No Control Sample | 0 | off | off | 48606.9164 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9408 | 2.8362 | 0.98966 | 79.9565 | -8.4697 | -0.67722 | 0 | 0 | 0 | 0 | 0 | 1.0181 | FALSE | 1 | 1.095
3557 | 3558 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | L-PFHxS | 3.24087280663932 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 3.244 | 4 | 0.811 | 0 | 7.34 | 7.35 | Baseline | Baseline | 37502.0005 | 61948.2234 | 0.60538 | 9298.6727 | 15348.8035 | 0.60582 | 4.033 | 4.036 | 9418.2394 | 15370.5409 | 0.77395 | 0.98912 | 7.3697 | 7.3688 | 0.40466 | 0.029702 | 7.281 | 7.2389 | 7.4943 | 7.496 | 0.2133 | 0.25713 | 0.06351 | 0.062308 | 130.7237 | 210.0935 | 0.019151 | 0.0015666 | FALSE | 1.0001 | TRUE | [M-H]- | 99.9036 | 7.306 | 7.4385 | 0.13247 | 7.3134 | 7.4292 | 0.11584 | -834.8573 | 1.0882 | 1.164 | 15.4372 | 0 | 26 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 37502.0005 | No Control Sample | 0 | off | off | 37502.0005 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9379 | 3.1626 | 1.2617 | 79.9571 | -1.0054 | -0.080392 | 0 | 0 | 0 | 0 | 0 | 4.0881 | FALSE | 0.811 | 0.60538
3559 | 3560 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | Br-PFHxS | 0.56906556560567 | Qualifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.756 | 4 | 0.189 | 0 | 7.16 | 7.35 | Manual | Baseline | 8946.611 | 61948.2234 | 0.14442 | 1576.8632 | 15348.8035 | 0.10274 | 5.6737 | 4.036 | 9418.2394 | 15370.5409 | 0.18474 | 0.98912 | 7.1868 | 7.3688 | 0.37457 | 0.026819 | 7.0803 | 7.2389 | 7.2867 | 7.496 | 0.20637 | 0.25713 | 0.089405 | 0.062308 | 22.1681 | 210.0935 | 0.026803 | 0.0015666 | TRUE | 0.9753 | TRUE | [M-H]- | 75.2732 | 7.0995 | 7.2896 | 0.19011 | 7.1072 | 7.2763 | 0.1691 | 204.7998 | 0.95349 | 0.83912 | 0.23856 | 0 | 24 | 12 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 8946.611 | No Control Sample | 0 | off | off | 37502.0005 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9379 | 3.1626 | 1.2617 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4.0881 | FALSE | 0.189 | 0.14442
3561 | 3562 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFHpS | 4.36828028329706 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7.98 | 7.35 | Baseline | Baseline | 41184.7596 | 61948.2234 | 0.66483 | 9601.8789 | 15348.8035 | 0.62558 | 4.2892 | 4.036 | 9613.8548 | 15370.5409 | 0.99113 | 0.98912 | 8.0028 | 7.3688 | 0.28618 | 0.022837 | 7.8785 | 7.2389 | 8.1457 | 7.496 | 0.26715 | 0.25713 | 0.067253 | 0.062308 | 120.2104 | 210.0935 | 0.002431 | 0.0015666 | FALSE | 1.086 | TRUE | [M-H]- | 109.207 | 7.94 | 8.0816 | 0.14156 | 7.947 | 8.0719 | 0.12495 | 87.3753 | 1.1592 | 1.3115 | 14.4989 | 0 | 20 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 41184.7596 | No Control Sample | 0 | off | off | 41184.7596 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9343 | 1.9184 | 0.86123 | 79.9569 | -3.2173 | -0.25724 | 0 | 0 | 0 | 0 | 0 | 1.1228 | FALSE | 1 | 0.66483
3563 | 3564 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | L-PFOS | 2.70495855550445 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 3.152 | 4 | 0.788 | 0 | 8.46 | 8.45 | Manual | Baseline | 29617.4828 | 42124.2521 | 0.7031 | 6178.8968 | 9147.6884 | 0.67546 | 4.7933 | 4.6049 | 6258.1565 | 9162.0411 | 0.6733 | 0.99012 | 8.4916 | 8.4908 | 0.37397 | 0.031638 | 8.4104 | 8.3666 | 8.5948 | 8.6228 | 0.18443 | 0.25618 | 0.074485 | 0.072618 | 54.6029 | 219.6948 | 0.0050896 | 0.0016772 | TRUE | 1.0001 | TRUE | [M-H]- | 85.8172 | 8.3962 | 8.5668 | 0.17064 | 8.3962 | 8.5585 | 0.16236 | -170.514 | 0.83088 | 0.58121 | 15.726 | 0 | 12 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 29617.4828 | No Control Sample | 0 | off | off | 29617.4828 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9316 | 2.6877 | 1.341 | 79.9569 | -4.9396 | -0.39496 | 0 | 0 | 0 | 0 | 0 | 1.3938 | FALSE | 0.788 | 0.7031
3565 | 3566 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | Br-PFOS | 0.532670209484639 | Qualifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.844 | 4 | 0.211 | 0 | 8.3 | 8.45 | Manual | Baseline | 12486.1335 | 42124.2521 | 0.29641 | 2019.1421 | 9147.6884 | 0.22073 | 6.1839 | 4.6049 | 6258.1565 | 9162.0411 | 0.28385 | 0.99012 | 8.344 | 8.4908 | 0.52988 | 0.04398 | 8.2318 | 8.3666 | 8.4146 | 8.6228 | 0.18275 | 0.25618 | 0.10249 | 0.072618 | 17.8431 | 219.6948 | 0.0045638 | 0.0016772 | TRUE | 0.98271 | TRUE | [M-H]- | 63.1126 | 8.2467 | 8.4149 | 0.16822 | 8.2588 | 8.4149 | 0.15615 | -50.4251 | 0.89988 | 0.91825 | 0.42158 | 0 | 12 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 12486.1335 | No Control Sample | 0 | off | off | 29617.4828 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9316 | 2.6877 | 1.341 | 79.9589 | 19.9347 | 1.5939 | 0 | 0 | 0 | 0 | 0 | 1.3938 | FALSE | 0.211 | 0.29641
3567 | 3568 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFNS | 3.85817237606618 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.85 | 8.45 | Baseline | Baseline | 36022.5091 | 42124.2521 | 0.85515 | 7444.6215 | 9147.6884 | 0.81383 | 4.8387 | 4.6049 | 7452.2249 | 9162.0411 | 0.9905 | 0.99012 | 8.8909 | 8.4908 | 0.462 | 0.040887 | 8.7233 | 8.3666 | 9.0701 | 8.6228 | 0.34673 | 0.25618 | 0.075338 | 0.072618 | 146.0196 | 219.6948 | 0.0021421 | 0.0016772 | FALSE | 1.0471 | TRUE | [M-H]- | 96.4543 | 8.8113 | 8.9766 | 0.16524 | 8.8245 | 8.965 | 0.14054 | 45.9922 | 1.0977 | 1.2629 | 15.4162 | 0 | 18 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 36022.5091 | No Control Sample | 0 | off | off | 36022.5091 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9279 | 1.6166 | 0.88739 | 79.9568 | -4.1769 | -0.33398 | 0 | 0 | 0 | 0 | 0 | 1.5736 | FALSE | 1 | 0.85515
3569 | 3570 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFDS | 3.72881903312147 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.19 | 8.45 | Baseline | Baseline | 35770.7248 | 42124.2521 | 0.84917 | 7740.036 | 9147.6884 | 0.84612 | 4.6215 | 4.6049 | 7740.0374 | 9162.0411 | 0.99295 | 0.99012 | 9.2206 | 8.4908 | 0.3329 | 0.030594 | 9.0335 | 8.3666 | 9.4119 | 8.6228 | 0.37833 | 0.25618 | 0.072316 | 0.072618 | 208.1451 | 219.6948 | 3.4953e-07 | 0.0016772 | FALSE | 1.086 | TRUE | [M-H]- | 93.2205 | 9.1472 | 9.2994 | 0.15228 | 9.1547 | 9.2863 | 0.13161 | 0.0071507 | 0.95267 | 0.81927 | 14.7518 | 0 | 21 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 35770.7248 | No Control Sample | 0 | off | off | 35770.7248 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9243 | 0.75153 | 0.45011 | 79.9567 | -2.9556 | -0.23632 | 0 | 0 | 0 | 0 | 0 | 2.9109 | FALSE | 1 | 0.84917
3573 | 3574 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FBSA | 3.03405091678149 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 4 | 4 | 1 | 0 | 4.6 | 8.86 | Manual | Baseline | 59301.0482 | 74125.3059 | 0.80001 | 5424.157 | 16003.0031 | 0.33895 | 10.9328 | 4.632 | 5479.2246 | 16011.0204 | 0.96392 | 0.99629 | 4.9778 | 9.0334 | 8.2131 | 0.3778 | 4.7458 | 8.8948 | 5.2101 | 9.2128 | 0.46432 | 0.31798 | 0.16745 | 0.071183 | 222.6855 | 115.8681 | 0 | 0.00022925 | TRUE | 0.55105 | TRUE | [M-H]- | 75.8513 | 4.7818 | 5.1761 | 0.39433 | 4.8196 | 5.1531 | 0.33352 | 0 | 1.0077 | 1.1127 | 5.8418 | 0 | 83 | 31 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 59301.0482 | No Control Sample | 0 | off | off | 59301.0482 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9607 | 5.6219 | 1.6751 | 77.965 | -3.3656 | -0.2624 | 0 | 0 | 0 | 0 | 0 | 0.27596 | FALSE | 1 | 0.80001
3575 | 3576 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FPeSA | 3.9495747438416 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7 | 8.86 | Baseline | Baseline | 119877.0858 | 74125.3059 | 1.6172 | 25146.7436 | 16003.0031 | 1.5714 | 4.7671 | 4.632 | 25178.222 | 16011.0204 | 0.99111 | 0.99629 | 7.1194 | 9.0334 | 1.7055 | 0.11939 | 6.9578 | 8.8948 | 7.2715 | 9.2128 | 0.31368 | 0.31798 | 0.07324 | 0.071183 | 212.9491 | 115.8681 | 0.0023716 | 0.00022925 | FALSE | 0.78812 | TRUE | [M-H]- | 98.7394 | 7.0384 | 7.1991 | 0.16074 | 7.0499 | 7.1871 | 0.13722 | 190.1215 | 0.95669 | 0.89219 | 3.8109 | 0 | 34 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 119877.0858 | No Control Sample | 0 | off | off | 119877.0858 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9569 | 3.085 | 1.0735 | 77.9648 | -5.7855 | -0.45107 | 0 | 0 | 0 | 0 | 0 | 1.6809 | FALSE | 1 | 1.6172
3577 | 3578 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FHxSA | 4.19339484000733 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7.86 | 8.86 | Baseline | Baseline | 63514.2221 | 74125.3059 | 0.85685 | 13582.6906 | 16003.0031 | 0.84876 | 4.6761 | 4.632 | 13586.9375 | 16011.0204 | 0.99536 | 0.99629 | 8.007 | 9.0334 | 1.87 | 0.14699 | 7.8768 | 8.8948 | 8.1916 | 9.2128 | 0.31478 | 0.31798 | 0.071552 | 0.071183 | 89.8479 | 115.8681 | 0.00037405 | 0.00022925 | FALSE | 0.88638 | TRUE | [M-H]- | 104.8349 | 7.9373 | 8.0905 | 0.15325 | 7.9465 | 8.079 | 0.13242 | 16.1401 | 1.2525 | 1.552 | 7.2807 | 0 | 23 | 6 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 63514.2221 | No Control Sample | 0 | off | off | 63514.2221 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9539 | 3.2449 | 1.2913 | 77.9649 | -6.7278 | -0.52453 | 0 | 0 | 0 | 0 | 0 | 0.68409 | FALSE | 1 | 0.85685
3579 | 3580 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FHpSA | 3.82975905995247 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.58 | 8.86 | Baseline | Baseline | 120147.364 | 74125.3059 | 1.6209 | 26548.4234 | 16003.0031 | 1.659 | 4.5256 | 4.632 | 26589.7584 | 16011.0204 | 0.98901 | 0.99629 | 8.5918 | 9.0334 | 0.13776 | 0.01182 | 8.4559 | 8.8948 | 8.7397 | 9.2128 | 0.28375 | 0.31798 | 0.070743 | 0.071183 | 146.8918 | 115.8681 | 0.00015758 | 0.00022925 | FALSE | 0.95112 | TRUE | [M-H]- | 95.744 | 8.5176 | 8.666 | 0.1483 | 8.5281 | 8.6576 | 0.12949 | 14.7438 | 0.91734 | 0.84077 | 3.7156 | 0 | 17 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 120147.364 | No Control Sample | 0 | off | off | 120147.364 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9505 | 2.5371 | 1.1365 | 77.9649 | -6.1447 | -0.47907 | 0 | 0 | 0 | 0 | 0 | 1.0551 | FALSE | 1 | 1.6209
3581 | 3582 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FOSA | 4.16301012899506 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 8.88 | 8.86 | Baseline | Baseline | 70213.3085 | 74125.3059 | 0.94722 | 15142.1119 | 16003.0031 | 0.9462 | 4.637 | 4.632 | 15142.1119 | 16011.0204 | 0.99609 | 0.99629 | 9.0333 | 9.0334 | 1.7262 | 0.15328 | 8.7803 | 8.8948 | 9.3418 | 9.2128 | 0.5615 | 0.31798 | 0.071799 | 0.071183 | 221.6872 | 115.8681 | 8.0493e-12 | 0.00022925 | FALSE | 0.99999 | TRUE | [M-H]- | 104.0753 | 8.9532 | 9.1123 | 0.15908 | 8.9651 | 9.1019 | 0.13675 | -2.1707e-07 | 1.0233 | 1.0785 | 7.5733 | 0 | 29 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 70213.3085 | No Control Sample | 0 | off | off | 70213.3085 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9472 | 2.0252 | 1.0084 | 77.9651 | -4.3805 | -0.34153 | 0 | 0 | 0 | 0 | 0 | 1.78 | FALSE | 1 | 0.94722
3600 | 3601 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 4:2 FTS | 4.28684480025353 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 5.94 | 5.92 | Baseline | Baseline | 36375.4669 | 16711.7354 | 2.1766 | 6287.2672 | 2857.5734 | 2.2002 | 5.7856 | 5.8482 | 6290.8925 | 2913.1152 | 0.99588 | 0.9422 | 5.9649 | 5.9655 | 0.41839 | 0.024853 | 5.7814 | 5.8544 | 6.155 | 6.1388 | 0.37367 | 0.28442 | 0.087527 | 0.088386 | 185.6773 | 188.4649 | 0.0011132 | 0.03026 | FALSE | 0.9999 | TRUE | [M-H]- | 107.1711 | 5.8633 | 6.0625 | 0.1992 | 5.8798 | 6.0506 | 0.17073 | 18.7309 | 0.93339 | 0.89294 | 4.8104 | 0 | 44 | 12 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 36375.4669 | No Control Sample | 0 | off | off | 36375.4669 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9763 | 6.083 | 1.989 | 306.9691 | 4.8921 | 1.5017 | 0 | 0 | 0 | 0 | 0 | 5.3291 | FALSE | 1 | 2.1766
3602 | 3603 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 6:2 FTS | 3.83058436256091 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7.92 | 7.92 | Baseline | Baseline | 35782.8334 | 15396.3289 | 2.3241 | 8537.1304 | 3862.9662 | 2.21 | 4.1914 | 3.9856 | 8537.6341 | 3862.9662 | 0.99794 | 1 | 7.9418 | 7.9421 | 0.27527 | 0.021802 | 7.8005 | 7.7199 | 8.1316 | 8.1851 | 0.33102 | 0.46523 | 0.064248 | 0.059339 | 96.0264 | 77.2081 | 0.00013716 | 0 | FALSE | 0.99996 | TRUE | [M-H]- | 95.7646 | 7.8726 | 8.0168 | 0.14428 | 7.8822 | 8.0047 | 0.12254 | 3.5375 | 1.0371 | 1.0453 | 4.518 | 0 | 25 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 35782.8334 | No Control Sample | 0 | off | off | 35782.8334 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9699 | 4.5838 | 1.9571 | 406.9638 | 5.7447 | 2.3379 | 0 | 0 | 0 | 0 | 0 | 5.5018 | FALSE | 1 | 2.3241
3604 | 3605 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 8:2 FTS | 4.27040015155737 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 8.88 | 8.89 | Baseline | Baseline | 33969.553 | 16212.4131 | 2.0953 | 6743.9302 | 3145.1555 | 2.1442 | 5.0371 | 5.1547 | 6743.9302 | 3145.1555 | 0.99882 | 1 | 8.9035 | 8.9039 | 0.26485 | 0.023519 | 8.6678 | 8.6579 | 9.2081 | 9.159 | 0.54023 | 0.50107 | 0.078927 | 0.080537 | 118.3349 | 160.6063 | 0 | 0 | FALSE | 0.99996 | TRUE | [M-H]- | 106.76 | 8.8187 | 8.9884 | 0.16973 | 8.8295 | 8.9748 | 0.14536 | 0 | 0.9418 | 0.83234 | 3.7723 | 0 | 28 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Fail | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 33969.553 | No Control Sample | 0 | off | off | 33969.553 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9641 | 4.8449 | 2.5531 | 506.9573 | 3.2393 | 1.6422 | 0 | 0 | 0 | 0 | 0 | 5.7212 | FALSE | 1 | 2.0953
3606 | 3607 | PFAS CS6 4ng/ml | 26 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 27) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 10:2 FTS | 3.51929628989541 | Quantifiers | 10/27/2023 02:02:41 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.5 | 8.89 | Baseline | Baseline | 29143.1526 | 16212.4131 | 1.7976 | 7190.5838 | 3145.1555 | 2.2862 | 4.053 | 5.1547 | 7190.5838 | 3145.1555 | 1 | 1 | 9.5205 | 8.9039 | 0.21565 | 0.020487 | 9.3076 | 8.6579 | 9.7177 | 9.159 | 0.41012 | 0.50107 | 0.063854 | 0.080537 | 490.9148 | 160.6063 | 0 | 0 | FALSE | 1.0693 | TRUE | [M-H]- | 87.9824 | 9.4537 | 9.5872 | 0.13351 | 9.4631 | 9.5786 | 0.1155 | 0 | 0.94183 | 0.88046 | 4.5091 | 0 | 28 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 29143.1526 | No Control Sample | 0 | off | off | 29143.1526 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9571 | 3.101 | 1.9442 | 606.9514 | 5.7761 | 3.5058 | 0 | 0 | 0 | 0 | 0 | 1.6181 | FALSE | 1 | 1.7976
3667 | 3668 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFBA | 4.11685885399904 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 1.49 | 1.48 | Baseline | Baseline | 72849.3967 | 98650.3094 | 0.73846 | 10673.5854 | 14531.1712 | 0.73453 | 6.8252 | 6.7889 | 10804.7522 | 14752.2306 | 0.94263 | 0.94166 | 1.4652 | 1.467 | 1.6626 | 0.024773 | 1.3241 | 1.3305 | 1.6587 | 1.6397 | 0.3346 | 0.30925 | 0.1072 | 0.10922 | 195.7942 | 175.9467 | 0.024265 | 0.033897 | FALSE | 0.99878 | TRUE | [M-H]- | 102.9215 | 1.3802 | 1.619 | 0.23878 | 1.3885 | 1.5979 | 0.20938 | 774.0557 | 1.4796 | 1.8902 | 1.7557 | 0 | 67 | 23 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 72849.3967 | No Control Sample | 0 | off | off | 72849.3967 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9807 | 6.8772 | 1.4647 | 168.9891 | 0.84044 | 0.14203 | 0 | 0 | 0 | 0 | 0 | 1.6264 | FALSE | 1 | 0.73846
3669 | 3670 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFPeA | 3.87618500985129 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 3.18 | 3.21 | Manual | Manual | 65542.4944 | 75608.4676 | 0.86687 | 6749.7126 | 8186.3015 | 0.82451 | 9.7104 | 9.236 | 6765.0573 | 8211.2103 | 0.98459 | 0.98577 | 3.2022 | 3.2043 | 0.69967 | 0.022249 | 2.9995 | 2.9936 | 3.5181 | 3.476 | 0.51864 | 0.48235 | 0.1588 | 0.142 | 445.1358 | 576.2379 | 0.00063273 | 0.0012411 | TRUE | 0.99937 | TRUE | [M-H]- | 96.9046 | 3.0614 | 3.3986 | 0.33716 | 3.0809 | 3.3774 | 0.29647 | -8.2345 | 1.1301 | 1.2861 | 1.9997 | 0 | 90 | 30 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 65542.4944 | No Control Sample | 0 | off | off | 65542.4944 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9777 | 6.3249 | 1.6633 | 218.9865 | 3.0582 | 0.6697 | 0 | 0 | 0 | 0 | 0 | 0.20434 | FALSE | 1 | 0.86687
3671 | 3672 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFHxA | 3.95913827357092 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 68.3907 | 6.13 | 6.12 | Baseline | Baseline | 139410.0747 | 151302.2724 | 0.9214 | 25845.7331 | 26464.7157 | 0.97661 | 5.3939 | 5.7171 | 25878.7073 | 26471.8017 | 0.99051 | 0.99659 | 6.1537 | 6.1524 | 0.38626 | 0.023678 | 5.9669 | 5.9526 | 6.3251 | 6.3465 | 0.35812 | 0.3939 | 0.083565 | 0.087736 | 254.9732 | 287.0046 | 0.0019401 | 0.00051523 | FALSE | 1.0002 | TRUE | [M-H]- | 98.9785 | 6.0597 | 6.2444 | 0.18472 | 6.0722 | 6.233 | 0.16084 | 140.0178 | 1.0089 | 1.036 | 1.5172 | 0 | 39 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 139410.0747 | No Control Sample | 203843.5852 | off | off | 139410.0747 | 203843.5852 | off | 0.25 | off | 4 | off | 2.7356 | off | 312.9745 | 5.3238 | 1.6662 | 268.9832 | 3.1295 | 0.84179 | 0 | 0 | 0 | 0 | 0 | 45.6794 | FALSE | 1 | 0.9214
3673 | 3674 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFHpA | 4.47077653990078 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 81.1256 | 7.26 | 7.27 | Baseline | Baseline | 185233.1393 | 165579.7345 | 1.1187 | 45606.9024 | 38454.9145 | 1.186 | 4.0615 | 4.3058 | 45623.7127 | 38498.5957 | 0.99667 | 0.99295 | 7.2981 | 7.2969 | 0.52443 | 0.038073 | 7.1595 | 7.1527 | 7.4586 | 7.4356 | 0.29905 | 0.28295 | 0.06279 | 0.06698 | 254.1502 | 240.1562 | 0.00078674 | 0.0021797 | FALSE | 1.0002 | TRUE | [M-H]- | 111.7694 | 7.2294 | 7.3685 | 0.13914 | 7.2387 | 7.3583 | 0.11954 | 119.982 | 1.0354 | 1.0668 | 1.5587 | 0 | 35 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Pass | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 185233.1393 | No Control Sample | 228328.7703 | off | off | 185233.1393 | 228328.7703 | off | 0.25 | off | 4 | off | 3.245 | off | 362.9709 | 3.3938 | 1.2319 | 318.98 | 2.9666 | 0.94628 | 0 | 0 | 0 | 0 | 0 | 0.66139 | FALSE | 1 | 1.1187
3675 | 3676 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFOA | 4.22886169788416 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 86.7341 | 7.96 | 7.94 | Baseline | Baseline | 166074.419 | 161931.2028 | 1.0256 | 38735.565 | 38033.5784 | 1.0185 | 4.2874 | 4.2576 | 38765.9218 | 38033.6867 | 0.99261 | 0.99954 | 7.9804 | 7.9799 | 0.2557 | 0.020354 | 7.8628 | 7.803 | 8.1738 | 8.2448 | 0.31103 | 0.44175 | 0.066237 | 0.066036 | 124.6394 | 97.6109 | 1.096e-05 | 6.9227e-06 | FALSE | 1.0001 | TRUE | [M-H]- | 105.7215 | 7.916 | 8.0594 | 0.14339 | 7.9244 | 8.048 | 0.12358 | 1.3649 | 1.0313 | 1.0236 | 1.5436 | 0 | 23 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 166074.419 | No Control Sample | 191475.4056 | off | off | 166074.419 | 191475.4056 | off | 0.25 | off | 4 | off | 3.4694 | off | 412.9679 | 3.6337 | 1.5006 | 368.9771 | 2.6474 | 0.97683 | 0 | 0 | 0 | 0 | 0 | 0.60043 | FALSE | 1 | 1.0256
3677 | 3678 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFNA | 3.82102450187645 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 77.0704 | 8.46 | 8.47 | Baseline | Baseline | 122652.4934 | 138176.1994 | 0.88765 | 26073.0375 | 28476.9992 | 0.91558 | 4.7042 | 4.8522 | 26075.5064 | 28477.7023 | 0.99838 | 0.99849 | 8.4927 | 8.4901 | 0.38694 | 0.032735 | 8.302 | 8.2783 | 8.7774 | 8.7269 | 0.47543 | 0.44855 | 0.072816 | 0.076322 | 136.7545 | 170.2276 | 0.00022287 | 5.3175e-05 | FALSE | 1.0003 | TRUE | [M-H]- | 95.5256 | 8.412 | 8.5697 | 0.15778 | 8.4197 | 8.5589 | 0.13916 | 12.2223 | 0.95142 | 0.85193 | 1.4633 | 0 | 27 | 5 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 122652.4934 | No Control Sample | 159143.3885 | off | off | 122652.4934 | 159143.3885 | off | 0.25 | off | 4 | off | 3.0828 | off | 462.9651 | 3.9575 | 1.8322 | 418.9745 | 3.0116 | 1.2618 | 0 | 0 | 0 | 0 | 0 | 0.57744 | FALSE | 1 | 0.88765
3679 | 3680 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFDA | 4.23849289647315 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.87 | 8.86 | Baseline | Baseline | 113756.2325 | 108114.514 | 1.0522 | 23498.3493 | 21033.5641 | 1.1172 | 4.841 | 5.1401 | 23502.9224 | 21040.3585 | 0.99666 | 0.99768 | 8.9023 | 8.9027 | 0.36434 | 0.032317 | 8.7575 | 8.7286 | 9.0858 | 9.1132 | 0.32828 | 0.38457 | 0.075067 | 0.080177 | 179.3952 | 135.4678 | 0.00031028 | 0.00073734 | FALSE | 0.99996 | TRUE | [M-H]- | 105.9623 | 8.8226 | 8.9869 | 0.16431 | 8.8309 | 8.9766 | 0.14569 | 22.2094 | 1.0164 | 1.0086 | 1.2906 | 0 | 17 | 5 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 113756.2325 | No Control Sample | 0 | off | off | 113756.2325 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9623 | 4.4635 | 2.2896 | 468.9709 | 1.8715 | 0.87768 | 0 | 0 | 0 | 0 | 0 | 23.4678 | FALSE | 1 | 1.0522
3681 | 3682 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFUdA | 4.62826392316386 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 4 | 4 | 1 | 0 | 9.21 | 9.2 | Baseline | Baseline | 87986.0582 | 85723.4293 | 1.0264 | 19255.9016 | 18263.3182 | 1.0543 | 4.5693 | 4.6937 | 19274.9285 | 18265.1303 | 0.99296 | 0.99854 | 9.24 | 9.2374 | 0.32586 | 0.030012 | 9.0489 | 9.0759 | 9.3752 | 9.4644 | 0.3263 | 0.38852 | 0.071601 | 0.074303 | 227.4814 | 196.3496 | 0.0015542 | 0.00024603 | FALSE | 1.0003 | TRUE | [M-H]- | 115.7066 | 9.1668 | 9.3192 | 0.15232 | 9.1734 | 9.3079 | 0.13449 | 91.7165 | 0.99192 | 0.91462 | 1.5533 | 0 | 18 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 87986.0582 | No Control Sample | 0 | off | off | 87986.0582 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9595 | 4.6341 | 2.6088 | 518.9689 | 3.0475 | 1.5815 | 0 | 0 | 0 | 0 | 0 | 0.94326 | FALSE | 1 | 1.0264
3683 | 3684 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFDoA | 3.71214896586107 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.47 | 9.47 | Baseline | Baseline | 76697.9071 | 79841.8654 | 0.96062 | 18816.0115 | 19377.014 | 0.97105 | 4.0762 | 4.1204 | 18821.6709 | 19392.5585 | 0.9976 | 0.99427 | 9.5049 | 9.5061 | 0.36875 | 0.034921 | 9.3932 | 9.3653 | 9.6931 | 9.6293 | 0.29988 | 0.26397 | 0.063813 | 0.064265 | 287.6293 | 257.6181 | 0.00032659 | 0.0014122 | FALSE | 0.99987 | TRUE | [M-H]- | 92.8037 | 9.438 | 9.5735 | 0.13548 | 9.4464 | 9.5638 | 0.11738 | -20.4917 | 0.95471 | 0.87708 | 1.6268 | 0 | 21 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 76697.9071 | No Control Sample | 0 | off | off | 76697.9071 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9556 | 3.2395 | 1.9857 | 568.9658 | 4.1985 | 2.3888 | 0 | 0 | 0 | 0 | 0 | 104.3174 | FALSE | 1 | 0.96062
3685 | 3686 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFTrDA | 3.70980370671094 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 9.68 | 9.47 | Baseline | Baseline | 55872.0118 | 79841.8654 | 0.69978 | 15909.5265 | 19377.014 | 0.82105 | 3.5119 | 4.1204 | 15921.2599 | 19392.5585 | 0.99256 | 0.99427 | 9.7216 | 9.5061 | 0.42971 | 0.041596 | 9.5961 | 9.3653 | 9.8231 | 9.6293 | 0.22695 | 0.26397 | 0.053988 | 0.064265 | 538.1864 | 257.6181 | 0.001239 | 0.0014122 | FALSE | 1.0227 | TRUE | [M-H]- | 92.7451 | 9.6621 | 9.7824 | 0.12036 | 9.6707 | 9.7727 | 0.10198 | 86.8592 | 1.0838 | 1.1755 | 1.5202 | 0 | 20 | 6 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 55872.0118 | No Control Sample | 0 | off | off | 55872.0118 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9526 | 3.2214 | 2.1356 | 618.9621 | 0.65546 | 0.4057 | 0 | 0 | 0 | 0 | 0 | 0.91967 | FALSE | 1 | 0.69978
3687 | 3688 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFTeDA | 3.64198032962639 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 9.87 | 9.86 | Baseline | Baseline | 25140.4382 | 19469.8065 | 1.2913 | 7361.6489 | 5786.3254 | 1.2722 | 3.4151 | 3.3648 | 7366.9644 | 5792.8953 | 0.98927 | 0.99378 | 9.9 | 9.903 | 0.30389 | 0.029994 | 9.8038 | 9.816 | 10.0157 | 10.0027 | 0.21182 | 0.1867 | 0.052499 | 0.05326 | 623.6216 | 760.2198 | 0.00099693 | 0.0021939 | FALSE | 0.9997 | TRUE | [M-H]- | 91.0495 | 9.8466 | 9.9599 | 0.11332 | 9.854 | 9.9521 | 0.098139 | 34.6482 | 1.1863 | 1.4289 | 1.7379 | 0 | 24 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 25140.4382 | No Control Sample | 0 | off | off | 25140.4382 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9495 | 3.1521 | 2.2472 | 668.9608 | 4.3779 | 2.9286 | 0 | 0 | 0 | 0 | 0 | 112.2697 | FALSE | 1 | 1.2913
3694 | 3695 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFBS | 3.68559962547997 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 3.5 | 3.5 | Manual | Manual | 42620.6844 | 43689.6255 | 0.97553 | 4388.9588 | 3960.5028 | 1.1082 | 9.7109 | 11.0313 | 4398.7174 | 3973.3212 | 0.98908 | 0.98592 | 3.8468 | 3.8503 | 9.9081 | 0.34678 | 3.6153 | 3.5996 | 4.1592 | 4.1232 | 0.54391 | 0.52359 | 0.15231 | 0.16386 | 300.6842 | 737.5467 | 0.00086047 | 0.00093734 | TRUE | 0.99909 | TRUE | [M-H]- | 92.14 | 3.6771 | 4.0605 | 0.38338 | 3.7075 | 4.0172 | 0.30964 | -6.9434 | 1.1654 | 1.3089 | 12.4389 | 0 | 85 | 26 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 42620.6844 | No Control Sample | 0 | off | off | 42620.6844 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9444 | 4.7386 | 1.4166 | 79.9569 | -3.3751 | -0.26986 | 0 | 0 | 0 | 0 | 0 | 1.168 | FALSE | 1 | 0.97553
3696 | 3697 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFPeS | 4.30699658210358 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 6.38 | 3.5 | Baseline | Manual | 50149.544 | 43689.6255 | 1.1479 | 9907.0747 | 3960.5028 | 2.5015 | 5.062 | 11.0313 | 9927.0921 | 3973.3212 | 0.98576 | 0.98592 | 6.4047 | 3.8503 | 0.38694 | 0.024687 | 6.2471 | 3.5996 | 6.5418 | 4.1232 | 0.29468 | 0.52359 | 0.079723 | 0.16386 | 247.5719 | 737.5467 | 0.003831 | 0.00093734 | FALSE | 1.6634 | TRUE | [M-H]- | 107.6749 | 6.3229 | 6.4921 | 0.16924 | 6.3359 | 6.4796 | 0.14367 | 128.7968 | 1.0816 | 1.2042 | 13.151 | 0 | 32 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 50149.544 | No Control Sample | 0 | off | off | 50149.544 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9408 | 2.8239 | 0.98538 | 79.9567 | -6.0734 | -0.48561 | 0 | 0 | 0 | 0 | 0 | 1.3786 | FALSE | 1 | 1.1479
3698 | 3699 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | L-PFHxS | 2.99296932068414 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 3.244 | 4 | 0.811 | 0 | 7.34 | 7.35 | Baseline | Baseline | 34468.7518 | 61653.7787 | 0.55907 | 8373.1043 | 15053.0141 | 0.55624 | 4.1166 | 4.0958 | 8453.1463 | 15064.525 | 0.76476 | 0.99119 | 7.3728 | 7.3721 | 0.44743 | 0.032841 | 7.281 | 7.257 | 7.5053 | 7.5536 | 0.22428 | 0.29653 | 0.064659 | 0.06335 | 107.6726 | 196.0744 | 0.012674 | 0.00091686 | FALSE | 1.0001 | TRUE | [M-H]- | 92.2617 | 7.3065 | 7.4412 | 0.13474 | 7.3161 | 7.4331 | 0.11697 | -473.1652 | 1.1118 | 1.2951 | 17.3303 | 0 | 27 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 34468.7518 | No Control Sample | 0 | off | off | 34468.7518 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9378 | 3.0003 | 1.1969 | 79.9569 | -3.7758 | -0.3019 | 0 | 0 | 0 | 0 | 0 | 2.549 | FALSE | 0.811 | 0.55907
3700 | 3701 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | Br-PFHxS | 0.575507855500215 | Qualifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.756 | 4 | 0.189 | 0 | 7.16 | 7.35 | Manual | Baseline | 9004.8886 | 61653.7787 | 0.14606 | 1648.4709 | 15053.0141 | 0.10951 | 5.4626 | 4.0958 | 8453.1463 | 15064.525 | 0.19999 | 0.99119 | 7.1916 | 7.3721 | 0.4414 | 0.031605 | 7.0857 | 7.257 | 7.2794 | 7.5536 | 0.19374 | 0.29653 | 0.087645 | 0.06335 | 21.1982 | 196.0744 | 0.021366 | 0.00091686 | TRUE | 0.97551 | TRUE | [M-H]- | 76.1254 | 7.101 | 7.281 | 0.18009 | 7.1074 | 7.2681 | 0.16067 | 181.794 | 1.0138 | 0.95008 | 0.26125 | 0 | 22 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9004.8886 | No Control Sample | 0 | off | off | 34468.7518 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9378 | 3.0003 | 1.1969 | 79.9652 | 100.1001 | 8.0037 | 0 | 0 | 0 | 0 | 0 | 2.549 | FALSE | 0.189 | 0.14606
3702 | 3703 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFHpS | 4.60853559931082 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 4 | 4 | 1 | 0 | 7.98 | 7.35 | Baseline | Baseline | 43243.3997 | 61653.7787 | 0.70139 | 9934.9762 | 15053.0141 | 0.66 | 4.3526 | 4.0958 | 9935.2103 | 15064.525 | 0.99791 | 0.99119 | 8.0059 | 7.3721 | 0.32462 | 0.025905 | 7.879 | 7.257 | 8.2091 | 7.5536 | 0.3301 | 0.29653 | 0.067778 | 0.06335 | 107.0777 | 196.0744 | 3.7088e-05 | 0.00091686 | FALSE | 1.086 | TRUE | [M-H]- | 115.2134 | 7.9391 | 8.0832 | 0.14406 | 7.9476 | 8.0736 | 0.12599 | -1.1162 | 1.1496 | 1.3288 | 12.5925 | 0 | 24 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 43243.3997 | No Control Sample | 0 | off | off | 43243.3997 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9344 | 2.2869 | 1.0267 | 79.9568 | -5.422 | -0.43353 | 0 | 0 | 0 | 0 | 0 | 1.7042 | FALSE | 1 | 0.70139
3704 | 3705 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | L-PFOS | 3.071266156453 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 3.152 | 4 | 0.788 | 0 | 8.46 | 8.45 | Manual | Baseline | 30376.4252 | 38050.808 | 0.79831 | 6309.5304 | 8136.0338 | 0.7755 | 4.8144 | 4.6768 | 6374.0436 | 8144.8527 | 0.67803 | 0.99081 | 8.4899 | 8.4894 | 0.35386 | 0.029937 | 8.412 | 8.3114 | 8.589 | 8.6742 | 0.17697 | 0.3628 | 0.075573 | 0.07286 | 50.29 | 169.8594 | 0.010153 | 0.0021724 | TRUE | 1.0001 | TRUE | [M-H]- | 97.4386 | 8.3965 | 8.5642 | 0.16773 | 8.3965 | 8.5546 | 0.15813 | -361.985 | 0.81592 | 0.53849 | 15.7879 | 0 | 12 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 30376.4252 | No Control Sample | 0 | off | off | 30376.4252 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9316 | 2.6862 | 1.3402 | 79.9571 | -2.5961 | -0.20758 | 0 | 0 | 0 | 0 | 0 | 1.7664 | FALSE | 0.788 | 0.79831
3706 | 3707 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | Br-PFOS | 0.584518976279883 | Qualifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.844 | 4 | 0.211 | 0 | 8.3 | 8.45 | Manual | Baseline | 12376.5574 | 38050.808 | 0.32526 | 2052.6248 | 8136.0338 | 0.25229 | 6.0296 | 4.6768 | 6374.0436 | 8144.8527 | 0.27637 | 0.99081 | 8.3408 | 8.4894 | 0.4918 | 0.040819 | 8.23 | 8.3114 | 8.412 | 8.6742 | 0.18203 | 0.3628 | 0.10068 | 0.07286 | 16.3604 | 169.8594 | 0.011976 | 0.0021724 | TRUE | 0.9825 | TRUE | [M-H]- | 69.2558 | 8.2538 | 8.4153 | 0.16152 | 8.262 | 8.4153 | 0.1533 | 135.0428 | 0.93362 | 0.95821 | 0.40744 | 0 | 12 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 12376.5574 | No Control Sample | 0 | off | off | 30376.4252 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9316 | 2.6862 | 1.3402 | 79.9655 | 102.591 | 8.2029 | 0 | 0 | 0 | 0 | 0 | 1.7664 | FALSE | 0.211 | 0.32526
3708 | 3709 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFNS | 4.26177397975665 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.85 | 8.45 | Baseline | Baseline | 35943.0085 | 38050.808 | 0.94461 | 7612.6084 | 8136.0338 | 0.93567 | 4.7215 | 4.6768 | 7612.6124 | 8144.8527 | 0.993 | 0.99081 | 8.8884 | 8.4894 | 0.43378 | 0.03839 | 8.7242 | 8.3114 | 9.0712 | 8.6742 | 0.34693 | 0.3628 | 0.073971 | 0.07286 | 175.2727 | 169.8594 | 3.0805e-07 | 0.0021724 | FALSE | 1.047 | TRUE | [M-H]- | 106.5443 | 8.8107 | 8.9695 | 0.15881 | 8.8228 | 8.9609 | 0.13816 | 0.0067593 | 1.0364 | 1.141 | 15.8309 | 0 | 18 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 35943.0085 | No Control Sample | 0 | off | off | 35943.0085 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9278 | 1.381 | 0.75805 | 79.9568 | -4.3657 | -0.34907 | 0 | 0 | 0 | 0 | 0 | 1.6857 | FALSE | 1 | 0.94461
3710 | 3711 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFDS | 3.97150539619593 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.19 | 8.45 | Baseline | Baseline | 34414.6431 | 38050.808 | 0.90444 | 7508.7891 | 8136.0338 | 0.92291 | 4.5832 | 4.6768 | 7511.7024 | 8144.8527 | 0.99187 | 0.99081 | 9.2157 | 8.4894 | 0.27958 | 0.025694 | 9.0533 | 8.3114 | 9.3796 | 8.6742 | 0.32632 | 0.3628 | 0.071985 | 0.07286 | 205.8902 | 169.8594 | 0.00078903 | 0.0021724 | FALSE | 1.0856 | TRUE | [M-H]- | 99.2876 | 9.1441 | 9.2966 | 0.15253 | 9.1535 | 9.2848 | 0.13132 | 18.1562 | 1.1825 | 1.3855 | 14.7444 | 0 | 18 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 34414.6431 | No Control Sample | 0 | off | off | 34414.6431 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9249 | 1.7208 | 1.0306 | 79.9567 | -2.2737 | -0.18179 | 0 | 0 | 0 | 0 | 0 | 2.1134 | FALSE | 1 | 0.90444
3714 | 3715 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FBSA | 3.26226010033251 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 4 | 4 | 1 | 0 | 4.6 | 8.86 | Manual | Baseline | 59506.3896 | 69178.6384 | 0.86018 | 5570.7017 | 14179.7203 | 0.39286 | 10.682 | 4.8787 | 5640.7233 | 14185.326 | 0.9553 | 0.99494 | 4.9615 | 9.0311 | 7.858 | 0.36147 | 4.7409 | 8.8118 | 5.1994 | 9.2142 | 0.45854 | 0.40233 | 0.17353 | 0.075443 | 201.5586 | 140.1391 | 0.0050883 | 0.00074763 | TRUE | 0.54938 | TRUE | [M-H]- | 81.5565 | 4.7821 | 5.1629 | 0.38081 | 4.7996 | 5.1348 | 0.33522 | 61.8161 | 1.0915 | 1.1366 | 5.6734 | 0 | 82 | 33 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 59506.3896 | No Control Sample | 0 | off | off | 59506.3896 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9607 | 5.7384 | 1.7098 | 77.965 | -3.2974 | -0.25708 | 0 | 0 | 0 | 0 | 0 | 1.0531 | FALSE | 1 | 0.86018
3716 | 3717 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FPeSA | 4.11378981811313 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7 | 8.86 | Baseline | Baseline | 116528.8472 | 69178.6384 | 1.6845 | 25006.1595 | 14179.7203 | 1.7635 | 4.66 | 4.8787 | 25026.812 | 14185.326 | 0.99218 | 0.99494 | 7.1221 | 9.0311 | 1.7444 | 0.12211 | 6.9684 | 8.8118 | 7.3061 | 9.2142 | 0.33765 | 0.40233 | 0.071711 | 0.075443 | 296.9238 | 140.1391 | 0.0017934 | 0.00074763 | FALSE | 0.78862 | TRUE | [M-H]- | 102.8447 | 7.0441 | 7.2024 | 0.15823 | 7.055 | 7.1903 | 0.13532 | 132.8163 | 1.009 | 1.0026 | 3.9252 | 0 | 37 | 9 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 116528.8472 | No Control Sample | 0 | off | off | 116528.8472 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9568 | 2.8405 | 0.98837 | 77.9648 | -6.6384 | -0.51757 | 0 | 0 | 0 | 0 | 0 | 1.5792 | FALSE | 1 | 1.6845
3718 | 3719 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FHxSA | 4.48571783831132 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7.86 | 8.86 | Baseline | Baseline | 63407.8007 | 69178.6384 | 0.91658 | 13869.599 | 14179.7203 | 0.97813 | 4.5717 | 4.8787 | 13880.4364 | 14185.326 | 0.99217 | 0.99494 | 8.0089 | 9.0311 | 1.8941 | 0.14887 | 7.8636 | 8.8118 | 8.1755 | 9.2142 | 0.3119 | 0.40233 | 0.071032 | 0.075443 | 113.776 | 140.1391 | 0.0006999 | 0.00074763 | FALSE | 0.88682 | TRUE | [M-H]- | 112.1429 | 7.9402 | 8.0885 | 0.14834 | 7.949 | 8.0779 | 0.1289 | 31.1231 | 0.98211 | 0.93396 | 6.0704 | 0 | 23 | 6 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 63407.8007 | No Control Sample | 0 | off | off | 63407.8007 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9539 | 3.2681 | 1.3005 | 77.9649 | -6.1751 | -0.48144 | 0 | 0 | 0 | 0 | 0 | 2.2114 | FALSE | 1 | 0.91658
3720 | 3721 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FHpSA | 4.14395797947232 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.58 | 8.86 | Baseline | Baseline | 121328.7292 | 69178.6384 | 1.7538 | 26455.5105 | 14179.7203 | 1.8657 | 4.5861 | 4.8787 | 26463.2967 | 14185.326 | 0.99519 | 0.99494 | 8.5906 | 9.0311 | 0.12326 | 0.010575 | 8.3572 | 8.8118 | 8.78 | 9.2142 | 0.42282 | 0.40233 | 0.070695 | 0.075443 | 133.3493 | 140.1391 | 0.00054588 | 0.00074763 | FALSE | 0.95123 | TRUE | [M-H]- | 103.5989 | 8.5168 | 8.666 | 0.14924 | 8.5263 | 8.6573 | 0.13101 | 34.1555 | 1.1382 | 1.3378 | 3.7085 | 0 | 24 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 121328.7292 | No Control Sample | 0 | off | off | 121328.7292 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9504 | 2.217 | 0.99312 | 77.9648 | -7.7173 | -0.60168 | 0 | 0 | 0 | 0 | 0 | 1.1552 | FALSE | 1 | 1.7538
3722 | 3723 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FOSA | 4.54775700537307 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.88 | 8.86 | Baseline | Baseline | 71583.7974 | 69178.6384 | 1.0348 | 15690.6944 | 14179.7203 | 1.1066 | 4.5622 | 4.8787 | 15709.8775 | 14185.326 | 0.9847 | 0.99494 | 9.0312 | 9.0311 | 1.7026 | 0.15119 | 8.8856 | 8.8118 | 9.1677 | 9.2142 | 0.28208 | 0.40233 | 0.070012 | 0.075443 | 126.0909 | 140.1391 | 0.0022003 | 0.00074763 | FALSE | 1 | TRUE | [M-H]- | 113.6939 | 8.9514 | 9.1083 | 0.15688 | 8.9611 | 9.0994 | 0.13835 | 122.3903 | 0.97347 | 0.95105 | 7.497 | 0 | 15 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 71583.7974 | No Control Sample | 0 | off | off | 71583.7974 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9469 | 1.4655 | 0.72973 | 77.965 | -5.5571 | -0.43326 | 0 | 0 | 0 | 0 | 0 | 1.4442 | FALSE | 1 | 1.0348
3741 | 3742 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 4:2 FTS | 3.99608560847874 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 5.94 | 5.92 | Baseline | Baseline | 36421.635 | 17950.452 | 2.029 | 6334.8021 | 2845.4938 | 2.2263 | 5.7495 | 6.3084 | 6347.9805 | 2848.2804 | 0.98884 | 0.99308 | 5.9588 | 5.9585 | 0.3168 | 0.018818 | 5.7929 | 5.7938 | 6.1255 | 6.1282 | 0.33267 | 0.3344 | 0.087149 | 0.093626 | 216.3562 | 138.0862 | 0.002249 | 0.0020765 | FALSE | 1.0001 | TRUE | [M-H]- | 99.9021 | 5.8551 | 6.0519 | 0.1968 | 5.8699 | 6.0399 | 0.17001 | 42.827 | 0.93666 | 0.8826 | 4.8245 | 0 | 39 | 12 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 36421.635 | No Control Sample | 0 | off | off | 36421.635 | 0 | off | 0 | off | 0 | off | 0 | off | 326.976 | 5.2967 | 1.7319 | 306.9689 | 4.2257 | 1.2972 | 0 | 0 | 0 | 0 | 0 | 5.3291 | FALSE | 1 | 2.029
3743 | 3744 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 6:2 FTS | 4.13233363995829 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7.92 | 7.92 | Baseline | Baseline | 37503.2394 | 14958.2535 | 2.5072 | 9087.8779 | 3493.9846 | 2.601 | 4.1267 | 4.2811 | 9092.2127 | 3496.2647 | 0.99711 | 0.99415 | 7.941 | 7.9426 | 0.26536 | 0.021016 | 7.8009 | 7.8056 | 8.0839 | 8.0587 | 0.28298 | 0.2531 | 0.062162 | 0.066992 | 94.3129 | 93.6625 | 0.00093105 | 0.0011312 | FALSE | 0.9998 | TRUE | [M-H]- | 103.3083 | 7.8729 | 8.019 | 0.14614 | 7.8832 | 8.0065 | 0.12329 | 29.9001 | 1.0444 | 1.068 | 4.496 | 0 | 22 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 37503.2394 | No Control Sample | 0 | off | off | 37503.2394 | 0 | off | 0 | off | 0 | off | 0 | off | 426.97 | 4.8323 | 2.0632 | 406.963 | 3.7859 | 1.5407 | 0 | 0 | 0 | 0 | 0 | 5.5018 | FALSE | 1 | 2.5072
3745 | 3746 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 8:2 FTS | 3.91790696588672 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.88 | 8.89 | Baseline | Baseline | 34695.2439 | 18048.5463 | 1.9223 | 6673.5056 | 3665.5621 | 1.8206 | 5.199 | 4.9238 | 6679.1956 | 3665.5621 | 0.99435 | 0.99884 | 8.8985 | 8.8999 | 0.20822 | 0.01849 | 8.7439 | 8.6586 | 9.0541 | 9.1402 | 0.3103 | 0.48155 | 0.081643 | 0.07711 | 138.3848 | 181.5892 | 0.001789 | 0 | FALSE | 0.99984 | TRUE | [M-H]- | 97.9477 | 8.8168 | 8.988 | 0.17118 | 8.8282 | 8.9735 | 0.14537 | 38.4764 | 1.1842 | 1.3877 | 3.798 | 0 | 16 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 34695.2439 | No Control Sample | 0 | off | off | 34695.2439 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9639 | 4.4942 | 2.3683 | 506.9574 | 3.4147 | 1.7311 | 0 | 0 | 0 | 0 | 0 | 5.7212 | FALSE | 1 | 1.9223
3747 | 3748 | PFAS CS6 4ng/ml | 27 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 28) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 10:2 FTS | 2.98838153914577 | Quantifiers | 10/27/2023 02:19:16 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 4 | 4 | 1 | 0 | 9.5 | 8.89 | Baseline | Baseline | 27549.3471 | 18048.5463 | 1.5264 | 6558.6585 | 3665.5621 | 1.7893 | 4.2005 | 4.9238 | 6558.6585 | 3665.5621 | 0.9967 | 0.99884 | 9.5193 | 8.8999 | 0.20279 | 0.019265 | 9.329 | 8.6586 | 9.6833 | 9.1402 | 0.35428 | 0.48155 | 0.065642 | 0.07711 | 213.3988 | 181.5892 | 6.2039e-13 | 0 | FALSE | 1.0696 | TRUE | [M-H]- | 74.7095 | 9.4511 | 9.5905 | 0.13945 | 9.4602 | 9.5809 | 0.12076 | 1.1485e-08 | 1.146 | 1.3332 | 4.715 | 0 | 24 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 27549.3471 | No Control Sample | 0 | off | off | 27549.3471 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9573 | 3.5022 | 2.1958 | 606.9531 | 8.6062 | 5.2235 | 0 | 0 | 0 | 0 | 0 | 1.6135 | FALSE | 1 | 1.5264
3808 | 3809 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFBA | 4.06285174749355 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 1.49 | 1.48 | Baseline | Baseline | 76448.036 | 104899.5987 | 0.72877 | 10504.0596 | 15130.1532 | 0.69425 | 7.278 | 6.9331 | 10552.8756 | 15169.9371 | 0.97308 | 0.98489 | 1.4673 | 1.4682 | 1.5236 | 0.022702 | 1.3467 | 1.3327 | 1.6916 | 1.6776 | 0.34487 | 0.34487 | 0.11897 | 0.10034 | 174.6763 | 226.5145 | 0.0021673 | 0.0066084 | FALSE | 0.99939 | TRUE | [M-H]- | 101.5713 | 1.3787 | 1.6256 | 0.24689 | 1.3862 | 1.6006 | 0.21447 | 66.0109 | 1.5531 | 1.9764 | 1.6629 | 0 | 69 | 25 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 76448.036 | No Control Sample | 0 | off | off | 76448.036 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9807 | 6.9919 | 1.4891 | 168.9891 | 0.32462 | 0.054857 | 0 | 0 | 0 | 0 | 0 | 2.9915 | FALSE | 1 | 0.72877
3810 | 3811 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFPeA | 3.86043598523104 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 3.18 | 3.21 | Baseline | Manual | 63793.9863 | 73891.6482 | 0.86335 | 7040.2072 | 7859.1465 | 0.8958 | 9.0614 | 9.402 | 7121.117 | 7887.479 | 0.95435 | 0.98292 | 3.2159 | 3.2161 | 1.1305 | 0.035949 | 3 | 3.0052 | 3.4388 | 3.49 | 0.43882 | 0.48479 | 0.13711 | 0.14919 | 452.9808 | 587.449 | 0.016818 | 0.0018754 | FALSE | 0.99996 | TRUE | [M-H]- | 96.5109 | 3.0679 | 3.398 | 0.3301 | 3.0862 | 3.3658 | 0.27961 | 269.8149 | 1.1238 | 1.1758 | 2.1119 | 0 | 77 | 26 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 63793.9863 | No Control Sample | 0 | off | off | 63793.9863 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9775 | 5.7718 | 1.5178 | 218.9859 | 0.61358 | 0.13437 | 0 | 0 | 0 | 0 | 0 | 1.4807 | FALSE | 1 | 0.86335
3812 | 3813 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFHxA | 4.0364607071337 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 68.5121 | 6.13 | 6.12 | Baseline | Baseline | 139657.4779 | 148667.2905 | 0.9394 | 24405.2869 | 27558.9902 | 0.88557 | 5.7224 | 5.3945 | 24461.2884 | 27595.2156 | 0.98586 | 0.99203 | 6.1338 | 6.1338 | 0.061297 | 0.0037575 | 5.9605 | 5.9631 | 6.3002 | 6.2934 | 0.33975 | 0.33035 | 0.090473 | 0.082087 | 254.0008 | 243.2548 | 0.0040902 | 0.002382 | FALSE | 1 | TRUE | [M-H]- | 100.9115 | 6.0424 | 6.2271 | 0.18472 | 6.055 | 6.2156 | 0.16061 | 293.8092 | 1.1043 | 1.2596 | 1.4805 | 0 | 37 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 139657.4779 | No Control Sample | 203843.5852 | off | off | 139657.4779 | 203843.5852 | off | 0.25 | off | 4 | off | 2.7405 | off | 312.9743 | 4.899 | 1.5332 | 268.9833 | 3.3367 | 0.89752 | 0 | 0 | 0 | 0 | 0 | 46.5466 | FALSE | 1 | 0.9394
3814 | 3815 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFHpA | 4.53824736405035 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 78.8833 | 7.26 | 7.27 | Baseline | Baseline | 180113.1644 | 158609.3373 | 1.1356 | 41885.7721 | 36258.7286 | 1.1552 | 4.3001 | 4.3744 | 41899.8852 | 36289.3632 | 0.99566 | 0.99438 | 7.2741 | 7.2757 | 0.1937 | 0.014062 | 7.136 | 7.1292 | 7.4379 | 7.4235 | 0.3019 | 0.29432 | 0.067694 | 0.068188 | 258.2489 | 273.6108 | 0.00055344 | 0.0017 | FALSE | 0.99977 | TRUE | [M-H]- | 113.4562 | 7.2046 | 7.3467 | 0.14214 | 7.2143 | 7.337 | 0.1227 | 76.784 | 0.99733 | 0.99208 | 1.5973 | 0 | 35 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Pass | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 180113.1644 | No Control Sample | 228328.7703 | off | off | 180113.1644 | 228328.7703 | off | 0.25 | off | 4 | off | 3.1553 | off | 362.971 | 3.6854 | 1.3377 | 318.9798 | 2.242 | 0.71516 | 0 | 0 | 0 | 0 | 0 | 0.17404 | FALSE | 1 | 1.1356
3816 | 3817 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFOA | 4.1707045388866 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 85.7356 | 7.96 | 7.94 | Baseline | Baseline | 164162.5311 | 162299.0198 | 1.0115 | 38064.3842 | 36986.4287 | 1.0291 | 4.3128 | 4.3881 | 38102.1135 | 36993.193 | 0.99138 | 0.99756 | 7.9613 | 7.961 | 0.015707 | 0.0012503 | 7.8472 | 7.8137 | 8.1735 | 8.2127 | 0.32623 | 0.39902 | 0.065833 | 0.066704 | 99.1915 | 107.8814 | 0.00088541 | 0.00050922 | FALSE | 1 | TRUE | [M-H]- | 104.2676 | 7.8928 | 8.0396 | 0.14684 | 7.9025 | 8.0276 | 0.12505 | -103.3086 | 1.0859 | 1.1625 | 1.5497 | 0 | 24 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 164162.5311 | No Control Sample | 191475.4056 | off | off | 164162.5311 | 191475.4056 | off | 0.25 | off | 4 | off | 3.4294 | off | 412.9679 | 3.5348 | 1.4598 | 368.977 | 2.3971 | 0.88448 | 0 | 0 | 0 | 0 | 0 | 1.1348 | FALSE | 1 | 1.0115
3818 | 3819 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFNA | 4.26380987663093 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 84.5208 | 8.46 | 8.47 | Baseline | Baseline | 134509.2983 | 135797.3077 | 0.99052 | 28990.9951 | 28481.5274 | 1.0179 | 4.6397 | 4.7679 | 29012.6599 | 28485.3818 | 0.99434 | 0.99835 | 8.474 | 8.4709 | 0.16495 | 0.013954 | 8.302 | 8.2623 | 8.6133 | 8.653 | 0.31128 | 0.39067 | 0.07283 | 0.075095 | 184.4818 | 183.3481 | 0.001281 | 0.00025546 | FALSE | 1.0004 | TRUE | [M-H]- | 106.5952 | 8.3954 | 8.5484 | 0.15295 | 8.4038 | 8.5397 | 0.13595 | 119.3047 | 0.9669 | 0.92275 | 1.3906 | 0 | 18 | 5 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 134509.2983 | No Control Sample | 159143.3885 | off | off | 134509.2983 | 159143.3885 | off | 0.25 | off | 4 | off | 3.3808 | off | 462.9655 | 4.7941 | 2.2195 | 418.9742 | 2.2862 | 0.95787 | 0 | 0 | 0 | 0 | 0 | 1.3662 | FALSE | 1 | 0.99052
3820 | 3821 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFDA | 4.01580705797313 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.87 | 8.86 | Baseline | Baseline | 123764.7561 | 124149.3389 | 0.9969 | 26355.8095 | 24787.8101 | 1.0633 | 4.6959 | 5.0085 | 26355.8906 | 24796.6005 | 0.99922 | 0.99741 | 8.8847 | 8.8844 | 0.16623 | 0.014745 | 8.6824 | 8.6903 | 9.1472 | 9.0582 | 0.46482 | 0.36792 | 0.072946 | 0.078331 | 202.6501 | 130.1223 | 6.8957e-06 | 0.00063872 | FALSE | 1 | TRUE | [M-H]- | 100.3952 | 8.8072 | 8.9669 | 0.15963 | 8.82 | 8.9577 | 0.1377 | 0.391 | 1.0557 | 1.1907 | 1.1887 | 0 | 24 | 5 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 123764.7561 | No Control Sample | 0 | off | off | 123764.7561 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9623 | 4.4344 | 2.2747 | 468.9709 | 1.936 | 0.90793 | 0 | 0 | 0 | 0 | 0 | 22.2024 | FALSE | 1 | 0.9969
3822 | 3823 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFUdA | 4.47315709484911 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.21 | 9.2 | Baseline | Baseline | 88822.9865 | 89539.5715 | 0.992 | 19345.0437 | 19634.7388 | 0.98525 | 4.5915 | 4.5603 | 19345.6382 | 19637.7453 | 0.99807 | 0.9985 | 9.2212 | 9.218 | 0.12126 | 0.011168 | 9.0509 | 9.0591 | 9.4755 | 9.4316 | 0.42467 | 0.37257 | 0.071199 | 0.071534 | 181.3917 | 162.9661 | 7.4269e-05 | 0.00035658 | FALSE | 1.0003 | TRUE | [M-H]- | 111.8289 | 9.1486 | 9.3007 | 0.15207 | 9.1557 | 9.2888 | 0.1331 | 3.3832 | 0.98655 | 0.90182 | 1.5498 | 0 | 24 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 88822.9865 | No Control Sample | 0 | off | off | 88822.9865 | 0 | off | 0 | off | 0 | off | 0 | off | 562.959 | 3.7462 | 2.109 | 518.9684 | 2.0725 | 1.0756 | 0 | 0 | 0 | 0 | 0 | 0.32614 | FALSE | 1 | 0.992
3824 | 3825 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFDoA | 3.41057569768022 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.47 | 9.47 | Baseline | Baseline | 68406.7952 | 77507.5679 | 0.88258 | 16221.868 | 19097.3108 | 0.84943 | 4.2169 | 4.0586 | 16222.0647 | 19097.3666 | 0.99849 | 0.99823 | 9.484 | 9.4858 | 0.14788 | 0.014004 | 9.3064 | 9.3129 | 9.7404 | 9.7207 | 0.43398 | 0.40773 | 0.065428 | 0.063593 | 214.3503 | 251.8625 | 3.0587e-06 | 4.8299e-06 | FALSE | 0.99982 | TRUE | [M-H]- | 85.2644 | 9.4181 | 9.5573 | 0.13916 | 9.4279 | 9.5475 | 0.1196 | 0.11433 | 1.1938 | 1.464 | 2.052 | 0 | 30 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 68406.7952 | No Control Sample | 0 | off | off | 68406.7952 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9558 | 3.4459 | 2.1122 | 568.9661 | 4.8296 | 2.7479 | 0 | 0 | 0 | 0 | 0 | 101.0462 | FALSE | 1 | 0.88258
3826 | 3827 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFTrDA | 3.85586397669999 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.68 | 9.47 | Baseline | Baseline | 56373.9567 | 77507.5679 | 0.72733 | 14909.5194 | 19097.3108 | 0.78071 | 3.7811 | 4.0586 | 14914.941 | 19097.3666 | 0.99467 | 0.99823 | 9.7035 | 9.4858 | 0.24242 | 0.023466 | 9.5673 | 9.3129 | 9.8639 | 9.7207 | 0.29663 | 0.40773 | 0.059447 | 0.063593 | 518.4477 | 251.8625 | 0.00060768 | 4.8299e-06 | FALSE | 1.023 | TRUE | [M-H]- | 96.3966 | 9.6418 | 9.763 | 0.12126 | 9.6486 | 9.7567 | 0.10812 | 30.5433 | 0.94377 | 0.88397 | 1.1793 | 0 | 26 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 56373.9567 | No Control Sample | 0 | off | off | 56373.9567 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9527 | 3.4264 | 2.2715 | 618.963 | 2.0564 | 1.2729 | 0 | 0 | 0 | 0 | 0 | 4.4596 | FALSE | 1 | 0.72733
3828 | 3829 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFTeDA | 3.56021444178906 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 9.87 | 9.86 | Baseline | Baseline | 24978.2577 | 19788.4771 | 1.2623 | 7707.4864 | 5726.1262 | 1.346 | 3.2408 | 3.4558 | 7712.5072 | 5739.5675 | 0.986 | 0.98564 | 9.8783 | 9.881 | 0.084173 | 0.0083079 | 9.7943 | 9.7863 | 9.994 | 9.9622 | 0.19975 | 0.17592 | 0.049495 | 0.053631 | 658.1035 | 687.9749 | 0.00046709 | 0.0040946 | FALSE | 0.99973 | TRUE | [M-H]- | 89.0054 | 9.8277 | 9.9412 | 0.11346 | 9.8348 | 9.9329 | 0.098188 | 18.0231 | 1.2029 | 1.4469 | 1.7695 | 0 | 22 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 24978.2577 | No Control Sample | 0 | off | off | 24978.2577 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9504 | 4.413 | 3.1463 | 668.9603 | 3.6101 | 2.415 | 0 | 0 | 0 | 0 | 0 | 97.3756 | FALSE | 1 | 1.2623
3835 | 3836 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFBS | 3.86350817970953 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 3.5 | 3.5 | Manual | Manual | 42768.5854 | 41822.4149 | 1.0226 | 4113.8119 | 4054.7985 | 1.0146 | 10.3963 | 10.3143 | 4123.1825 | 4068.9779 | 0.98931 | 0.98407 | 3.8574 | 3.8592 | 10.2101 | 0.35735 | 3.6107 | 3.6421 | 4.1764 | 4.1552 | 0.56578 | 0.51308 | 0.16728 | 0.15804 | 296.5572 | 840.9207 | 0 | 0 | TRUE | 0.99953 | TRUE | [M-H]- | 96.5877 | 3.6893 | 4.0618 | 0.37251 | 3.7024 | 4.0333 | 0.33092 | 0 | 1.2346 | 1.4021 | 12.3666 | 0 | 88 | 28 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 42768.5854 | No Control Sample | 0 | off | off | 42768.5854 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9445 | 4.9598 | 1.4827 | 79.9568 | -5.1594 | -0.41253 | 0 | 0 | 0 | 0 | 0 | 0.9867 | FALSE | 1 | 1.0226
3837 | 3838 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFPeS | 4.39938227035024 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Total Width | 0 | 0 | 4 | 4 | 1 | 0 | 6.38 | 3.5 | Baseline | Manual | 49035.9898 | 41822.4149 | 1.1725 | 9831.0874 | 4054.7985 | 2.4246 | 4.9879 | 10.3143 | 9867.5308 | 4068.9779 | 0.98052 | 0.98407 | 6.3848 | 3.8592 | 0.075553 | 0.0048203 | 6.2607 | 3.6421 | 6.5082 | 4.1552 | 0.24747 | 0.51308 | 0.078475 | 0.15804 | 271.3143 | 840.9207 | 0.0058681 | 0 | FALSE | 1.6545 | TRUE | [M-H]- | 109.9846 | 6.3049 | 6.4698 | 0.16497 | 6.3157 | 6.4581 | 0.14235 | 233.1204 | 1.0165 | 1.0261 | 13.4815 | 0 | 27 | 11 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 49035.9898 | No Control Sample | 0 | off | off | 49035.9898 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9407 | 2.6832 | 0.93627 | 79.9568 | -5.2716 | -0.4215 | 0 | 0 | 0 | 0 | 0 | 1.6419 | FALSE | 1 | 1.1725
3839 | 3840 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | L-PFHxS | 3.12122586348818 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 3.244 | 4 | 0.811 | 0 | 7.34 | 7.35 | Baseline | Baseline | 36133.4695 | 61975.6192 | 0.58303 | 9214.7937 | 15366.5486 | 0.59967 | 3.9212 | 4.0332 | 9346.4579 | 15375.9169 | 0.76112 | 0.9932 | 7.35 | 7.3477 | 0.13603 | 0.0099844 | 7.2517 | 7.2155 | 7.4741 | 7.5068 | 0.22235 | 0.29128 | 0.061224 | 0.06239 | 127.9913 | 221.5978 | 0.016897 | 0.0013251 | FALSE | 1.0003 | TRUE | [M-H]- | 96.2153 | 7.2874 | 7.4151 | 0.12769 | 7.2952 | 7.4075 | 0.11222 | -700.2377 | 1.0824 | 1.1942 | 16.7218 | 0 | 27 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 36133.4695 | No Control Sample | 0 | off | off | 36133.4695 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9377 | 2.6951 | 1.0752 | 79.9567 | -6.394 | -0.51125 | 0 | 0 | 0 | 0 | 0 | 3.4579 | FALSE | 0.811 | 0.58303
3841 | 3842 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | Br-PFHxS | 0.573395329759308 | Qualifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.756 | 4 | 0.189 | 0 | 7.16 | 7.35 | Manual | Baseline | 9018.6684 | 61975.6192 | 0.14552 | 1730.2183 | 15366.5486 | 0.1126 | 5.2124 | 4.0332 | 9346.4579 | 15375.9169 | 0.19003 | 0.9932 | 7.1633 | 7.3477 | 0.045998 | 0.0032935 | 7.0613 | 7.2155 | 7.2586 | 7.5068 | 0.19733 | 0.29128 | 0.083314 | 0.06239 | 24.0323 | 221.5978 | 0.03001 | 0.0013251 | TRUE | 0.9749 | TRUE | [M-H]- | 75.8459 | 7.082 | 7.2603 | 0.17831 | 7.0916 | 7.2466 | 0.15496 | 263.1263 | 1.057 | 1.0732 | 0.24959 | 0 | 23 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 9018.6684 | No Control Sample | 0 | off | off | 36133.4695 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9377 | 2.6951 | 1.0752 | 79.9631 | 73.2983 | 5.8607 | 0 | 0 | 0 | 0 | 0 | 3.4579 | FALSE | 0.189 | 0.14552
3843 | 3844 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFHpS | 4.40974076372719 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7.98 | 7.35 | Baseline | Baseline | 41594.0412 | 61975.6192 | 0.67114 | 9855.3552 | 15366.5486 | 0.64135 | 4.2205 | 4.0332 | 9861.2246 | 15375.9169 | 0.99307 | 0.9932 | 7.9823 | 7.3477 | 0.028899 | 0.0023061 | 7.8635 | 7.2155 | 8.145 | 7.5068 | 0.28155 | 0.29128 | 0.06552 | 0.06239 | 146.0214 | 221.5978 | 0.0013189 | 0.0013251 | FALSE | 1.0864 | TRUE | [M-H]- | 110.2435 | 7.9202 | 8.0616 | 0.1414 | 7.9288 | 8.0501 | 0.12135 | 46.1679 | 1.0716 | 1.1126 | 14.8194 | 0 | 21 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 41594.0412 | No Control Sample | 0 | off | off | 41594.0412 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9343 | 1.966 | 0.88262 | 79.9569 | -3.5112 | -0.28074 | 0 | 0 | 0 | 0 | 0 | 1.2026 | FALSE | 1 | 0.67114
3845 | 3846 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | L-PFOS | 2.77498070475876 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 3.152 | 4 | 0.788 | 0 | 8.46 | 8.45 | Manual | Baseline | 28466.2385 | 39465.2438 | 0.7213 | 5975.1209 | 8363.8018 | 0.7144 | 4.7641 | 4.7186 | 6005.2983 | 8376.1095 | 0.6714 | 0.98896 | 8.4716 | 8.471 | 0.13735 | 0.01162 | 8.3992 | 8.3115 | 8.5666 | 8.607 | 0.16738 | 0.2956 | 0.075327 | 0.074795 | 45.2941 | 211.8907 | 0 | 0.002545 | TRUE | 1.0001 | TRUE | [M-H]- | 88.0387 | 8.3963 | 8.549 | 0.15263 | 8.3963 | 8.5398 | 0.14343 | 0 | 0.92483 | 0.73818 | 16.9371 | 0 | 10 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 28466.2385 | No Control Sample | 0 | off | off | 28466.2385 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9315 | 2.5908 | 1.2926 | 79.9568 | -6.0506 | -0.48379 | 0 | 0 | 0 | 0 | 0 | 1.6324 | FALSE | 0.788 | 0.7213
3847 | 3848 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | Br-PFOS | 0.554619957963015 | Qualifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 0.844 | 4 | 0.211 | 0 | 8.3 | 8.45 | Manual | Baseline | 12180.01 | 39465.2438 | 0.30863 | 2083.2678 | 8363.8018 | 0.24908 | 5.8466 | 4.7186 | 6005.2983 | 8376.1095 | 0.28743 | 0.98896 | 8.3203 | 8.471 | 0.24458 | 0.0203 | 8.2074 | 8.3115 | 8.3959 | 8.607 | 0.18851 | 0.2956 | 0.096357 | 0.074795 | 15.7921 | 211.8907 | 0.0068424 | 0.002545 | TRUE | 0.98221 | TRUE | [M-H]- | 65.7133 | 8.2314 | 8.3963 | 0.1649 | 8.2409 | 8.3963 | 0.15548 | 75.6176 | 0.91543 | 0.92797 | 0.42788 | 0 | 13 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 12180.01 | No Control Sample | 0 | off | off | 28466.2385 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9315 | 2.5908 | 1.2926 | 79.9601 | 34.7711 | 2.7802 | 0 | 0 | 0 | 0 | 0 | 1.6324 | FALSE | 0.211 | 0.30863
3849 | 3850 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFNS | 4.00505558645757 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.85 | 8.45 | Baseline | Baseline | 35033.4952 | 39465.2438 | 0.88771 | 7356.1524 | 8363.8018 | 0.87952 | 4.7625 | 4.7186 | 7363.4066 | 8376.1095 | 0.98926 | 0.98896 | 8.8701 | 8.471 | 0.22742 | 0.020126 | 8.743 | 8.3115 | 9.0733 | 8.607 | 0.33032 | 0.2956 | 0.073756 | 0.074795 | 119.3818 | 211.8907 | 0.00023059 | 0.002545 | FALSE | 1.0471 | TRUE | [M-H]- | 100.1264 | 8.7915 | 8.9513 | 0.15986 | 8.8044 | 8.943 | 0.13863 | -5.1353 | 1.0621 | 1.2229 | 15.9565 | 0 | 17 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 35033.4952 | No Control Sample | 0 | off | off | 35033.4952 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9276 | 1.1261 | 0.61816 | 79.9571 | 0.41537 | 0.033212 | 0 | 0 | 0 | 0 | 0 | 1.8624 | FALSE | 1 | 0.88771
3851 | 3852 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | PFDS | 4.06335857857317 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 9.19 | 8.45 | Baseline | Baseline | 36519.445 | 39465.2438 | 0.92536 | 7824.3745 | 8363.8018 | 0.9355 | 4.6674 | 4.7186 | 7831.431 | 8376.1095 | 0.98455 | 0.98896 | 9.1991 | 8.471 | 0.099558 | 0.0091494 | 9.0552 | 8.3115 | 9.3627 | 8.607 | 0.3075 | 0.2956 | 0.072678 | 0.074795 | 137.5239 | 211.8907 | 0.0020154 | 0.002545 | FALSE | 1.086 | TRUE | [M-H]- | 101.584 | 9.1246 | 9.2818 | 0.15723 | 9.135 | 9.2684 | 0.13337 | 51.2812 | 1.1886 | 1.3959 | 14.0032 | 0 | 17 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 36519.445 | No Control Sample | 0 | off | off | 36519.445 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9247 | 1.4375 | 0.86094 | 79.9568 | -1.155 | -0.092354 | 0 | 0 | 0 | 0 | 0 | 1.621 | FALSE | 1 | 0.92536
3855 | 3856 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FBSA | 3.39517598059977 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy | 0 | 0 | 4 | 4 | 1 | 0 | 4.6 | 8.86 | Manual | Baseline | 61467.7831 | 68661.3366 | 0.89523 | 5571.614 | 14098.5656 | 0.39519 | 11.0323 | 4.8701 | 5628.1786 | 14108.5744 | 0.96438 | 0.9945 | 4.9641 | 9.0137 | 7.915 | 0.36409 | 4.7286 | 8.854 | 5.2061 | 9.1581 | 0.4775 | 0.30408 | 0.16677 | 0.076087 | 210.3332 | 126.2639 | 0 | 0.0012597 | TRUE | 0.55072 | TRUE | [M-H]- | 84.8794 | 4.7748 | 5.1801 | 0.40529 | 4.7968 | 5.1406 | 0.34377 | 0 | 1.0945 | 1.1074 | 5.6531 | 0 | 86 | 31 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 61467.7831 | No Control Sample | 0 | off | off | 61467.7831 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9608 | 6.0819 | 1.8122 | 77.9651 | -3.0473 | -0.23758 | 0 | 0 | 0 | 0 | 0 | 0.98429 | FALSE | 1 | 0.89523
3857 | 3858 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FPeSA | 4.18930749513992 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7 | 8.86 | Baseline | Baseline | 117780.6192 | 68661.3366 | 1.7154 | 24707.745 | 14098.5656 | 1.7525 | 4.767 | 4.8701 | 24775.3777 | 14108.5744 | 0.98506 | 0.9945 | 7.1006 | 9.0137 | 1.4372 | 0.1006 | 6.943 | 8.854 | 7.2318 | 9.1581 | 0.28882 | 0.30408 | 0.074265 | 0.076087 | 312.8877 | 126.2639 | 0.0046949 | 0.0012597 | FALSE | 0.78775 | TRUE | [M-H]- | 104.7327 | 7.0246 | 7.1814 | 0.15678 | 7.0338 | 7.1702 | 0.13637 | 401.6416 | 1.0519 | 1.088 | 3.8813 | 0 | 31 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 117780.6192 | No Control Sample | 0 | off | off | 117780.6192 | 0 | off | 0 | off | 0 | off | 0 | off | 347.957 | 3.6349 | 1.2648 | 77.9649 | -5.6934 | -0.44389 | 0 | 0 | 0 | 0 | 0 | 1.5744 | FALSE | 1 | 1.7154
3859 | 3860 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FHxSA | 4.63007689139358 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 4 | 4 | 1 | 0 | 7.86 | 8.86 | Baseline | Baseline | 64958.9782 | 68661.3366 | 0.94608 | 14651.5113 | 14098.5656 | 1.0392 | 4.4336 | 4.8701 | 14653.2763 | 14108.5744 | 0.99701 | 0.9945 | 7.9876 | 9.0137 | 1.6235 | 0.1276 | 7.8481 | 8.854 | 8.1912 | 9.1581 | 0.34303 | 0.30408 | 0.068908 | 0.076087 | 97.2954 | 126.2639 | 0.00030028 | 0.0012597 | FALSE | 0.88616 | TRUE | [M-H]- | 115.7519 | 7.9191 | 8.0697 | 0.15062 | 7.9284 | 8.0593 | 0.13088 | 12.8256 | 1.1523 | 1.3351 | 7.3616 | 0 | 25 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 64958.9782 | No Control Sample | 0 | off | off | 64958.9782 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9538 | 3.1395 | 1.2494 | 77.9648 | -7.4485 | -0.58072 | 0 | 0 | 0 | 0 | 0 | 1.7691 | FALSE | 1 | 0.94608
3861 | 3862 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FHpSA | 4.09238719246259 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.58 | 8.86 | Baseline | Baseline | 118922.8384 | 68661.3366 | 1.732 | 26207.9825 | 14098.5656 | 1.8589 | 4.5377 | 4.8701 | 26267.82 | 14108.5744 | 0.98631 | 0.9945 | 8.5724 | 9.0137 | 0.088544 | 0.0075971 | 8.4348 | 8.854 | 8.8641 | 9.1581 | 0.42927 | 0.30408 | 0.070579 | 0.076087 | 135.1985 | 126.2639 | 0.0026081 | 0.0012597 | FALSE | 0.95104 | TRUE | [M-H]- | 102.3097 | 8.4992 | 8.6465 | 0.14733 | 8.5081 | 8.6376 | 0.12951 | -159.2301 | 1.1208 | 1.2781 | 3.7568 | 0 | 24 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 118922.8384 | No Control Sample | 0 | off | off | 118922.8384 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9507 | 2.8124 | 1.2598 | 77.9648 | -7.2873 | -0.56816 | 0 | 0 | 0 | 0 | 0 | 1.326 | FALSE | 1 | 1.732
3863 | 3864 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | FOSA | 4.30954018610787 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.88 | 8.86 | Baseline | Baseline | 67326.9064 | 68661.3366 | 0.98057 | 14772.6853 | 14098.5656 | 1.0478 | 4.5575 | 4.8701 | 14775.4427 | 14108.5744 | 0.99183 | 0.9945 | 9.0151 | 9.0137 | 1.5209 | 0.13506 | 8.8021 | 8.854 | 9.2682 | 9.1581 | 0.46607 | 0.30408 | 0.070549 | 0.076087 | 186.2136 | 126.2639 | 0.00039487 | 0.0012597 | FALSE | 1.0001 | TRUE | [M-H]- | 107.7385 | 8.9345 | 9.0904 | 0.15595 | 8.9472 | 9.0824 | 0.13524 | 12.516 | 0.96625 | 0.98986 | 7.7865 | 0 | 24 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 67326.9064 | No Control Sample | 0 | off | off | 67326.9064 | 0 | off | 0 | off | 0 | off | 0 | off | 497.947 | 1.6213 | 0.80731 | 77.965 | -4.8543 | -0.37847 | 0 | 0 | 0 | 0 | 0 | 1.6192 | FALSE | 1 | 0.98057
3882 | 3883 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 4:2 FTS | 3.95100411823245 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 5.94 | 5.92 | Baseline | Baseline | 36578.8817 | 18233.6526 | 2.0061 | 6172.657 | 3301.0824 | 1.8699 | 5.926 | 5.5235 | 6179.1651 | 3301.0825 | 0.99417 | 0.99876 | 5.9437 | 5.9428 | 0.061731 | 0.0036668 | 5.7947 | 5.7294 | 6.1287 | 6.1313 | 0.33392 | 0.40183 | 0.092006 | 0.083258 | 207.5063 | 191.7394 | 0.001893 | 9.1548e-09 | FALSE | 1.0002 | TRUE | [M-H]- | 98.7751 | 5.8419 | 6.0378 | 0.19597 | 5.8572 | 6.0248 | 0.16753 | -34.994 | 0.94718 | 0.90224 | 4.6877 | 0 | 39 | 13 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 36578.8817 | No Control Sample | 0 | off | off | 36578.8817 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9761 | 5.6755 | 1.8557 | 306.9687 | 3.4853 | 1.0699 | 0 | 0 | 0 | 0 | 0 | 5.3291 | FALSE | 1 | 2.0061
3884 | 3885 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 6:2 FTS | 3.76487850319385 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 7.92 | 7.92 | Baseline | Baseline | 36079.8777 | 15795.0708 | 2.2842 | 8780.4994 | 3986.4799 | 2.2026 | 4.1091 | 3.9622 | 8780.5007 | 3986.4799 | 0.99841 | 1 | 7.9214 | 7.9221 | 0.017301 | 0.0013702 | 7.7271 | 7.7182 | 8.1309 | 8.1528 | 0.40375 | 0.43453 | 0.06243 | 0.060735 | 97.9912 | 107.4966 | 3.0486e-07 | 0 | FALSE | 0.99991 | TRUE | [M-H]- | 94.122 | 7.8531 | 7.9933 | 0.14028 | 7.8617 | 7.9834 | 0.12177 | 0.00663 | 1.0872 | 1.1779 | 4.6807 | 0 | 30 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 36079.8777 | No Control Sample | 0 | off | off | 36079.8777 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9698 | 4.4914 | 1.9177 | 406.9635 | 4.9109 | 1.9985 | 0 | 0 | 0 | 0 | 0 | 5.5018 | FALSE | 1 | 2.2842
3886 | 3887 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 8:2 FTS | 4.19647648018425 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 4 | 4 | 1 | 0 | 8.88 | 8.89 | Baseline | Baseline | 35339.2157 | 17163.2099 | 2.059 | 6867.1216 | 3304.0166 | 2.0784 | 5.1461 | 5.1947 | 6869.0517 | 3304.683 | 0.99769 | 0.99522 | 8.8814 | 8.8816 | 0.015537 | 0.0013797 | 8.7249 | 8.7142 | 9.0373 | 9.0455 | 0.31237 | 0.33137 | 0.080505 | 0.081775 | 121.7039 | 150.0336 | 0.00050549 | 0.0004074 | FALSE | 0.99997 | TRUE | [M-H]- | 104.9119 | 8.7978 | 8.9693 | 0.17154 | 8.8088 | 8.956 | 0.14717 | 11.1129 | 0.94082 | 0.8373 | 3.7238 | 0 | 16 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 35339.2157 | No Control Sample | 0 | off | off | 35339.2157 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9635 | 3.7701 | 1.9867 | 506.9578 | 4.1202 | 2.0887 | 0 | 0 | 0 | 0 | 0 | 5.7212 | FALSE | 1 | 2.059
3888 | 3889 | PFAS CS6 4ng/ml | 28 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 29) | CS6 | 0 | 0 | Standard | 0 | 0 | 0 | 7 | 1 | 10:2 FTS | 3.16973303046909 | Quantifiers | 10/27/2023 02:35:53 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 4 | 4 | 1 | 0 | 9.5 | 8.89 | Baseline | Baseline | 27787.8048 | 17163.2099 | 1.619 | 6885.425 | 3304.0166 | 2.084 | 4.0357 | 5.1947 | 6885.425 | 3304.683 | 0.99675 | 0.99522 | 9.5018 | 8.8816 | 0.019258 | 0.0018295 | 9.3112 | 8.7142 | 9.671 | 9.0455 | 0.3598 | 0.33137 | 0.063132 | 0.081775 | 290.0071 | 150.0336 | 1.3267e-13 | 0.0004074 | FALSE | 1.0698 | TRUE | [M-H]- | 79.2433 | 9.4334 | 9.5679 | 0.13459 | 9.441 | 9.5606 | 0.11959 | 2.5388e-09 | 1.0618 | 1.1475 | 4.5033 | 0 | 24 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 27787.8048 | No Control Sample | 0 | off | off | 27787.8048 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9576 | 3.9226 | 2.4593 | 606.9514 | 5.7826 | 3.5097 | 0 | 0 | 0 | 0 | 0 | 1.697 | FALSE | 1 | 1.619
3949 | 3950 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFBA | 10.6688402547355 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 1.49 | 1.48 | Baseline | Baseline | 197703.6604 | 103308.479 | 1.9137 | 30413.0211 | 16033.4698 | 1.8968 | 6.5006 | 6.4433 | 30486.0959 | 16072.0389 | 0.98488 | 0.98508 | 1.4657 | 1.4653 | 1.6279 | 0.024255 | 1.3234 | 1.3145 | 1.6897 | 1.6709 | 0.36633 | 0.3564 | 0.10124 | 0.10284 | 312.9629 | 281.6962 | 0.0030779 | 0.006204 | FALSE | 1.0003 | TRUE | [M-H]- | 106.6884 | 1.3765 | 1.609 | 0.23248 | 1.3865 | 1.5852 | 0.19874 | 255.531 | 1.3769 | 1.6708 | 1.6428 | 0 | 73 | 22 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 197703.6604 | No Control Sample | 0 | off | off | 197703.6604 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9806 | 6.7194 | 1.4311 | 168.989 | -0.24721 | -0.041777 | 0 | 0 | 0 | 0 | 0 | 0.74008 | FALSE | 2.5 | 1.9137
3951 | 3952 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFPeA | 10.2704650910755 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 3.18 | 3.21 | Baseline | Baseline | 165332.7949 | 71981.4954 | 2.2969 | 18274.4595 | 7910.6346 | 2.3101 | 9.0472 | 9.0993 | 18361.2983 | 7987.9347 | 0.97849 | 0.96723 | 3.2011 | 3.1943 | 0.66403 | 0.021116 | 2.9676 | 3.0208 | 3.4333 | 3.4613 | 0.46563 | 0.44055 | 0.13413 | 0.13425 | 782.9519 | 693.0859 | 0.003593 | 0.0055009 | FALSE | 1.0021 | TRUE | [M-H]- | 102.7047 | 3.0587 | 3.3734 | 0.31477 | 3.0764 | 3.3468 | 0.27043 | 141.0138 | 1.0015 | 0.93906 | 2.1284 | 0 | 83 | 26 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 165332.7949 | No Control Sample | 0 | off | off | 165332.7949 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9775 | 5.4975 | 1.4457 | 218.9862 | 1.7083 | 0.3741 | 0 | 0 | 0 | 0 | 0 | 0.75753 | FALSE | 2.5 | 2.2969
3953 | 3954 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFHxA | 10.0303005278716 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 164.1925 | 6.13 | 6.12 | Baseline | Baseline | 334695.9634 | 143379.9715 | 2.3343 | 61242.0208 | 26686.3799 | 2.2949 | 5.4651 | 5.3728 | 61378.7483 | 26708.2891 | 0.98836 | 0.99455 | 6.1388 | 6.1342 | 0.1434 | 0.0087904 | 5.9466 | 5.9577 | 6.2932 | 6.3147 | 0.34667 | 0.35695 | 0.08508 | 0.081998 | 286.4633 | 249.322 | 0.0037718 | 0.0015902 | FALSE | 1.0007 | TRUE | [M-H]- | 100.303 | 6.0434 | 6.2288 | 0.18538 | 6.0549 | 6.2169 | 0.16198 | 666.3307 | 0.88422 | 0.73496 | 1.6014 | 0 | 38 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 334695.9634 | No Control Sample | 203843.5852 | off | off | 334695.9634 | 203843.5852 | off | 0.1 | off | 10 | off | 16.4193 | off | 312.9738 | 3.0914 | 0.96753 | 268.9825 | 0.39205 | 0.10545 | 0 | 0 | 0 | 0 | 0 | 14.8867 | FALSE | 2.5 | 2.3343
3955 | 3956 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFHpA | 12.9543976782364 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 209.4003 | 7.26 | 7.27 | Baseline | Baseline | 478121.1461 | 147500.0559 | 3.2415 | 108354.1532 | 34878.0808 | 3.1067 | 4.4126 | 4.229 | 108407.0859 | 34894.638 | 0.99596 | 0.99536 | 7.2785 | 7.2792 | 0.25478 | 0.018497 | 7.1337 | 7.1086 | 7.4486 | 7.4428 | 0.31483 | 0.33413 | 0.068657 | 0.064981 | 255.1761 | 238.5253 | 0.00091847 | 0.0009442 | FALSE | 0.9999 | TRUE | [M-H]- | 129.544 | 7.2101 | 7.3501 | 0.14007 | 7.2189 | 7.3407 | 0.12174 | 316.1022 | 0.97929 | 0.94327 | 1.4542 | 0 | 37 | 10 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 478121.1461 | No Control Sample | 228328.7703 | off | off | 478121.1461 | 228328.7703 | off | 0.1 | off | 10 | off | 20.94 | off | 362.9704 | 2.1921 | 0.79567 | 318.9792 | 0.39399 | 0.12567 | 0 | 0 | 0 | 0 | 0 | 1.2289 | FALSE | 2.5 | 3.2415
3957 | 3958 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFOA | 11.8085542503627 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 204.31 | 7.96 | 7.94 | Baseline | Baseline | 391203.4628 | 136602.0582 | 2.8638 | 92858.5067 | 30703.3575 | 3.0244 | 4.2129 | 4.4491 | 92935.834 | 30717.4084 | 0.9931 | 0.99607 | 7.9647 | 7.9642 | 0.058703 | 0.0046727 | 7.8345 | 7.8274 | 8.235 | 8.1149 | 0.40055 | 0.28753 | 0.064147 | 0.068605 | 148.392 | 132.6346 | 0.00039845 | 0.00087475 | FALSE | 1.0001 | TRUE | [M-H]- | 118.0855 | 7.9002 | 8.043 | 0.1428 | 7.9084 | 8.034 | 0.12557 | -92.3719 | 1.2066 | 1.465 | 1.2961 | 0 | 29 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 391203.4628 | No Control Sample | 191475.4056 | off | off | 391203.4628 | 191475.4056 | off | 0.1 | off | 10 | off | 20.431 | off | 412.9674 | 2.2489 | 0.92873 | 368.976 | -0.19774 | -0.072963 | 0 | 0 | 0 | 0 | 0 | 1.1179 | FALSE | 2.5 | 2.8638
3959 | 3960 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFNA | 11.0974914979435 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 199.1771 | 8.46 | 8.47 | Baseline | Baseline | 316977.2391 | 122953.2298 | 2.578 | 63627.9771 | 25409.704 | 2.5041 | 4.9817 | 4.8388 | 63638.0086 | 25424.6841 | 0.99823 | 0.99662 | 8.4781 | 8.4781 | 0.21417 | 0.018119 | 8.3336 | 8.2601 | 8.6809 | 8.6198 | 0.34732 | 0.35968 | 0.079098 | 0.075958 | 188.4726 | 156.6569 | 0.00018254 | 0.00094928 | FALSE | 1 | TRUE | [M-H]- | 110.9749 | 8.3952 | 8.5572 | 0.16191 | 8.4068 | 8.5454 | 0.13861 | 33.4413 | 1.0648 | 1.1505 | 1.1418 | 0 | 20 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 316977.2391 | No Control Sample | 159143.3885 | off | off | 316977.2391 | 159143.3885 | off | 0.1 | off | 10 | off | 19.9177 | off | 462.9647 | 3.2351 | 1.4977 | 418.9737 | 1.2974 | 0.54356 | 0 | 0 | 0 | 0 | 0 | 1.1142 | FALSE | 2.5 | 2.578
3961 | 3962 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFDA | 10.5064769262301 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.87 | 8.86 | Baseline | Baseline | 269347.7806 | 103270.5643 | 2.6082 | 51543.3876 | 21013.804 | 2.4528 | 5.2257 | 4.9144 | 51563.1487 | 21017.1447 | 0.99645 | 0.99856 | 8.8901 | 8.8914 | 0.22651 | 0.020091 | 8.7186 | 8.7273 | 9.0634 | 9.1295 | 0.3448 | 0.40213 | 0.081874 | 0.076445 | 146.9893 | 146.5024 | 0.00076698 | 0.00037936 | FALSE | 0.99986 | TRUE | [M-H]- | 105.0648 | 8.8059 | 8.9766 | 0.17068 | 8.8188 | 8.9622 | 0.14335 | 114.6536 | 1.1432 | 1.3195 | 1.0549 | 0 | 18 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 269347.7806 | No Control Sample | 0 | off | off | 269347.7806 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9618 | 3.3593 | 1.7232 | 468.9703 | 0.728 | 0.34141 | 0 | 0 | 0 | 0 | 0 | 11.1114 | FALSE | 2.5 | 2.6082
3963 | 3964 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFUdA | 9.84657588043196 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.21 | 9.2 | Baseline | Baseline | 198217.3959 | 90773.7601 | 2.1836 | 42201.0933 | 20530.5724 | 2.0555 | 4.697 | 4.4214 | 42222.2772 | 20546.1738 | 0.99551 | 0.99476 | 9.2246 | 9.2267 | 0.15893 | 0.014637 | 9.0832 | 9.0914 | 9.3699 | 9.3596 | 0.28675 | 0.2682 | 0.074239 | 0.068766 | 212.0026 | 148.8835 | 0.00095544 | 0.0013279 | FALSE | 0.99978 | TRUE | [M-H]- | 98.4658 | 9.15 | 9.3059 | 0.15588 | 9.1623 | 9.295 | 0.13278 | 140.6129 | 1.1215 | 1.3188 | 1.359 | 0 | 16 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 198217.3959 | No Control Sample | 0 | off | off | 198217.3959 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9587 | 3.2785 | 1.8456 | 518.9677 | 0.85176 | 0.44204 | 0 | 0 | 0 | 0 | 0 | 1.6115 | FALSE | 2.5 | 2.1836
3965 | 3966 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFDoA | 9.87589391838497 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.47 | 9.47 | Baseline | Baseline | 191200.317 | 74814.3228 | 2.5557 | 46378.6198 | 18591.4 | 2.4946 | 4.1226 | 4.0241 | 46385.6835 | 18592.3515 | 0.99769 | 0.99643 | 9.4932 | 9.4918 | 0.24493 | 0.023195 | 9.354 | 9.3605 | 9.687 | 9.7027 | 0.33303 | 0.34223 | 0.06438 | 0.062369 | 273.8112 | 298.1981 | 0.00035792 | 0.00011739 | FALSE | 1.0001 | TRUE | [M-H]- | 98.7589 | 9.4264 | 9.5621 | 0.13574 | 9.4364 | 9.5539 | 0.1175 | 49.8446 | 1.1267 | 1.3427 | 1.6714 | 0 | 23 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 191200.317 | No Control Sample | 0 | off | off | 191200.317 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9552 | 2.5416 | 1.5579 | 568.9642 | 1.3855 | 0.78828 | 0 | 0 | 0 | 0 | 0 | 34.1905 | FALSE | 2.5 | 2.5557
3967 | 3968 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFTrDA | 10.3342903019396 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.68 | 9.47 | Baseline | Baseline | 145840.4883 | 74814.3228 | 1.9494 | 40987.3098 | 18591.4 | 2.2046 | 3.5582 | 4.0241 | 41020.5939 | 18592.3515 | 0.99461 | 0.99643 | 9.7113 | 9.4918 | 0.32284 | 0.031251 | 9.5901 | 9.3605 | 9.816 | 9.7027 | 0.2259 | 0.34223 | 0.055785 | 0.062369 | 644.634 | 298.1981 | 0.0015006 | 0.00011739 | FALSE | 1.0231 | TRUE | [M-H]- | 103.3429 | 9.6505 | 9.771 | 0.12049 | 9.6576 | 9.7617 | 0.10414 | 272.261 | 0.99395 | 0.94461 | 1.047 | 0 | 20 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Fail | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 145840.4883 | No Control Sample | 0 | off | off | 145840.4883 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9523 | 2.823 | 1.8715 | 618.9613 | -0.70551 | -0.43668 | 0 | 0 | 0 | 0 | 0 | 0.73084 | FALSE | 2.5 | 1.9494
3969 | 3970 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFTeDA | 9.65367818850934 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.87 | 9.86 | Baseline | Baseline | 71948.8264 | 21021.1895 | 3.4227 | 20986.766 | 6169.2362 | 3.4018 | 3.4283 | 3.4074 | 21013.7693 | 6178.0784 | 0.98865 | 0.99016 | 9.8861 | 9.8859 | 0.16348 | 0.016136 | 9.7666 | 9.7687 | 9.9958 | 10.0036 | 0.22927 | 0.2349 | 0.053891 | 0.052697 | 921.2504 | 864.0916 | 0.0022487 | 0.0027613 | FALSE | 1 | TRUE | [M-H]- | 96.5368 | 9.8364 | 9.9466 | 0.11022 | 9.841 | 9.9385 | 0.097477 | 205.8423 | 1.084 | 1.1113 | 1.5975 | 0 | 25 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 71948.8264 | No Control Sample | 0 | off | off | 71948.8264 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9495 | 3.0843 | 2.1989 | 668.9601 | 3.2465 | 2.1718 | 0 | 0 | 0 | 0 | 0 | 38.7857 | FALSE | 2.5 | 3.4227
3976 | 3977 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFBS | 9.74009281375534 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 3.5 | 3.5 | Baseline | Manual | 105524.7738 | 40931.4747 | 2.5781 | 10787.8849 | 3585.0649 | 3.0091 | 9.7818 | 11.4172 | 11011.9552 | 3593.3359 | 0.93502 | 0.99061 | 3.8387 | 3.8379 | 9.6773 | 0.33871 | 3.5851 | 3.5928 | 4.0513 | 4.1379 | 0.46622 | 0.54507 | 0.14991 | 0.17152 | 444.1221 | 876.9932 | 0.032899 | 0.00096179 | FALSE | 1.0002 | TRUE | [M-H]- | 97.4009 | 3.68 | 4.0278 | 0.34779 | 3.6953 | 4.0003 | 0.30503 | 761.2456 | 1.1255 | 1.1908 | 13.1134 | 0 | 72 | 25 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 105524.7738 | No Control Sample | 0 | off | off | 105524.7738 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9437 | 2.5318 | 0.75685 | 79.9567 | -5.8343 | -0.46649 | 0 | 0 | 0 | 0 | 0 | 1.0316 | FALSE | 2.5 | 2.5781
3978 | 3979 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFPeS | 11.7185844262635 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 6.38 | 3.5 | Baseline | Manual | 127834.0901 | 40931.4747 | 3.1231 | 25684.312 | 3585.0649 | 7.1643 | 4.9771 | 11.4172 | 25707.8189 | 3593.3359 | 0.99148 | 0.99061 | 6.3881 | 3.8379 | 0.12734 | 0.0081244 | 6.2254 | 3.5928 | 6.5648 | 4.1379 | 0.33938 | 0.54507 | 0.07769 | 0.17152 | 332.7122 | 876.9932 | 0.001815 | 0.00096179 | FALSE | 1.6645 | TRUE | [M-H]- | 117.1858 | 6.3071 | 6.4688 | 0.16167 | 6.3179 | 6.4593 | 0.14135 | 137.3576 | 1.0069 | 1.0336 | 7.9 | 0 | 37 | 11 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 127834.0901 | No Control Sample | 0 | off | off | 127834.0901 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9406 | 2.2004 | 0.76781 | 79.9567 | -6.0555 | -0.48418 | 0 | 0 | 0 | 0 | 0 | 1.649 | FALSE | 2.5 | 3.1231
3980 | 3981 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | L-PFHxS | 8.72679433842753 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 8.11 | 4 | 2.0275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 91802.2026 | 56316.3847 | 1.6301 | 23098.2992 | 13926.5762 | 1.6586 | 3.9744 | 4.0438 | 23395.3437 | 13939.6022 | 0.76074 | 0.99031 | 7.3541 | 7.3539 | 0.19232 | 0.014117 | 7.2664 | 7.2032 | 7.4932 | 7.5086 | 0.22688 | 0.30538 | 0.062233 | 0.062476 | 123.7204 | 181.5782 | 0.016175 | 0.0015852 | FALSE | 1 | TRUE | [M-H]- | 107.6054 | 7.291 | 7.4209 | 0.12994 | 7.2983 | 7.4127 | 0.11446 | -1646.7173 | 1.0902 | 1.1899 | 13.2881 | 0 | 28 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 91802.2026 | No Control Sample | 0 | off | off | 91802.2026 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9374 | 1.9975 | 0.79686 | 79.9568 | -5.6091 | -0.44849 | 0 | 0 | 0 | 0 | 0 | 1.2961 | FALSE | 2.0275 | 1.6301
3982 | 3983 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | Br-PFHxS | 1.66780632479201 | Qualifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 1.89 | 4 | 0.4725 | 0 | 7.16 | 7.35 | Manual | Baseline | 23836.7903 | 56316.3847 | 0.42327 | 4248.4538 | 13926.5762 | 0.30506 | 5.6107 | 4.0438 | 23395.3437 | 13939.6022 | 0.19763 | 0.99031 | 7.1671 | 7.3539 | 0.099146 | 0.0070989 | 7.0577 | 7.2032 | 7.2666 | 7.5086 | 0.20892 | 0.30538 | 0.08872 | 0.062476 | 22.7558 | 181.5782 | 0.024196 | 0.0015852 | TRUE | 0.97459 | TRUE | [M-H]- | 88.2437 | 7.0851 | 7.2748 | 0.18979 | 7.0937 | 7.2558 | 0.16208 | 492.0301 | 1.0793 | 1.0447 | 0.25965 | 0 | 25 | 12 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 23836.7903 | No Control Sample | 0 | off | off | 91802.2026 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9374 | 1.9975 | 0.79686 | 79.9498 | -92.5403 | -7.3993 | 0 | 0 | 0 | 0 | 0 | 1.2961 | FALSE | 0.4725 | 0.42327
3984 | 3985 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFHpS | 12.0615562272874 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7.98 | 7.35 | Baseline | Baseline | 103379.7069 | 56316.3847 | 1.8357 | 24806.4686 | 13926.5762 | 1.7812 | 4.1674 | 4.0438 | 24826.3601 | 13939.6022 | 0.99336 | 0.99031 | 7.9881 | 7.3539 | 0.10157 | 0.0081049 | 7.8371 | 7.2032 | 8.1588 | 7.5086 | 0.32167 | 0.30538 | 0.064842 | 0.062476 | 138.4187 | 181.5782 | 0.00058798 | 0.0015852 | FALSE | 1.0862 | TRUE | [M-H]- | 120.6156 | 7.9237 | 8.0645 | 0.14078 | 7.9325 | 8.0561 | 0.12363 | 45.3442 | 1.151 | 1.358 | 10.1089 | 0 | 24 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 103379.7069 | No Control Sample | 0 | off | off | 103379.7069 | 0 | off | 0 | off | 0 | off | 0 | off | 448.934 | 1.3129 | 0.58942 | 79.9566 | -7.0202 | -0.56132 | 0 | 0 | 0 | 0 | 0 | 1.8708 | FALSE | 2.5 | 1.8357
3986 | 3987 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | L-PFOS | 7.68717217679825 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 7.88 | 4 | 1.97 | 0 | 8.46 | 8.45 | Manual | Baseline | 72006.6997 | 36037.1943 | 1.9981 | 15243.49 | 7285.4641 | 2.0923 | 4.7238 | 4.9465 | 15445.035 | 7294.5162 | 0.67039 | 0.98679 | 8.4751 | 8.4738 | 0.17865 | 0.015113 | 8.3938 | 8.327 | 8.5687 | 8.6068 | 0.1749 | 0.27977 | 0.074424 | 0.078291 | 73.7615 | 167.6877 | 0.015275 | 0.0025039 | TRUE | 1.0002 | TRUE | [M-H]- | 97.5529 | 8.3932 | 8.5495 | 0.15625 | 8.3932 | 8.5419 | 0.14869 | -1331.2444 | 0.9481 | 0.8045 | 14.29 | 0 | 11 | 5 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 72006.6997 | No Control Sample | 0 | off | off | 72006.6997 | 0 | off | 0 | off | 0 | off | 0 | off | 498.931 | 1.5212 | 0.75899 | 79.9567 | -6.8989 | -0.55162 | 0 | 0 | 0 | 0 | 0 | 1.4019 | FALSE | 1.97 | 1.9981
3988 | 3989 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | Br-PFOS | 1.48524407480862 | Qualifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 2.11 | 4 | 0.5275 | 0 | 8.3 | 8.45 | Manual | Baseline | 29784.2153 | 36037.1943 | 0.82649 | 4959.6735 | 7285.4641 | 0.68076 | 6.0053 | 4.9465 | 15445.0269 | 7294.5162 | 0.27732 | 0.98679 | 8.3262 | 8.4738 | 0.31591 | 0.02622 | 8.2189 | 8.327 | 8.393 | 8.6068 | 0.17406 | 0.27977 | 0.10395 | 0.078291 | 23.9993 | 167.6877 | 0.013526 | 0.0025039 | TRUE | 0.98259 | TRUE | [M-H]- | 70.3907 | 8.2371 | 8.3932 | 0.15615 | 8.2447 | 8.3932 | 0.1485 | 385.3939 | 0.78063 | 0.60773 | 0.41363 | 0 | 12 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 29784.2153 | No Control Sample | 0 | off | off | 72006.6997 | 0 | off | 0 | off | 0 | off | 0 | off | 498.931 | 1.5212 | 0.75899 | 79.9638 | 80.8043 | 6.4609 | 0 | 0 | 0 | 0 | 0 | 1.4019 | FALSE | 0.5275 | 0.82649
3990 | 3991 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFNS | 10.9671765549398 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.85 | 8.45 | Baseline | Baseline | 87600.3691 | 36037.1943 | 2.4308 | 17740.2693 | 7285.4641 | 2.435 | 4.9379 | 4.9465 | 17753.4035 | 7294.5162 | 0.99253 | 0.98679 | 8.8743 | 8.4738 | 0.27475 | 0.024315 | 8.723 | 8.327 | 9.049 | 8.6068 | 0.32608 | 0.27977 | 0.077322 | 0.078291 | 115.1888 | 167.6877 | 0.0016679 | 0.0025039 | FALSE | 1.0473 | TRUE | [M-H]- | 109.6718 | 8.7974 | 8.9594 | 0.16206 | 8.806 | 8.9453 | 0.1393 | 90.7396 | 1.2065 | 1.3808 | 12.2137 | 0 | 17 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 87600.3691 | No Control Sample | 0 | off | off | 87600.3691 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9274 | 0.64258 | 0.35273 | 79.9565 | -7.2257 | -0.57774 | 0 | 0 | 0 | 0 | 0 | 1.3028 | FALSE | 2.5 | 2.4308
3992 | 3993 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFDS | 10.6376743983953 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.19 | 8.45 | Baseline | Baseline | 87301.538 | 36037.1943 | 2.4225 | 19804.0877 | 7285.4641 | 2.7183 | 4.4083 | 4.9465 | 19812.6077 | 7294.5162 | 0.9923 | 0.98679 | 9.2039 | 8.4738 | 0.1517 | 0.013942 | 9.0686 | 8.327 | 9.3743 | 8.6068 | 0.30563 | 0.27977 | 0.06847 | 0.078291 | 234.3812 | 167.6877 | 0.00089736 | 0.0025039 | FALSE | 1.0862 | TRUE | [M-H]- | 106.3767 | 9.1313 | 9.282 | 0.15065 | 9.1417 | 9.273 | 0.13135 | 58.1459 | 1.0328 | 1.098 | 12.2484 | 0 | 17 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 87301.538 | No Control Sample | 0 | off | off | 87301.538 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9241 | 0.47393 | 0.28385 | 79.9566 | -3.2086 | -0.25655 | 0 | 0 | 0 | 0 | 0 | 1.7171 | FALSE | 2.5 | 2.4225
3996 | 3997 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FBSA | 8.9588043326914 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 4.6 | 8.86 | Manual | Baseline | 151336.1305 | 64064.8136 | 2.3622 | 14148.7656 | 14062.9996 | 1.0061 | 10.6961 | 4.5556 | 14257.4597 | 14067.4348 | 0.96994 | 0.99452 | 4.96 | 9.0176 | 7.8269 | 0.36004 | 4.7298 | 8.831 | 5.1977 | 9.1903 | 0.46786 | 0.35932 | 0.16745 | 0.070442 | 354.1345 | 152.7234 | 0.0050637 | 0.0005803 | TRUE | 0.55004 | TRUE | [M-H]- | 89.588 | 4.7662 | 5.149 | 0.38283 | 4.7976 | 5.1205 | 0.3229 | -153.1331 | 1.0101 | 1.0416 | 5.9276 | 0 | 84 | 32 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 151336.1305 | No Control Sample | 0 | off | off | 151336.1305 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9602 | 4.0064 | 1.1937 | 77.9649 | -5.1549 | -0.40191 | 0 | 0 | 0 | 0 | 0 | 1.1937 | FALSE | 2.5 | 2.3622
3998 | 3999 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FPeSA | 11.8289838133561 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7 | 8.86 | Baseline | Baseline | 310303.2651 | 64064.8136 | 4.8436 | 65054.255 | 14062.9996 | 4.6259 | 4.7699 | 4.5556 | 65104.668 | 14067.4348 | 0.99314 | 0.99452 | 7.1052 | 9.0176 | 1.5024 | 0.10517 | 6.9216 | 8.831 | 7.3163 | 9.1903 | 0.39472 | 0.35932 | 0.073911 | 0.070442 | 234.8596 | 152.7234 | 0.0016156 | 0.0005803 | FALSE | 0.78792 | TRUE | [M-H]- | 118.2898 | 7.0255 | 7.1854 | 0.15995 | 7.0361 | 7.1737 | 0.13758 | 266.2676 | 0.98194 | 0.94437 | 3.1195 | 0 | 43 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 310303.2651 | No Control Sample | 0 | off | off | 310303.2651 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9566 | 2.2524 | 0.78373 | 77.9646 | -8.7591 | -0.68291 | 0 | 0 | 0 | 0 | 0 | 1.1204 | FALSE | 2.5 | 4.8436
4000 | 4001 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FHxSA | 12.1494361104293 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7.86 | 8.86 | Baseline | Baseline | 159042.9478 | 64064.8136 | 2.4825 | 35543.01 | 14062.9996 | 2.5274 | 4.4747 | 4.5556 | 35550.2374 | 14067.4348 | 0.99749 | 0.99452 | 7.992 | 9.0176 | 1.6791 | 0.13198 | 7.8354 | 8.831 | 8.2527 | 9.1903 | 0.41737 | 0.35932 | 0.069218 | 0.070442 | 150.5872 | 152.7234 | 8.8722e-05 | 0.0005803 | FALSE | 0.88627 | TRUE | [M-H]- | 121.4944 | 7.9242 | 8.0729 | 0.14872 | 7.933 | 8.0609 | 0.12796 | 7.5556 | 1.0402 | 1.0404 | 4.9555 | 0 | 30 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 159042.9478 | No Control Sample | 0 | off | off | 159042.9478 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9536 | 2.5895 | 1.0305 | 77.9647 | -8.7059 | -0.67876 | 0 | 0 | 0 | 0 | 0 | 1.1973 | FALSE | 2.5 | 2.4825
4002 | 4003 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FHpSA | 11.3333944041156 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.58 | 8.86 | Baseline | Baseline | 307295.2629 | 64064.8136 | 4.7966 | 67578.5872 | 14062.9996 | 4.8054 | 4.5472 | 4.5556 | 67590.0051 | 14067.4348 | 0.99689 | 0.99452 | 8.575 | 9.0176 | 0.058855 | 0.0050498 | 8.3176 | 8.831 | 8.8201 | 9.1903 | 0.50258 | 0.35932 | 0.070454 | 0.070442 | 167.2314 | 152.7234 | 0.00031873 | 0.0005803 | FALSE | 0.95091 | TRUE | [M-H]- | 113.3339 | 8.5006 | 8.651 | 0.15039 | 8.5125 | 8.6411 | 0.1286 | 42.8575 | 0.92602 | 0.85575 | 2.7921 | 0 | 28 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 307295.2629 | No Control Sample | 0 | off | off | 307295.2629 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9503 | 2.0055 | 0.89838 | 77.9646 | -10.5606 | -0.82336 | 0 | 0 | 0 | 0 | 0 | 1.4255 | FALSE | 2.5 | 4.7966
4004 | 4005 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FOSA | 11.9361301089256 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.88 | 8.86 | Baseline | Baseline | 173991.7206 | 64064.8136 | 2.7159 | 35935.9602 | 14062.9996 | 2.5554 | 4.8417 | 4.5556 | 35942.5445 | 14067.4348 | 0.99308 | 0.99452 | 9.0195 | 9.0176 | 1.5706 | 0.13947 | 8.8628 | 8.831 | 9.3212 | 9.1903 | 0.4584 | 0.35932 | 0.075491 | 0.070442 | 125.5977 | 152.7234 | 2.4378e-05 | 0.0005803 | FALSE | 1.0002 | TRUE | [M-H]- | 119.3613 | 8.9394 | 9.0979 | 0.15851 | 8.9488 | 9.0852 | 0.13646 | 1.9111 | 1.1205 | 1.2243 | 5.8869 | 0 | 24 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 173991.7206 | No Control Sample | 0 | off | off | 173991.7206 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9467 | 0.9171 | 0.45666 | 77.9649 | -6.7904 | -0.52942 | 0 | 0 | 0 | 0 | 0 | 0.58193 | FALSE | 2.5 | 2.7159
4023 | 4024 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 4:2 FTS | 10.8221523126095 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 5.94 | 5.92 | Baseline | Baseline | 92533.8791 | 16839.8362 | 5.4949 | 15910.9892 | 2747.5432 | 5.791 | 5.8157 | 6.1291 | 15929.5831 | 2747.5432 | 0.99302 | 0.99935 | 5.9445 | 5.9435 | 0.075378 | 0.0044775 | 5.757 | 5.7173 | 6.132 | 6.1629 | 0.37493 | 0.44558 | 0.088515 | 0.095775 | 192.0373 | 181.7502 | 0.0022678 | 0 | FALSE | 1.0002 | TRUE | [M-H]- | 108.2215 | 5.8389 | 6.0429 | 0.20399 | 5.8566 | 6.029 | 0.17243 | 96.2369 | 0.93197 | 0.87928 | 4.786 | 0 | 45 | 12 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 92533.8791 | No Control Sample | 0 | off | off | 92533.8791 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9756 | 4.0444 | 1.3224 | 306.9686 | 3.2254 | 0.9901 | 0 | 0 | 0 | 0 | 0 | 5.3291 | FALSE | 2.5 | 5.4949
4025 | 4026 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 6:2 FTS | 10.8691616805106 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 7.92 | 7.92 | Baseline | Baseline | 85658.4302 | 12989.1744 | 6.5946 | 21059.0207 | 3221.6018 | 6.5368 | 4.0675 | 4.0319 | 21065.4908 | 3224.5738 | 0.9966 | 0.99349 | 7.9263 | 7.9276 | 0.079992 | 0.0063354 | 7.7863 | 7.8046 | 8.1128 | 8.0715 | 0.32652 | 0.26683 | 0.062546 | 0.05961 | 122.1653 | 108.3564 | 0.00067289 | 0.0018649 | FALSE | 0.99983 | TRUE | [M-H]- | 108.6916 | 7.858 | 7.9972 | 0.13914 | 7.8679 | 7.9855 | 0.11755 | 43.3988 | 0.9809 | 0.92561 | 3.7043 | 0 | 25 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 85658.4302 | No Control Sample | 0 | off | off | 85658.4302 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9696 | 3.9036 | 1.6667 | 406.9632 | 4.2758 | 1.7401 | 0 | 0 | 0 | 0 | 0 | 5.5018 | FALSE | 2.5 | 6.5946
4027 | 4028 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 8:2 FTS | 12.0823680120544 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.88 | 8.89 | Baseline | Baseline | 77128.7491 | 13010.3994 | 5.9282 | 15082.2229 | 2430.1441 | 6.2063 | 5.1139 | 5.3538 | 15090.5732 | 2432.0535 | 0.99474 | 0.99383 | 8.8851 | 8.8854 | 0.057417 | 0.0050986 | 8.7239 | 8.7321 | 9.0322 | 9.0218 | 0.30832 | 0.28975 | 0.080657 | 0.083833 | 139.7013 | 120.2137 | 0.0010356 | 0.001577 | FALSE | 0.99996 | TRUE | [M-H]- | 120.8237 | 8.8043 | 8.9693 | 0.16503 | 8.8114 | 8.9608 | 0.14939 | 50.6578 | 1.0015 | 0.98408 | 3.0115 | 0 | 16 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 77128.7491 | No Control Sample | 0 | off | off | 77128.7491 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9633 | 3.4566 | 1.8215 | 506.9569 | 2.2736 | 1.1526 | 0 | 0 | 0 | 0 | 0 | 5.7212 | FALSE | 2.5 | 5.9282
4029 | 4030 | PFAS CS7 10ng/ml | 29 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 30) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 10:2 FTS | 10.651663138859 | Quantifiers | 10/27/2023 02:52:28 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.5 | 8.89 | Baseline | Baseline | 70784.9763 | 13010.3994 | 5.4406 | 16888.8801 | 2430.1441 | 6.9497 | 4.1912 | 5.3538 | 16888.8803 | 2432.0535 | 0.99834 | 0.99383 | 9.5091 | 8.8854 | 0.095935 | 0.0091138 | 9.3242 | 8.7321 | 9.6772 | 9.0218 | 0.35305 | 0.28975 | 0.065598 | 0.083833 | 369.6664 | 120.2137 | 1.3207e-08 | 0.001577 | FALSE | 1.0702 | TRUE | [M-H]- | 106.5166 | 9.4405 | 9.579 | 0.13852 | 9.4471 | 9.5712 | 0.12417 | 0.00063179 | 1.0591 | 1.1106 | 4.6249 | 0 | 24 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 70784.9763 | No Control Sample | 0 | off | off | 70784.9763 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9571 | 3.1013 | 1.9444 | 606.9517 | 6.262 | 3.8007 | 0 | 0 | 0 | 0 | 0 | 1.5397 | FALSE | 2.5 | 5.4406
4090 | 4091 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFBA | 10.9646615594601 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 1.49 | 1.48 | Baseline | Baseline | 201977.5837 | 102694.317 | 1.9668 | 28375.742 | 15292.2007 | 1.8556 | 7.118 | 6.7155 | 28495.1409 | 15324.4307 | 0.97779 | 0.98754 | 1.4548 | 1.4536 | 2.3605 | 0.035171 | 1.3063 | 1.3075 | 1.6621 | 1.6683 | 0.3558 | 0.36082 | 0.11034 | 0.10018 | 280.8163 | 244.502 | 0.0093423 | 0.0053459 | FALSE | 1.0008 | TRUE | [M-H]- | 109.6466 | 1.3663 | 1.6042 | 0.23793 | 1.3742 | 1.5837 | 0.20949 | 745.0664 | 1.4483 | 1.8227 | 1.6154 | 0 | 71 | 24 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 201977.5837 | No Control Sample | 0 | off | off | 201977.5837 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9806 | 6.6151 | 1.4089 | 168.9891 | 0.36259 | 0.061274 | 0 | 0 | 0 | 0 | 0 | 0.37118 | FALSE | 2.5 | 1.9668
4092 | 4093 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFPeA | 10.8451798371674 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 3.18 | 3.21 | Baseline | Baseline | 172994.027 | 71325.7454 | 2.4254 | 19134.854 | 8090.0537 | 2.3652 | 9.0408 | 8.8165 | 19246.4977 | 8208.504 | 0.97476 | 0.94942 | 3.1838 | 3.1806 | 0.118 | 0.0037523 | 2.974 | 2.9843 | 3.4345 | 3.3903 | 0.46047 | 0.406 | 0.13778 | 0.13895 | 683.2833 | 638.8832 | 0.0024879 | 0.015247 | FALSE | 1.001 | TRUE | [M-H]- | 108.4518 | 3.0381 | 3.3634 | 0.32525 | 3.058 | 3.3266 | 0.2686 | 103.3841 | 1.0779 | 1.0514 | 2.01 | 0 | 82 | 26 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 172994.027 | No Control Sample | 0 | off | off | 172994.027 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9775 | 5.5605 | 1.4623 | 218.9865 | 3.3192 | 0.72686 | 0 | 0 | 0 | 0 | 0 | 0.083922 | FALSE | 2.5 | 2.4254
4094 | 4095 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFHxA | 11.0553968184708 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 179.0728 | 6.13 | 6.12 | Baseline | Baseline | 365028.4472 | 141874.4974 | 2.5729 | 65728.059 | 25541.1292 | 2.5734 | 5.5536 | 5.5547 | 65797.1649 | 25560.2031 | 0.99225 | 0.99417 | 6.119 | 6.1172 | 0.17985 | 0.011025 | 5.9254 | 5.928 | 6.2976 | 6.3096 | 0.3722 | 0.38158 | 0.086363 | 0.087219 | 294.1922 | 300.3243 | 0.0018409 | 0.0014858 | FALSE | 1.0003 | TRUE | [M-H]- | 110.554 | 6.0277 | 6.2129 | 0.18515 | 6.0399 | 6.2 | 0.1601 | 325.0953 | 0.96307 | 0.90825 | 1.5123 | 0 | 41 | 12 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 365028.4472 | No Control Sample | 203843.5852 | off | off | 365028.4472 | 203843.5852 | off | 0.1 | off | 10 | off | 17.9073 | off | 312.9739 | 3.4844 | 1.0905 | 268.9828 | 1.4475 | 0.38935 | 0 | 0 | 0 | 0 | 0 | 13.5547 | FALSE | 2.5 | 2.5729
4096 | 4097 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFHpA | 11.5587917201521 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 204.7366 | 7.26 | 7.27 | Baseline | Baseline | 467472.5814 | 161627.4657 | 2.8923 | 109159.1489 | 36915.321 | 2.957 | 4.2825 | 4.3783 | 109169.6435 | 36933.523 | 0.99833 | 0.99586 | 7.2744 | 7.2726 | 0.19823 | 0.014392 | 7.1255 | 7.1096 | 7.5443 | 7.4266 | 0.41882 | 0.317 | 0.066854 | 0.068309 | 251.3288 | 234.0614 | 0.000178 | 0.00098157 | FALSE | 1.0002 | TRUE | [M-H]- | 115.5879 | 7.2022 | 7.3434 | 0.14123 | 7.2118 | 7.3338 | 0.12203 | 46.3945 | 0.98595 | 0.96763 | 1.401 | 0 | 47 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 467472.5814 | No Control Sample | 228328.7703 | off | off | 467472.5814 | 228328.7703 | off | 0.1 | off | 10 | off | 20.4737 | off | 362.9707 | 2.8646 | 1.0397 | 318.9793 | 0.74754 | 0.23845 | 0 | 0 | 0 | 0 | 0 | 1.4342 | FALSE | 2.5 | 2.8923
4098 | 4099 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFOA | 11.6723401639454 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 213.1959 | 7.96 | 7.94 | Baseline | Baseline | 408217.7733 | 144206.6392 | 2.8308 | 99762.8408 | 33764.2105 | 2.9547 | 4.0919 | 4.271 | 99829.6483 | 33779.7946 | 0.99363 | 0.99481 | 7.9626 | 7.9632 | 0.032919 | 0.0026204 | 7.7965 | 7.8275 | 8.134 | 8.1413 | 0.33752 | 0.31388 | 0.062581 | 0.065418 | 119.3386 | 133.0665 | 0.00049554 | 0.0010141 | FALSE | 0.99993 | TRUE | [M-H]- | 116.7234 | 7.8978 | 8.0406 | 0.14282 | 7.9066 | 8.0302 | 0.12356 | 146.4711 | 1.1607 | 1.3449 | 1.2561 | 0 | 26 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 408217.7733 | No Control Sample | 191475.4056 | off | off | 408217.7733 | 191475.4056 | off | 0.1 | off | 10 | off | 21.3196 | off | 412.9677 | 3.0333 | 1.2526 | 368.9764 | 0.70662 | 0.26072 | 0 | 0 | 0 | 0 | 0 | 0.66178 | FALSE | 2.5 | 2.8308
4100 | 4101 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFNA | 11.0434217072199 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 196.1776 | 8.46 | 8.47 | Baseline | Baseline | 312203.7412 | 121694.5504 | 2.5655 | 62293.3897 | 25339.2252 | 2.4584 | 5.0118 | 4.8026 | 62321.2106 | 25345.4206 | 0.99635 | 0.99739 | 8.4773 | 8.4762 | 0.2043 | 0.017284 | 8.3278 | 8.2863 | 8.7289 | 8.6631 | 0.4011 | 0.37675 | 0.078715 | 0.074871 | 102.9718 | 138.9598 | 0.00037222 | 0.00047871 | FALSE | 1.0001 | TRUE | [M-H]- | 110.4342 | 8.3962 | 8.5557 | 0.15948 | 8.4056 | 8.5465 | 0.14088 | 57.8077 | 0.88604 | 0.74704 | 1.2972 | 0 | 23 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 312203.7412 | No Control Sample | 159143.3885 | off | off | 312203.7412 | 159143.3885 | off | 0.1 | off | 10 | off | 19.6178 | off | 462.9652 | 4.1915 | 1.9405 | 418.9736 | 1.0573 | 0.44299 | 0 | 0 | 0 | 0 | 0 | 0.13254 | FALSE | 2.5 | 2.5655
4102 | 4103 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFDA | 9.68716187178898 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.87 | 8.86 | Baseline | Baseline | 260639.831 | 108383.8286 | 2.4048 | 52222.4562 | 21782.2681 | 2.3975 | 4.991 | 4.9758 | 52260.0241 | 21785.0568 | 0.99472 | 0.99821 | 8.8913 | 8.8884 | 0.24009 | 0.021296 | 8.7298 | 8.683 | 9.0551 | 9.1196 | 0.32533 | 0.43665 | 0.078046 | 0.078015 | 134.0671 | 163.1277 | 0.0013496 | 0.00027187 | FALSE | 1.0003 | TRUE | [M-H]- | 96.8716 | 8.8112 | 8.9755 | 0.16438 | 8.8185 | 8.9656 | 0.14713 | 216.6387 | 1.005 | 0.97642 | 1.0351 | 0 | 17 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 260639.831 | No Control Sample | 0 | off | off | 260639.831 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9623 | 4.4732 | 2.2946 | 468.9707 | 1.4828 | 0.69539 | 0 | 0 | 0 | 0 | 0 | 8.5418 | FALSE | 2.5 | 2.4048
4104 | 4105 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFUdA | 11.1702976133072 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.21 | 9.2 | Baseline | Baseline | 228689.1797 | 92317.6336 | 2.4772 | 49669.9306 | 20281.072 | 2.4491 | 4.6042 | 4.5519 | 49715.5124 | 20294.2946 | 0.99461 | 0.99514 | 9.227 | 9.2224 | 0.1851 | 0.017047 | 9.0747 | 9.0829 | 9.361 | 9.3844 | 0.28628 | 0.30147 | 0.072107 | 0.071636 | 271.012 | 227.139 | 0.0016369 | 0.0014092 | FALSE | 1.0005 | TRUE | [M-H]- | 111.703 | 9.1523 | 9.3047 | 0.15239 | 9.1591 | 9.2947 | 0.13564 | 284.003 | 0.98926 | 0.92912 | 0.85934 | 0 | 16 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 228689.1797 | No Control Sample | 0 | off | off | 228689.1797 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9596 | 4.8182 | 2.7124 | 518.9677 | 0.71023 | 0.36859 | 0 | 0 | 0 | 0 | 0 | 0.7929 | FALSE | 2.5 | 2.4772
4106 | 4107 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFDoA | 10.7345771368448 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.47 | 9.47 | Baseline | Baseline | 202960.0363 | 73063.109 | 2.7779 | 48307.7046 | 17541.4563 | 2.7539 | 4.2014 | 4.1652 | 48343.6297 | 17545.1991 | 0.99507 | 0.99618 | 9.4891 | 9.4919 | 0.20201 | 0.019131 | 9.3619 | 9.3342 | 9.6108 | 9.6593 | 0.24897 | 0.32515 | 0.066487 | 0.065147 | 389.0427 | 271.0588 | 0.0013866 | 0.00040947 | FALSE | 0.99971 | TRUE | [M-H]- | 107.3458 | 9.4251 | 9.5604 | 0.13532 | 9.4321 | 9.5511 | 0.11901 | 269.0441 | 0.98179 | 0.92274 | 1.6817 | 0 | 17 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 202960.0363 | No Control Sample | 0 | off | off | 202960.0363 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9553 | 2.7469 | 1.6837 | 568.9652 | 3.1069 | 1.7677 | 0 | 0 | 0 | 0 | 0 | 35.6888 | FALSE | 2.5 | 2.7779
4108 | 4109 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFTrDA | 10.6442737950081 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.68 | 9.47 | Baseline | Baseline | 146698.9106 | 73063.109 | 2.0078 | 39117.2518 | 17541.4563 | 2.23 | 3.7502 | 4.1652 | 39126.7928 | 17545.1991 | 0.99669 | 0.99618 | 9.7071 | 9.4919 | 0.28043 | 0.027145 | 9.5682 | 9.3342 | 9.843 | 9.6593 | 0.27483 | 0.32515 | 0.059609 | 0.065147 | 670.3622 | 271.0588 | 0.00048061 | 0.00040947 | FALSE | 1.0227 | TRUE | [M-H]- | 106.4427 | 9.6453 | 9.7681 | 0.12282 | 9.654 | 9.7605 | 0.10641 | 68.4057 | 1.0097 | 1.0448 | 1.2558 | 0 | 24 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 146698.9106 | No Control Sample | 0 | off | off | 146698.9106 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9522 | 2.7022 | 1.7914 | 618.9625 | 1.2829 | 0.7941 | 0 | 0 | 0 | 0 | 0 | 1.8802 | FALSE | 2.5 | 2.0078
4110 | 4111 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFTeDA | 11.4376881974809 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.87 | 9.86 | Baseline | Baseline | 81331.4485 | 20056.1072 | 4.0552 | 22869.51 | 5948.5968 | 3.8445 | 3.5563 | 3.3716 | 22874.9888 | 5952.4971 | 0.99316 | 0.99236 | 9.884 | 9.8831 | 0.14151 | 0.013967 | 9.7536 | 9.7757 | 10.0427 | 9.9877 | 0.2892 | 0.21202 | 0.056537 | 0.051356 | 770.2036 | 710.8659 | 0.00052164 | 0.0012287 | FALSE | 1.0001 | TRUE | [M-H]- | 114.3769 | 9.8325 | 9.9445 | 0.11204 | 9.8378 | 9.937 | 0.099141 | 41.2509 | 1.0845 | 1.1395 | 1.4798 | 0 | 33 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 81331.4485 | No Control Sample | 0 | off | off | 81331.4485 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9493 | 2.8343 | 2.0207 | 668.9593 | 2.0779 | 1.3901 | 0 | 0 | 0 | 0 | 0 | 40.5978 | FALSE | 2.5 | 4.0552
4117 | 4118 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFBS | 10.1040664553445 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 3.5 | 3.5 | Baseline | Manual | 112164.2671 | 41939.6098 | 2.6744 | 10725.9305 | 4030.7818 | 2.661 | 10.4573 | 10.4048 | 10800.0006 | 4036.864 | 0.97501 | 0.98923 | 3.8179 | 3.816 | 9.0821 | 0.31787 | 3.5861 | 3.567 | 4.1045 | 4.0893 | 0.51842 | 0.52234 | 0.16231 | 0.15944 | 434.2132 | 768.0211 | 0.0017194 | 0.0010116 | FALSE | 1.0005 | TRUE | [M-H]- | 101.0407 | 3.647 | 4.0251 | 0.37808 | 3.678 | 3.9954 | 0.31739 | -35.5738 | 1.1215 | 1.3071 | 12.2001 | 0 | 80 | 27 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 112164.2671 | No Control Sample | 0 | off | off | 112164.2671 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9439 | 2.9943 | 0.89512 | 79.9569 | -3.5262 | -0.28195 | 0 | 0 | 0 | 0 | 0 | 1.1992 | FALSE | 2.5 | 2.6744
4119 | 4120 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFPeS | 10.7210454875483 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 6.38 | 3.5 | Baseline | Manual | 119832.7882 | 41939.6098 | 2.8573 | 24056.9235 | 4030.7818 | 5.9683 | 4.9812 | 10.4048 | 24120.8688 | 4036.864 | 0.9846 | 0.98923 | 6.3742 | 3.816 | 0.091579 | 0.0058427 | 6.2008 | 3.567 | 6.5117 | 4.0893 | 0.31088 | 0.52234 | 0.078411 | 0.15944 | 262.7285 | 768.0211 | 0.0044866 | 0.0010116 | FALSE | 1.6704 | TRUE | [M-H]- | 107.2105 | 6.293 | 6.4603 | 0.16723 | 6.3029 | 6.4479 | 0.14498 | 347.1804 | 0.94088 | 0.83569 | 9.9312 | 0 | 34 | 11 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 119832.7882 | No Control Sample | 0 | off | off | 119832.7882 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9406 | 2.3701 | 0.82701 | 79.9567 | -6.017 | -0.4811 | 0 | 0 | 0 | 0 | 0 | 1.6714 | FALSE | 2.5 | 2.8573
4121 | 4122 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | L-PFHxS | 8.84507405899535 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 8.11 | 4 | 2.0275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 92249.2125 | 55833.8533 | 1.6522 | 22892.7743 | 13628.5797 | 1.6798 | 4.0296 | 4.0968 | 23129.9443 | 13643.9894 | 0.77445 | 0.9876 | 7.348 | 7.3477 | 0.10938 | 0.0080282 | 7.2589 | 7.2141 | 7.4679 | 7.4954 | 0.20895 | 0.28137 | 0.062978 | 0.063028 | 119.6068 | 172.4205 | 0.013679 | 0.0020808 | FALSE | 1 | TRUE | [M-H]- | 109.0638 | 7.2841 | 7.4158 | 0.13177 | 7.2917 | 7.4073 | 0.11555 | -1498.7104 | 0.98275 | 0.9458 | 13.2072 | 0 | 26 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 92249.2125 | No Control Sample | 0 | off | off | 92249.2125 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9374 | 1.9576 | 0.78095 | 79.9569 | -4.3487 | -0.34771 | 0 | 0 | 0 | 0 | 0 | 1.6218 | FALSE | 2.0275 | 1.6522
4123 | 4124 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | Br-PFHxS | 1.59459809918863 | Qualifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 1.89 | 4 | 0.4725 | 0 | 7.16 | 7.35 | Manual | Baseline | 22595.2023 | 55833.8533 | 0.40469 | 4006.7311 | 13628.5797 | 0.29399 | 5.6393 | 4.0968 | 23129.9443 | 13643.9894 | 0.18977 | 0.9876 | 7.1628 | 7.3477 | 0.039229 | 0.0028088 | 7.0602 | 7.2141 | 7.2579 | 7.4954 | 0.19771 | 0.28137 | 0.090497 | 0.063028 | 20.9338 | 172.4205 | 0.018076 | 0.0020808 | TRUE | 0.97484 | TRUE | [M-H]- | 84.3703 | 7.0797 | 7.2589 | 0.17917 | 7.0857 | 7.2446 | 0.15892 | 366.3252 | 1.1857 | 1.2833 | 0.24494 | 0 | 23 | 12 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 22595.2023 | No Control Sample | 0 | off | off | 92249.2125 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9374 | 1.9576 | 0.78095 | 79.9508 | -79.5727 | -6.3624 | 0 | 0 | 0 | 0 | 0 | 1.6218 | FALSE | 0.4725 | 0.40469
4125 | 4126 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFHpS | 12.0488794670427 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7.98 | 7.35 | Baseline | Baseline | 102386.2046 | 55833.8533 | 1.8338 | 23817.9826 | 13628.5797 | 1.7476 | 4.2987 | 4.0968 | 23879.0534 | 13643.9894 | 0.9822 | 0.9876 | 7.9868 | 7.3477 | 0.084652 | 0.0067552 | 7.8644 | 7.2141 | 8.1217 | 7.4954 | 0.25728 | 0.28137 | 0.067145 | 0.063028 | 95.9859 | 172.4205 | 0.0022875 | 0.0020808 | FALSE | 1.087 | TRUE | [M-H]- | 120.4888 | 7.9224 | 8.0665 | 0.14416 | 7.9308 | 8.0558 | 0.12499 | 211.7687 | 1.1545 | 1.314 | 12.03 | 0 | 20 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 102386.2046 | No Control Sample | 0 | off | off | 102386.2046 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9344 | 2.2187 | 0.99603 | 79.9566 | -7.4185 | -0.59317 | 0 | 0 | 0 | 0 | 0 | 1.5695 | FALSE | 2.5 | 1.8338
4127 | 4128 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | L-PFOS | 7.85053776104655 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 7.88 | 4 | 1.97 | 0 | 8.46 | 8.45 | Manual | Baseline | 72497.6746 | 35527.8841 | 2.0406 | 15093.0615 | 7420.1436 | 2.0341 | 4.8034 | 4.788 | 15365.2832 | 7425.9839 | 0.65442 | 0.99026 | 8.4757 | 8.4769 | 0.18599 | 0.015735 | 8.3998 | 8.3213 | 8.5693 | 8.6175 | 0.16947 | 0.29623 | 0.076659 | 0.075127 | 48.043 | 122.8477 | 0.030695 | 0.0014302 | TRUE | 0.99986 | TRUE | [M-H]- | 99.6261 | 8.3873 | 8.5514 | 0.16409 | 8.3873 | 8.5422 | 0.15484 | -2733.7351 | 0.99589 | 0.87947 | 14.7773 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 72497.6746 | No Control Sample | 0 | off | off | 72497.6746 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9314 | 2.3589 | 1.1769 | 79.9569 | -5.6055 | -0.4482 | 0 | 0 | 0 | 0 | 0 | 1.3362 | FALSE | 1.97 | 2.0406
4129 | 4130 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | Br-PFOS | 1.65691798214728 | Qualifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 2.11 | 4 | 0.5275 | 0 | 8.3 | 8.45 | Manual | Baseline | 32757.2713 | 35527.8841 | 0.92202 | 5323.6073 | 7420.1436 | 0.71745 | 6.1532 | 4.788 | 15365.2832 | 7425.9839 | 0.29595 | 0.99026 | 8.3285 | 8.4769 | 0.34336 | 0.028499 | 8.2134 | 8.3213 | 8.4015 | 8.6175 | 0.18811 | 0.29623 | 0.10297 | 0.075127 | 16.9457 | 122.8477 | 0.0096206 | 0.0014302 | TRUE | 0.98249 | TRUE | [M-H]- | 78.5269 | 8.2373 | 8.4061 | 0.16885 | 8.2478 | 8.4061 | 0.15829 | 272.272 | 0.89797 | 0.89663 | 0.45184 | 0 | 13 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 32757.2713 | No Control Sample | 0 | off | off | 72497.6746 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9314 | 2.3589 | 1.1769 | 79.9555 | -22.5666 | -1.8044 | 0 | 0 | 0 | 0 | 0 | 1.3362 | FALSE | 0.5275 | 0.92202
4131 | 4132 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFNS | 11.5952179998697 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.85 | 8.45 | Baseline | Baseline | 91307.9077 | 35527.8841 | 2.57 | 19492.1083 | 7420.1436 | 2.6269 | 4.6844 | 4.788 | 19516.7568 | 7425.9839 | 0.98755 | 0.99026 | 8.8756 | 8.4769 | 0.28954 | 0.025625 | 8.7529 | 8.3213 | 9.1158 | 8.6175 | 0.36288 | 0.29623 | 0.072529 | 0.075127 | 149.7816 | 122.8477 | 0.00020772 | 0.0014302 | FALSE | 1.047 | TRUE | [M-H]- | 115.9522 | 8.7974 | 8.9577 | 0.16028 | 8.8068 | 8.947 | 0.14015 | -11.1576 | 1.0089 | 1.0016 | 11.9603 | 0 | 19 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Fail | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 91307.9077 | No Control Sample | 0 | off | off | 91307.9077 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9277 | 1.2638 | 0.69371 | 79.9569 | -3.05 | -0.24387 | 0 | 0 | 0 | 0 | 0 | 1.614 | FALSE | 2.5 | 2.57
4133 | 4134 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFDS | 11.3978521213255 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.19 | 8.45 | Baseline | Baseline | 92218.1895 | 35527.8841 | 2.5957 | 19934.2059 | 7420.1436 | 2.6865 | 4.6261 | 4.788 | 19939.7428 | 7425.9839 | 0.99232 | 0.99026 | 9.203 | 8.4769 | 0.14129 | 0.012984 | 9.0604 | 8.3213 | 9.3992 | 8.6175 | 0.33878 | 0.29623 | 0.072372 | 0.075127 | 212.8116 | 122.8477 | 0.00053861 | 0.0014302 | FALSE | 1.0857 | TRUE | [M-H]- | 113.9785 | 9.1297 | 9.2829 | 0.15329 | 9.1394 | 9.2709 | 0.13143 | 31.692 | 1.1796 | 1.381 | 10.5824 | 0 | 19 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 92218.1895 | No Control Sample | 0 | off | off | 92218.1895 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9241 | 0.50582 | 0.30295 | 79.9567 | -1.8989 | -0.15183 | 0 | 0 | 0 | 0 | 0 | 1.6541 | FALSE | 2.5 | 2.5957
4137 | 4138 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FBSA | 9.12488926963311 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 4.6 | 8.86 | Baseline | Baseline | 157602.3511 | 65503.1344 | 2.406 | 14309.6888 | 13874.0197 | 1.0314 | 11.0137 | 4.7213 | 14449.2359 | 13875.3196 | 0.96585 | 0.99629 | 4.9382 | 9.0164 | 7.3529 | 0.33823 | 4.672 | 8.8229 | 5.2117 | 9.2002 | 0.53965 | 0.37735 | 0.16908 | 0.073439 | 351.5942 | 140.3764 | 0.00096519 | 0.00017445 | FALSE | 0.5477 | TRUE | [M-H]- | 91.2489 | 4.7452 | 5.1339 | 0.38873 | 4.7744 | 5.1091 | 0.33472 | 25.5935 | 1.0661 | 1.1854 | 5.8759 | 0 | 96 | 32 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 157602.3511 | No Control Sample | 0 | off | off | 157602.3511 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9601 | 3.8741 | 1.1543 | 77.9651 | -3.1715 | -0.24727 | 0 | 0 | 0 | 0 | 0 | 1.3536 | FALSE | 2.5 | 2.406
4139 | 4140 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FPeSA | 11.8945588883814 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7 | 8.86 | Baseline | Baseline | 319028.7072 | 65503.1344 | 4.8704 | 67462.2291 | 13874.0197 | 4.8625 | 4.729 | 4.7213 | 67529.3331 | 13875.3196 | 0.99159 | 0.99629 | 7.0959 | 9.0164 | 1.3701 | 0.095905 | 6.9332 | 8.8229 | 7.2663 | 9.2002 | 0.33308 | 0.37735 | 0.073487 | 0.073439 | 228.5199 | 140.3764 | 0.0019323 | 0.00017445 | FALSE | 0.787 | TRUE | [M-H]- | 118.9456 | 7.0177 | 7.1754 | 0.15771 | 7.0293 | 7.1647 | 0.13542 | 391.3699 | 0.97321 | 0.9505 | 3.049 | 0 | 36 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 319028.7072 | No Control Sample | 0 | off | off | 319028.7072 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9568 | 2.8959 | 1.0076 | 77.9646 | -8.7833 | -0.68479 | 0 | 0 | 0 | 0 | 0 | 1.3129 | FALSE | 2.5 | 4.8704
4141 | 4142 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FHxSA | 12.1461487960474 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7.86 | 8.86 | Baseline | Baseline | 162569.6262 | 65503.1344 | 2.4819 | 36758.4862 | 13874.0197 | 2.6494 | 4.4226 | 4.7213 | 36782.4876 | 13875.3196 | 0.99292 | 0.99629 | 7.991 | 9.0164 | 1.6666 | 0.13099 | 7.8227 | 8.8229 | 8.1675 | 9.2002 | 0.34475 | 0.37735 | 0.068219 | 0.073439 | 124.3535 | 140.3764 | 0.0012468 | 0.00017445 | FALSE | 0.88628 | TRUE | [M-H]- | 121.4615 | 7.9221 | 8.072 | 0.14982 | 7.9306 | 8.0622 | 0.13153 | 132.9399 | 1.0357 | 1.0609 | 6.1369 | 0 | 26 | 8 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 162569.6262 | No Control Sample | 0 | off | off | 162569.6262 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9539 | 3.2936 | 1.3107 | 77.9648 | -8.0504 | -0.62765 | 0 | 0 | 0 | 0 | 0 | 1.9067 | FALSE | 2.5 | 2.4819
4143 | 4144 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FHpSA | 11.5236543644069 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.58 | 8.86 | Baseline | Baseline | 319468.9114 | 65503.1344 | 4.8772 | 73536.1592 | 13874.0197 | 5.3003 | 4.3444 | 4.7213 | 73682.3314 | 13875.3196 | 0.9871 | 0.99629 | 8.5753 | 9.0164 | 0.054958 | 0.0047154 | 8.4471 | 8.8229 | 8.7128 | 9.2002 | 0.26577 | 0.37735 | 0.067691 | 0.073439 | 109.3067 | 140.3764 | 0.00020258 | 0.00017445 | FALSE | 0.95108 | TRUE | [M-H]- | 115.2365 | 8.5039 | 8.6509 | 0.14702 | 8.5124 | 8.6397 | 0.12723 | -56.0532 | 1.0263 | 1.0163 | 2.7248 | 0 | 16 | 7 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 319468.9114 | No Control Sample | 0 | off | off | 319468.9114 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9502 | 1.7785 | 0.7967 | 77.9646 | -10.3255 | -0.80504 | 0 | 0 | 0 | 0 | 0 | 2.1016 | FALSE | 2.5 | 4.8772
4145 | 4146 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FOSA | 11.5338686582742 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.88 | 8.86 | Baseline | Baseline | 171902.6435 | 65503.1344 | 2.6243 | 37316.0292 | 13874.0197 | 2.6896 | 4.6067 | 4.7213 | 37324.7482 | 13875.3196 | 0.99237 | 0.99629 | 9.0184 | 9.0164 | 1.5587 | 0.13841 | 8.8338 | 8.8229 | 9.2124 | 9.2002 | 0.3786 | 0.37735 | 0.070795 | 0.073439 | 135.8279 | 140.3764 | 0.00046875 | 0.00017445 | FALSE | 1.0002 | TRUE | [M-H]- | 115.3387 | 8.9386 | 9.0939 | 0.15537 | 8.9469 | 9.085 | 0.13806 | 46.2018 | 0.9489 | 0.87804 | 6.1206 | 0 | 20 | 5 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 171902.6435 | No Control Sample | 0 | off | off | 171902.6435 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9472 | 1.9728 | 0.98237 | 77.9647 | -8.9849 | -0.70051 | 0 | 0 | 0 | 0 | 0 | 1.5588 | FALSE | 2.5 | 2.6243
4164 | 4165 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 4:2 FTS | 11.0286973351006 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 5.94 | 5.92 | Baseline | Baseline | 93825.3497 | 16755.0878 | 5.5998 | 16254.1355 | 2877.9774 | 5.6478 | 5.7724 | 5.8218 | 16263.376 | 2878.6903 | 0.99507 | 0.99714 | 5.926 | 5.9326 | 0.23518 | 0.01397 | 5.7124 | 5.7755 | 6.117 | 6.1291 | 0.40457 | 0.3536 | 0.084266 | 0.088266 | 212.7503 | 180.5515 | 0.0010611 | 0.00043777 | FALSE | 0.9989 | TRUE | [M-H]- | 110.287 | 5.8226 | 6.0241 | 0.20152 | 5.8396 | 6.0105 | 0.17092 | 42.6328 | 0.96394 | 0.95257 | 4.683 | 0 | 50 | 12 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 93825.3497 | No Control Sample | 0 | off | off | 93825.3497 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9754 | 3.4956 | 1.143 | 306.9685 | 3.0882 | 0.94798 | 0 | 0 | 0 | 0 | 0 | 5.3291 | FALSE | 2.5 | 5.5998
4166 | 4167 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 6:2 FTS | 11.5981625977489 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7.92 | 7.92 | Baseline | Baseline | 83532.133 | 11870.5779 | 7.0369 | 19618.4485 | 2886.0051 | 6.7978 | 4.2578 | 4.1132 | 19625.9272 | 2888.6563 | 0.9956 | 0.99295 | 7.9254 | 7.9252 | 0.067759 | 0.0053665 | 7.7864 | 7.8174 | 8.1078 | 8.0666 | 0.32142 | 0.24927 | 0.065858 | 0.062339 | 134.0348 | 116.9335 | 0.00082741 | 0.0020404 | FALSE | 1 | TRUE | [M-H]- | 115.9816 | 7.8537 | 7.995 | 0.14126 | 7.8625 | 7.9853 | 0.12282 | 50.5027 | 0.93956 | 0.84935 | 3.8948 | 0 | 25 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 83532.133 | No Control Sample | 0 | off | off | 83532.133 | 0 | off | 0 | off | 0 | off | 0 | off | 426.97 | 4.8285 | 2.0616 | 406.9635 | 5.0091 | 2.0385 | 0 | 0 | 0 | 0 | 0 | 5.5018 | FALSE | 2.5 | 7.0369
4168 | 4169 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 8:2 FTS | 10.9167710312708 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.88 | 8.89 | Baseline | Baseline | 78186.5302 | 14597.0179 | 5.3563 | 15184.3536 | 2839.9447 | 5.3467 | 5.1492 | 5.1399 | 15184.3536 | 2839.9447 | 1 | 1 | 8.8862 | 8.8879 | 0.06945 | 0.0061672 | 8.6449 | 8.6514 | 9.2543 | 9.1266 | 0.60932 | 0.4752 | 0.080496 | 0.079739 | 156.0573 | 115.1939 | 0 | 0 | FALSE | 0.9998 | TRUE | [M-H]- | 109.1677 | 8.8023 | 8.975 | 0.17275 | 8.8147 | 8.9606 | 0.14583 | 0 | 1.132 | 1.2843 | 3.1029 | 0 | 32 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Fail | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 78186.5302 | No Control Sample | 0 | off | off | 78186.5302 | 0 | off | 0 | off | 0 | off | 0 | off | 526.964 | 4.7438 | 2.4998 | 506.9566 | 1.8575 | 0.94167 | 0 | 0 | 0 | 0 | 0 | 5.7212 | FALSE | 2.5 | 5.3563
4170 | 4171 | PFAS CS7 10ng/ml | 30 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 31) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 10:2 FTS | 9.29190494135466 | Quantifiers | 10/27/2023 03:09:05 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.5 | 8.89 | Baseline | Baseline | 69279.0521 | 14597.0179 | 4.7461 | 16287.9562 | 2839.9447 | 5.7353 | 4.2534 | 5.1399 | 16287.9629 | 2839.9447 | 0.99733 | 1 | 9.506 | 8.8879 | 0.063241 | 0.0060079 | 9.3325 | 8.6514 | 9.684 | 9.1266 | 0.35157 | 0.4752 | 0.065972 | 0.079739 | 319.1366 | 115.1939 | 7.9225e-07 | 0 | FALSE | 1.0695 | TRUE | [M-H]- | 92.919 | 9.4372 | 9.5767 | 0.13954 | 9.4471 | 9.5673 | 0.12022 | 0.036705 | 0.91648 | 0.81655 | 4.8183 | 0 | 24 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 69279.0521 | No Control Sample | 0 | off | off | 69279.0521 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9568 | 2.6274 | 1.6473 | 606.951 | 5.0907 | 3.0898 | 0 | 0 | 0 | 0 | 0 | 2.1276 | FALSE | 2.5 | 4.7461
4231 | 4232 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFBA | 10.8498888025582 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 1.49 | 1.48 | Baseline | Baseline | 199016.1653 | 102258.9988 | 1.9462 | 26879.1317 | 14201.539 | 1.8927 | 7.4041 | 7.2006 | 27054.9655 | 14217.5586 | 0.96921 | 0.99244 | 1.4564 | 1.4577 | 2.2568 | 0.033627 | 1.3034 | 1.3098 | 1.6544 | 1.7107 | 0.35093 | 0.40087 | 0.11634 | 0.10818 | 228.8478 | 233.9008 | 0.015336 | 0.0031266 | FALSE | 0.99911 | TRUE | [M-H]- | 108.4989 | 1.365 | 1.6176 | 0.25261 | 1.3739 | 1.5906 | 0.21669 | 1174.6325 | 1.5633 | 2.0127 | 1.6614 | 0 | 70 | 25 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 199016.1653 | No Control Sample | 0 | off | off | 199016.1653 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9807 | 6.88 | 1.4653 | 168.9891 | 0.39378 | 0.066545 | 0 | 0 | 0 | 0 | 0 | 1.3357 | FALSE | 2.5 | 1.9462
4233 | 4234 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFPeA | 9.59364224103092 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 3.18 | 3.21 | Baseline | Manual | 169446.2506 | 78976.96 | 2.1455 | 17536.8562 | 8468.7607 | 2.0708 | 9.6623 | 9.3257 | 17595.939 | 8481.221 | 0.98343 | 0.99161 | 3.1791 | 3.1805 | 0.027243 | 0.00086631 | 2.9633 | 2.9527 | 3.4731 | 3.4472 | 0.50977 | 0.49454 | 0.14643 | 0.13906 | 596.6631 | 503.5753 | 0.0022174 | 0.0010116 | FALSE | 0.99956 | TRUE | [M-H]- | 95.9364 | 3.0297 | 3.3686 | 0.33888 | 3.0552 | 3.3408 | 0.2856 | 76.2818 | 1.1535 | 1.3533 | 2.0811 | 0 | 90 | 28 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 169446.2506 | No Control Sample | 0 | off | off | 169446.2506 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9776 | 5.9132 | 1.555 | 218.9863 | 2.2288 | 0.48807 | 0 | 0 | 0 | 0 | 0 | 0.80761 | FALSE | 2.5 | 2.1455
4235 | 4236 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFHxA | 11.1183418647112 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 179.6538 | 6.13 | 6.12 | Baseline | Baseline | 366212.8214 | 141529.0144 | 2.5875 | 62944.5885 | 24884.1946 | 2.5295 | 5.818 | 5.6875 | 62969.4793 | 24905.6542 | 0.99608 | 0.99418 | 6.1206 | 6.1192 | 0.15401 | 0.0094406 | 5.9334 | 5.9445 | 6.3718 | 6.319 | 0.43835 | 0.37447 | 0.0887 | 0.08947 | 289.8847 | 243.51 | 0.00068826 | 0.0018487 | FALSE | 1.0002 | TRUE | [M-H]- | 111.1834 | 6.0256 | 6.2151 | 0.18955 | 6.0384 | 6.2029 | 0.16456 | 98.8305 | 0.96706 | 0.93083 | 1.5096 | 0 | 48 | 12 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 366212.8214 | No Control Sample | 203843.5852 | off | off | 366212.8214 | 203843.5852 | off | 0.1 | off | 10 | off | 17.9654 | off | 312.974 | 3.7303 | 1.1675 | 268.9826 | 0.80091 | 0.21543 | 0 | 0 | 0 | 0 | 0 | 15.4799 | FALSE | 2.5 | 2.5875
4237 | 4238 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFHpA | 12.6124416795151 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 211.531 | 7.26 | 7.27 | Baseline | Baseline | 482986.1216 | 153040.6983 | 3.1559 | 116353.1267 | 36354.1097 | 3.2005 | 4.151 | 4.2097 | 116458.4681 | 36417.2602 | 0.99336 | 0.99001 | 7.2723 | 7.2713 | 0.16966 | 0.012317 | 7.1189 | 7.1216 | 7.4014 | 7.3791 | 0.28252 | 0.25757 | 0.064122 | 0.064466 | 274.6787 | 214.6503 | 0.0016194 | 0.0029534 | FALSE | 1.0001 | TRUE | [M-H]- | 126.1244 | 7.2031 | 7.3427 | 0.13956 | 7.2122 | 7.3331 | 0.12094 | 666.9339 | 1.0799 | 1.1771 | 1.3399 | 0 | 33 | 10 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 482986.1216 | No Control Sample | 228328.7703 | off | off | 482986.1216 | 228328.7703 | off | 0.1 | off | 10 | off | 21.1531 | off | 362.9705 | 2.3505 | 0.85316 | 318.9792 | 0.22724 | 0.072485 | 0 | 0 | 0 | 0 | 0 | 0.72984 | FALSE | 2.5 | 3.1559
4239 | 4240 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFOA | 11.6289642707788 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 207.2467 | 7.96 | 7.94 | Baseline | Baseline | 396826.5076 | 140705.4504 | 2.8203 | 94077.2563 | 32831.7904 | 2.8654 | 4.2181 | 4.2856 | 94147.9539 | 32849.1706 | 0.99331 | 0.99495 | 7.9599 | 7.9612 | 0.001672 | 0.00013309 | 7.7055 | 7.812 | 8.1166 | 8.1239 | 0.41103 | 0.31192 | 0.064344 | 0.065914 | 124.023 | 132.4528 | 0.0011577 | 0.001086 | FALSE | 0.99983 | TRUE | [M-H]- | 116.2896 | 7.8937 | 8.0376 | 0.14385 | 7.9029 | 8.0271 | 0.12422 | 264.9692 | 1.144 | 1.3125 | 1.2891 | 0 | 31 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 396826.5076 | No Control Sample | 191475.4056 | off | off | 396826.5076 | 191475.4056 | off | 0.1 | off | 10 | off | 20.7247 | off | 412.9678 | 3.4183 | 1.4116 | 368.9762 | 0.19168 | 0.070724 | 0 | 0 | 0 | 0 | 0 | 0.078656 | FALSE | 2.5 | 2.8203
4241 | 4242 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFNA | 9.9666740905423 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 184.1726 | 8.46 | 8.47 | Baseline | Baseline | 293098.4699 | 126590.1686 | 2.3153 | 58141.1243 | 25461.1965 | 2.2835 | 5.0412 | 4.9719 | 58148.6923 | 25468.901 | 0.99825 | 0.99666 | 8.4717 | 8.4725 | 0.1381 | 0.011684 | 8.2927 | 8.2851 | 8.7679 | 8.6808 | 0.47513 | 0.39578 | 0.079984 | 0.07826 | 110.8609 | 150.8071 | 0.00035618 | 0.00061513 | FALSE | 0.9999 | TRUE | [M-H]- | 99.6667 | 8.3892 | 8.5526 | 0.16337 | 8.4008 | 8.5414 | 0.14058 | 43.5849 | 1.0876 | 1.2122 | 1.3959 | 0 | 27 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 293098.4699 | No Control Sample | 159143.3885 | off | off | 293098.4699 | 159143.3885 | off | 0.1 | off | 10 | off | 18.4173 | off | 462.9652 | 4.1608 | 1.9263 | 418.9739 | 1.5866 | 0.66475 | 0 | 0 | 0 | 0 | 0 | 1.1613 | FALSE | 2.5 | 2.3153
4243 | 4244 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFDA | 10.9770380986566 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.87 | 8.86 | Baseline | Baseline | 270768.6123 | 99364.9933 | 2.725 | 53629.0121 | 20053.2208 | 2.6743 | 5.0489 | 4.9551 | 53639.1143 | 20062.1903 | 0.99788 | 0.99668 | 8.8849 | 8.8819 | 0.16851 | 0.014947 | 8.6747 | 8.6826 | 9.0933 | 9.0638 | 0.41862 | 0.3812 | 0.07873 | 0.076871 | 124.2243 | 131.0982 | 0.00035778 | 0.00084475 | FALSE | 1.0003 | TRUE | [M-H]- | 109.7704 | 8.8037 | 8.9718 | 0.16804 | 8.8142 | 8.958 | 0.14387 | 45.8353 | 0.93907 | 0.8207 | 0.93252 | 0 | 22 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 270768.6123 | No Control Sample | 0 | off | off | 270768.6123 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9622 | 4.2772 | 2.194 | 468.971 | 2.1293 | 0.99859 | 0 | 0 | 0 | 0 | 0 | 7.4606 | FALSE | 2.5 | 2.725
4245 | 4246 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFUdA | 12.7389988016508 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.21 | 9.2 | Baseline | Baseline | 237437.0081 | 84045.9722 | 2.8251 | 49674.1632 | 17839.853 | 2.7844 | 4.7799 | 4.7111 | 49680.2485 | 17846.2507 | 0.99854 | 0.99726 | 9.2215 | 9.2218 | 0.12488 | 0.011502 | 9.0191 | 9.0835 | 9.4778 | 9.3685 | 0.45872 | 0.28498 | 0.07472 | 0.07452 | 257.7991 | 203.8572 | 0.00025264 | 0.00071999 | FALSE | 0.99996 | TRUE | [M-H]- | 127.39 | 9.144 | 9.301 | 0.15702 | 9.1554 | 9.2887 | 0.13324 | 27.3582 | 0.9071 | 0.77402 | 0.94075 | 0 | 26 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 237437.0081 | No Control Sample | 0 | off | off | 237437.0081 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9593 | 4.371 | 2.4607 | 518.9676 | 0.65259 | 0.33867 | 0 | 0 | 0 | 0 | 0 | 0.44984 | FALSE | 2.5 | 2.8251
4247 | 4248 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFDoA | 10.3465026027458 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.47 | 9.47 | Baseline | Baseline | 202375.8829 | 75585.3699 | 2.6774 | 48705.1574 | 18549.3135 | 2.6257 | 4.1551 | 4.0748 | 48722.99 | 18560.1783 | 0.99667 | 0.99493 | 9.4857 | 9.4881 | 0.16621 | 0.01574 | 9.3629 | 9.3526 | 9.6681 | 9.6299 | 0.30525 | 0.27737 | 0.06487 | 0.064216 | 306.4188 | 404.6498 | 1.435e-05 | 0.001098 | FALSE | 0.99975 | TRUE | [M-H]- | 103.465 | 9.4188 | 9.557 | 0.1382 | 9.4292 | 9.5476 | 0.11847 | -2.2896 | 1.1549 | 1.3939 | 1.6655 | 0 | 21 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 202375.8829 | No Control Sample | 0 | off | off | 202375.8829 | 0 | off | 0 | off | 0 | off | 0 | off | 612.955 | 2.2094 | 1.3542 | 568.9653 | 3.3751 | 1.9203 | 0 | 0 | 0 | 0 | 0 | 34.9897 | FALSE | 2.5 | 2.6774
4249 | 4250 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFTrDA | 10.7857640185925 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.68 | 9.47 | Baseline | Baseline | 153780.5324 | 75585.3699 | 2.0345 | 41561.5098 | 18549.3135 | 2.2406 | 3.7001 | 4.0748 | 41581.3615 | 18560.1783 | 0.99551 | 0.99493 | 9.7006 | 9.4881 | 0.21235 | 0.020556 | 9.5532 | 9.3526 | 9.8434 | 9.6299 | 0.29018 | 0.27737 | 0.056585 | 0.064216 | 680.8776 | 404.6498 | 0.00085784 | 0.001098 | FALSE | 1.0224 | TRUE | [M-H]- | 107.8576 | 9.6427 | 9.7655 | 0.12283 | 9.6496 | 9.7571 | 0.10759 | 122.8645 | 0.96999 | 0.90744 | 1.3443 | 0 | 25 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 153780.5324 | No Control Sample | 0 | off | off | 153780.5324 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9521 | 2.4759 | 1.6414 | 618.961 | -1.0543 | -0.65254 | 0 | 0 | 0 | 0 | 0 | 0.26988 | FALSE | 2.5 | 2.0345
4251 | 4252 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFTeDA | 9.45727752143316 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.87 | 9.86 | Baseline | Baseline | 73812.3673 | 22013.5148 | 3.353 | 21713.1518 | 6201.9414 | 3.501 | 3.3994 | 3.5495 | 21740.3582 | 6201.9812 | 0.98892 | 0.99634 | 9.8787 | 9.8797 | 0.0881 | 0.0086955 | 9.7537 | 9.7759 | 9.9785 | 10.0232 | 0.22482 | 0.24732 | 0.052743 | 0.055917 | 867.415 | 603.8355 | 0.0022416 | 1.4723e-05 | FALSE | 0.99989 | TRUE | [M-H]- | 94.5728 | 9.8263 | 9.9414 | 0.11511 | 9.8329 | 9.9325 | 0.09956 | 216.4961 | 1.1967 | 1.4006 | 1.8166 | 0 | 24 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 73812.3673 | No Control Sample | 0 | off | off | 73812.3673 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9494 | 2.9942 | 2.1347 | 668.9596 | 2.5034 | 1.6747 | 0 | 0 | 0 | 0 | 0 | 40.8099 | FALSE | 2.5 | 3.353
4258 | 4259 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFBS | 10.7161874321046 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 3.5 | 3.5 | Manual | Manual | 115195.9434 | 40612.8009 | 2.8364 | 10487.2982 | 3764.0055 | 2.7862 | 10.9843 | 10.7898 | 10516.807 | 3771.3544 | 0.98676 | 0.99098 | 3.8136 | 3.8173 | 8.9605 | 0.31362 | 3.5838 | 3.557 | 4.1278 | 4.1407 | 0.54399 | 0.58371 | 0.17002 | 0.16492 | 371.144 | 973.9289 | 0.0018282 | 0.0012651 | TRUE | 0.99904 | TRUE | [M-H]- | 107.1619 | 3.6416 | 4.0283 | 0.38676 | 3.6635 | 3.9955 | 0.33191 | -35.2448 | 1.0022 | 0.94113 | 11.2936 | 0 | 85 | 28 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 115195.9434 | No Control Sample | 0 | off | off | 115195.9434 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9438 | 2.8274 | 0.84523 | 79.9568 | -5.1728 | -0.41361 | 0 | 0 | 0 | 0 | 0 | 1.0629 | FALSE | 2.5 | 2.8364
4260 | 4261 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFPeS | 12.0717406472849 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 6.38 | 3.5 | Baseline | Manual | 130661.3009 | 40612.8009 | 3.2172 | 26283.3645 | 3764.0055 | 6.9828 | 4.9713 | 10.7898 | 26317.6593 | 3771.3544 | 0.9896 | 0.99098 | 6.3756 | 3.8173 | 0.069494 | 0.0044337 | 6.2195 | 3.557 | 6.5496 | 4.1407 | 0.33012 | 0.58371 | 0.075564 | 0.16492 | 264.3914 | 973.9289 | 0.0025693 | 0.0012651 | FALSE | 1.6702 | TRUE | [M-H]- | 120.7174 | 6.2915 | 6.4612 | 0.16971 | 6.3024 | 6.4494 | 0.14702 | 204.5622 | 0.93916 | 0.85106 | 7.0934 | 0 | 36 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 130661.3009 | No Control Sample | 0 | off | off | 130661.3009 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9406 | 2.382 | 0.83118 | 79.9568 | -5.4573 | -0.43635 | 0 | 0 | 0 | 0 | 0 | 1.0632 | FALSE | 2.5 | 3.2172
4262 | 4263 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | L-PFHxS | 8.47142780724467 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 8.11 | 4 | 2.0275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 92578.3828 | 58504.515 | 1.5824 | 22496.8738 | 14519.7375 | 1.5494 | 4.1152 | 4.0293 | 22835.5248 | 14535.8951 | 0.75657 | 0.98893 | 7.3474 | 7.3456 | 0.10066 | 0.0073888 | 7.2609 | 7.2256 | 7.4983 | 7.4894 | 0.2374 | 0.26387 | 0.065079 | 0.062657 | 121.1168 | 183.4428 | 0.018373 | 0.0023797 | FALSE | 1.0002 | TRUE | [M-H]- | 104.4566 | 7.2854 | 7.4155 | 0.13012 | 7.2927 | 7.4074 | 0.1147 | -1741.0827 | 1.0909 | 1.1904 | 13.1962 | 0 | 29 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 92578.3828 | No Control Sample | 0 | off | off | 92578.3828 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9377 | 2.6694 | 1.0649 | 79.9566 | -6.9444 | -0.55526 | 0 | 0 | 0 | 0 | 0 | 1.4985 | FALSE | 2.0275 | 1.5824
4264 | 4265 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | Br-PFHxS | 1.63172842177195 | Qualifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 1.89 | 4 | 0.4725 | 0 | 7.16 | 7.35 | Manual | Baseline | 24227.2797 | 58504.515 | 0.41411 | 4391.8533 | 14519.7375 | 0.30247 | 5.5164 | 4.0293 | 22835.5248 | 14535.8951 | 0.19801 | 0.98893 | 7.1611 | 7.3456 | 0.015975 | 0.0011438 | 7.0573 | 7.2256 | 7.2606 | 7.4894 | 0.20322 | 0.26387 | 0.085815 | 0.062657 | 23.6445 | 183.4428 | 0.03033 | 0.0023797 | TRUE | 0.97488 | TRUE | [M-H]- | 86.3348 | 7.079 | 7.2609 | 0.18195 | 7.0864 | 7.2571 | 0.17063 | 655.4689 | 1.1596 | 1.4033 | 0.26169 | 0 | 23 | 12 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 24227.2797 | No Control Sample | 0 | off | off | 92578.3828 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9377 | 2.6694 | 1.0649 | 79.9504 | -84.6542 | -6.7687 | 0 | 0 | 0 | 0 | 0 | 1.4985 | FALSE | 0.4725 | 0.41411
4266 | 4267 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFHpS | 11.9037553045622 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7.98 | 7.35 | Baseline | Baseline | 105991.3826 | 58504.515 | 1.8117 | 24943.097 | 14519.7375 | 1.7179 | 4.2493 | 4.0293 | 24952.7871 | 14535.8951 | 0.99475 | 0.98893 | 7.9845 | 7.3456 | 0.056782 | 0.0045312 | 7.848 | 7.2256 | 8.2157 | 7.4894 | 0.3677 | 0.26387 | 0.065933 | 0.062657 | 110.8187 | 183.4428 | 0.001035 | 0.0023797 | FALSE | 1.087 | TRUE | [M-H]- | 119.0376 | 7.9191 | 8.0616 | 0.14253 | 7.9277 | 8.0519 | 0.1242 | 70.2115 | 1.1315 | 1.2831 | 11.7217 | 0 | 27 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 105991.3826 | No Control Sample | 0 | off | off | 105991.3826 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9343 | 2.0404 | 0.916 | 79.9568 | -5.0651 | -0.40499 | 0 | 0 | 0 | 0 | 0 | 1.6644 | FALSE | 2.5 | 1.8117
4268 | 4269 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | L-PFOS | 8.24905308538985 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 7.88 | 4 | 1.97 | 0 | 8.46 | 8.45 | Manual | Baseline | 76599.3797 | 35724.4728 | 2.1442 | 16335.4524 | 7335.7734 | 2.2268 | 4.6891 | 4.8699 | 16532.1176 | 7342.4081 | 0.6682 | 0.98999 | 8.4698 | 8.4716 | 0.11532 | 0.009756 | 8.392 | 8.3378 | 8.5654 | 8.6007 | 0.17339 | 0.26292 | 0.074631 | 0.076804 | 65.0714 | 99.318 | 0.0047072 | 0.0015745 | TRUE | 0.99978 | TRUE | [M-H]- | 104.6834 | 8.3862 | 8.5442 | 0.15794 | 8.3862 | 8.5372 | 0.151 | -443.463 | 0.95741 | 0.83064 | 13.9816 | 0 | 11 | 5 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 76599.3797 | No Control Sample | 0 | off | off | 76599.3797 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9311 | 1.7958 | 0.89596 | 79.9568 | -5.771 | -0.46143 | 0 | 0 | 0 | 0 | 0 | 1.5229 | FALSE | 1.97 | 2.1442
4270 | 4271 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | Br-PFOS | 1.65601487345667 | Qualifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 2.11 | 4 | 0.5275 | 0 | 8.3 | 8.45 | Manual | Baseline | 32920.5759 | 35724.4728 | 0.92151 | 5793.3715 | 7335.7734 | 0.78974 | 5.6825 | 4.8699 | 16532.1176 | 7342.4081 | 0.28724 | 0.98999 | 8.3243 | 8.4716 | 0.29306 | 0.024324 | 8.2016 | 8.3378 | 8.3946 | 8.6007 | 0.19294 | 0.26292 | 0.093467 | 0.076804 | 23.0776 | 99.318 | 0.015565 | 0.0015745 | TRUE | 0.98262 | TRUE | [M-H]- | 78.4841 | 8.234 | 8.4051 | 0.17112 | 8.2431 | 8.4051 | 0.162 | 467.3559 | 0.88985 | 0.86134 | 0.42978 | 0 | 13 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 32920.5759 | No Control Sample | 0 | off | off | 76599.3797 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9311 | 1.7958 | 0.89596 | 79.9579 | 6.8881 | 0.55075 | 0 | 0 | 0 | 0 | 0 | 1.5229 | FALSE | 0.5275 | 0.92151
4272 | 4273 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFNS | 11.4217669988892 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.85 | 8.45 | Baseline | Baseline | 90439.7293 | 35724.4728 | 2.5316 | 18640.2349 | 7335.7734 | 2.541 | 4.8519 | 4.8699 | 18652.3457 | 7342.4081 | 0.99019 | 0.98999 | 8.8711 | 8.4716 | 0.23837 | 0.021096 | 8.715 | 8.3378 | 9.0412 | 8.6007 | 0.32613 | 0.26292 | 0.075682 | 0.076804 | 111.3722 | 99.318 | 0.00023241 | 0.0015745 | FALSE | 1.0472 | TRUE | [M-H]- | 114.2177 | 8.7943 | 8.9549 | 0.16054 | 8.8011 | 8.9416 | 0.14048 | 13.2833 | 0.96948 | 0.84695 | 10.7068 | 0 | 17 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 90439.7293 | No Control Sample | 0 | off | off | 90439.7293 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9272 | 0.298 | 0.16358 | 79.9571 | -0.2205 | -0.017631 | 0 | 0 | 0 | 0 | 0 | 1.4753 | FALSE | 2.5 | 2.5316
4274 | 4275 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFDS | 11.1912491560357 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.19 | 8.45 | Baseline | Baseline | 91047.6252 | 35724.4728 | 2.5486 | 19966.6664 | 7335.7734 | 2.7218 | 4.56 | 4.8699 | 19976.9635 | 7342.4081 | 0.99051 | 0.98999 | 9.1974 | 8.4716 | 0.080285 | 0.0073782 | 9.0608 | 8.3378 | 9.3496 | 8.6007 | 0.2888 | 0.26292 | 0.071023 | 0.076804 | 226.1444 | 99.318 | 0.0010595 | 0.0015745 | FALSE | 1.0857 | TRUE | [M-H]- | 111.9125 | 9.124 | 9.279 | 0.155 | 9.1349 | 9.2689 | 0.13396 | 73.2496 | 1.0807 | 1.204 | 12.0082 | 0 | 16 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 91047.6252 | No Control Sample | 0 | off | off | 91047.6252 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9245 | 1.0721 | 0.64214 | 79.9568 | -1.873 | -0.14976 | 0 | 0 | 0 | 0 | 0 | 1.3053 | FALSE | 2.5 | 2.5486
4278 | 4279 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FBSA | 8.69951908418324 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 4.6 | 8.86 | Baseline | Baseline | 151545.4397 | 66065.4839 | 2.2939 | 13722.6152 | 14545.1943 | 0.94345 | 11.0435 | 4.5421 | 14026.2704 | 14546.2679 | 0.93306 | 0.99802 | 4.9351 | 9.0122 | 7.2856 | 0.33514 | 4.6433 | 8.8021 | 5.1594 | 9.2601 | 0.51605 | 0.45795 | 0.17205 | 0.069787 | 334.8752 | 159.3354 | 0.032615 | 0.00014885 | FALSE | 0.5476 | TRUE | [M-H]- | 86.9952 | 4.7548 | 5.1323 | 0.37745 | 4.7764 | 5.1079 | 0.33155 | 867.2969 | 1.0957 | 1.1996 | 5.7941 | 0 | 92 | 32 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 151545.4397 | No Control Sample | 0 | off | off | 151545.4397 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9603 | 4.361 | 1.2994 | 77.965 | -4.4211 | -0.34469 | 0 | 0 | 0 | 0 | 0 | 0.54409 | FALSE | 2.5 | 2.2939
4280 | 4281 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FPeSA | 11.5513232156074 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7 | 8.86 | Baseline | Baseline | 312482.4974 | 66065.4839 | 4.7299 | 65499.8652 | 14545.1943 | 4.5032 | 4.7707 | 4.5421 | 65562.9964 | 14546.2679 | 0.99211 | 0.99802 | 7.0974 | 9.0122 | 1.3909 | 0.097366 | 6.946 | 8.8021 | 7.2601 | 9.2601 | 0.31402 | 0.45795 | 0.073536 | 0.069787 | 246.4012 | 159.3354 | 0.0017348 | 0.00014885 | FALSE | 0.78753 | TRUE | [M-H]- | 115.5132 | 7.0175 | 7.1771 | 0.15956 | 7.0289 | 7.1663 | 0.13736 | 361.8492 | 0.95154 | 0.8952 | 2.8746 | 0 | 34 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 312482.4974 | No Control Sample | 0 | off | off | 312482.4974 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9567 | 2.6187 | 0.91119 | 77.9647 | -7.5079 | -0.58535 | 0 | 0 | 0 | 0 | 0 | 1.8866 | FALSE | 2.5 | 4.7299
4282 | 4283 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FHxSA | 12.0308328038428 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7.86 | 8.86 | Baseline | Baseline | 162408.6062 | 66065.4839 | 2.4583 | 36338.2668 | 14545.1943 | 2.4983 | 4.4694 | 4.5421 | 36357.9767 | 14546.2679 | 0.99391 | 0.99802 | 7.9881 | 9.0122 | 1.6292 | 0.12806 | 7.8073 | 8.8021 | 8.1662 | 9.2601 | 0.3589 | 0.45795 | 0.069281 | 0.069787 | 145.5942 | 159.3354 | 0.00099259 | 0.00014885 | FALSE | 0.88636 | TRUE | [M-H]- | 120.3083 | 7.9183 | 8.0693 | 0.15096 | 7.9279 | 8.0593 | 0.13145 | 100.499 | 1.0279 | 1.0582 | 6.0892 | 0 | 27 | 8 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 162408.6062 | No Control Sample | 0 | off | off | 162408.6062 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9537 | 2.7903 | 1.1104 | 77.9648 | -7.826 | -0.61016 | 0 | 0 | 0 | 0 | 0 | 1.7783 | FALSE | 2.5 | 2.4583
4284 | 4285 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FHpSA | 11.0637746968558 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.58 | 8.86 | Baseline | Baseline | 309352.9339 | 66065.4839 | 4.6825 | 67991.656 | 14545.1943 | 4.6745 | 4.5499 | 4.5421 | 68012.5143 | 14546.2679 | 0.99626 | 0.99802 | 8.5695 | 9.0122 | 0.12251 | 0.010511 | 8.3284 | 8.8021 | 8.8332 | 9.2601 | 0.50485 | 0.45795 | 0.07111 | 0.069787 | 80.6573 | 159.3354 | 0.00061742 | 0.00014885 | FALSE | 0.95087 | TRUE | [M-H]- | 110.6377 | 8.4963 | 8.6474 | 0.15111 | 8.5079 | 8.6368 | 0.12893 | 83.1525 | 0.96209 | 0.92825 | 2.8485 | 0 | 28 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 309352.9339 | No Control Sample | 0 | off | off | 309352.9339 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9501 | 1.634 | 0.73193 | 77.9647 | -8.5391 | -0.66575 | 0 | 0 | 0 | 0 | 0 | 1.7545 | FALSE | 2.5 | 4.6825
4286 | 4287 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FOSA | 12.0244904095741 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.88 | 8.86 | Baseline | Baseline | 180753.5235 | 66065.4839 | 2.736 | 37808.7806 | 14545.1943 | 2.5994 | 4.7807 | 4.5421 | 37816.1757 | 14546.2679 | 0.99385 | 0.99802 | 9.013 | 9.0122 | 1.4973 | 0.13296 | 8.7922 | 8.8021 | 9.2332 | 9.2601 | 0.441 | 0.45795 | 0.074351 | 0.069787 | 114.3721 | 159.3354 | 0.00037258 | 0.00014885 | FALSE | 1.0001 | TRUE | [M-H]- | 120.2449 | 8.9345 | 9.0912 | 0.15669 | 8.9423 | 9.0797 | 0.13735 | 31.9431 | 0.90801 | 0.74937 | 5.5112 | 0 | 23 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 180753.5235 | No Control Sample | 0 | off | off | 180753.5235 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9464 | 0.44408 | 0.22113 | 77.9647 | -8.4861 | -0.66162 | 0 | 0 | 0 | 0 | 0 | 1.2783 | FALSE | 2.5 | 2.736
4305 | 4306 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 4:2 FTS | 11.0470189636853 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 5.94 | 5.92 | Baseline | Baseline | 89594.8447 | 15973.0797 | 5.6091 | 14930.0647 | 2655.3753 | 5.6226 | 6.001 | 6.0154 | 14931.3724 | 2655.3753 | 0.99817 | 0.99928 | 5.9277 | 5.9248 | 0.20699 | 0.012295 | 5.7194 | 5.7353 | 6.1639 | 6.1383 | 0.4445 | 0.40295 | 0.092959 | 0.0901 | 214.8769 | 180.2234 | 0.00018008 | 0 | FALSE | 1.0005 | TRUE | [M-H]- | 110.4702 | 5.8237 | 6.0282 | 0.2045 | 5.838 | 6.0155 | 0.17757 | 6.0485 | 0.91716 | 0.82653 | 4.9941 | 0 | 54 | 13 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 89594.8447 | No Control Sample | 0 | off | off | 89594.8447 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9757 | 4.1793 | 1.3665 | 306.9691 | 5.0205 | 1.5411 | 0 | 0 | 0 | 0 | 0 | 5.3291 | FALSE | 2.5 | 5.6091
4307 | 4308 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 6:2 FTS | 10.649344419283 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 7.92 | 7.92 | Baseline | Baseline | 84484.9103 | 13075.6642 | 6.4612 | 20666.1593 | 2980.6895 | 6.9333 | 4.0881 | 4.3868 | 20681.6841 | 2980.7323 | 0.99358 | 0.99676 | 7.9204 | 7.9228 | 0.0050554 | 0.00040038 | 7.7818 | 7.7437 | 8.074 | 8.1278 | 0.29218 | 0.38407 | 0.06242 | 0.066988 | 116.324 | 133.2221 | 0.0016306 | 2.9737e-05 | FALSE | 0.9997 | TRUE | [M-H]- | 106.4934 | 7.8519 | 7.9934 | 0.1415 | 7.8619 | 7.9823 | 0.12048 | 115.3294 | 1.1255 | 1.2786 | 3.741 | 0 | 23 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 84484.9103 | No Control Sample | 0 | off | off | 84484.9103 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9697 | 4.307 | 1.839 | 406.9632 | 4.2682 | 1.737 | 0 | 0 | 0 | 0 | 0 | 5.5018 | FALSE | 2.5 | 6.4612
4309 | 4310 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 8:2 FTS | 10.8823399564007 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.88 | 8.89 | Baseline | Baseline | 76974.9568 | 14416.2919 | 5.3394 | 15239.604 | 2842.7792 | 5.3608 | 5.051 | 5.0712 | 15248.2735 | 2843.4914 | 0.99522 | 0.99436 | 8.88 | 8.8814 | 0.00016758 | 1.4881e-05 | 8.6791 | 8.7055 | 9.0429 | 9.0325 | 0.36382 | 0.32705 | 0.078776 | 0.079337 | 145.34 | 119.4728 | 0.0010254 | 0.00044727 | FALSE | 0.99985 | TRUE | [M-H]- | 108.8234 | 8.7979 | 8.9668 | 0.1689 | 8.8069 | 8.956 | 0.14916 | 42.9507 | 1.0435 | 1.072 | 2.6669 | 0 | 19 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 76974.9568 | No Control Sample | 0 | off | off | 76974.9568 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9641 | 4.929 | 2.5974 | 506.9578 | 4.1886 | 2.1234 | 0 | 0 | 0 | 0 | 0 | 5.7212 | FALSE | 2.5 | 5.3394
4311 | 4312 | PFAS CS7 10ng/ml | 31 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 32) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 10:2 FTS | 9.1669743295948 | Quantifiers | 10/27/2023 03:25:43 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.5 | 8.89 | Baseline | Baseline | 67501.3751 | 14416.2919 | 4.6823 | 16674.1248 | 2842.7792 | 5.8654 | 4.0483 | 5.0712 | 16674.3194 | 2843.4914 | 0.99759 | 0.99436 | 9.5018 | 8.8814 | 0.019474 | 0.00185 | 9.3508 | 8.7055 | 9.6966 | 9.0325 | 0.34578 | 0.32705 | 0.062439 | 0.079337 | 314.046 | 119.4728 | 2.735e-05 | 0.00044727 | FALSE | 1.0699 | TRUE | [M-H]- | 91.6697 | 9.4336 | 9.5694 | 0.13583 | 9.441 | 9.5619 | 0.12092 | 1.3189 | 1.0558 | 1.1252 | 5.1862 | 0 | 24 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 67501.3751 | No Control Sample | 0 | off | off | 67501.3751 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9566 | 2.4162 | 1.5148 | 606.9513 | 5.6564 | 3.4331 | 0 | 0 | 0 | 0 | 0 | 2.3599 | FALSE | 2.5 | 4.6823
4372 | 4373 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFBA | 11.6689214667188 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 1.49 | 1.48 | Baseline | Manual | 206988.7011 | 98890.4588 | 2.0931 | 29428.3556 | 14281.4494 | 2.0606 | 7.0336 | 6.9244 | 29469.7939 | 14314.6524 | 0.98913 | 0.98818 | 1.4652 | 1.4643 | 1.6629 | 0.024776 | 1.3228 | 1.3443 | 1.7236 | 1.6778 | 0.40082 | 0.33351 | 0.10622 | 0.1035 | 295.22 | 195.0765 | 0.0011831 | 0.0017211 | FALSE | 1.0007 | TRUE | [M-H]- | 116.6892 | 1.3772 | 1.6209 | 0.24367 | 1.3852 | 1.5931 | 0.20788 | 86.866 | 1.3086 | 1.442 | 1.5711 | 0 | 80 | 23 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 206988.7011 | No Control Sample | 0 | off | off | 206988.7011 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9806 | 6.3621 | 1.355 | 168.9889 | -0.86128 | -0.14555 | 0 | 0 | 0 | 0 | 0 | 0.86646 | FALSE | 2.5 | 2.0931
4374 | 4375 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFPeA | 10.0391713344047 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 3.18 | 3.21 | Baseline | Manual | 165161.4811 | 73563.5825 | 2.2452 | 17805.2203 | 7647.7125 | 2.3282 | 9.276 | 9.619 | 18079.371 | 7675.3384 | 0.94868 | 0.98289 | 3.2042 | 3.1997 | 0.76083 | 0.024194 | 3.0217 | 3.0011 | 3.4358 | 3.4566 | 0.41405 | 0.45555 | 0.1446 | 0.15306 | 642.8464 | 543.531 | 0.0017689 | 0.00081098 | FALSE | 1.0014 | TRUE | [M-H]- | 100.3917 | 3.0616 | 3.3785 | 0.31691 | 3.0802 | 3.3517 | 0.27153 | 76.0692 | 1.0884 | 1.1386 | 2.0976 | 0 | 72 | 27 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Pass | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 165161.4811 | No Control Sample | 0 | off | off | 165161.4811 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9776 | 5.9134 | 1.5551 | 218.9862 | 2.0523 | 0.44942 | 0 | 0 | 0 | 0 | 0 | 0.042603 | FALSE | 2.5 | 2.2452
4376 | 4377 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFHxA | 10.6768820906297 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 175.66 | 6.13 | 6.12 | Baseline | Baseline | 358071.5585 | 144104.4328 | 2.4848 | 64073.7838 | 26789.0699 | 2.3918 | 5.5884 | 5.3792 | 64178.8124 | 26820.4005 | 0.99014 | 0.99228 | 6.1471 | 6.1416 | 0.2783 | 0.01706 | 5.9655 | 5.9596 | 6.3052 | 6.3172 | 0.3397 | 0.35755 | 0.086424 | 0.081628 | 285.0403 | 241.5461 | 0.003131 | 0.0022166 | FALSE | 1.0009 | TRUE | [M-H]- | 106.7688 | 6.051 | 6.2383 | 0.18733 | 6.0647 | 6.2256 | 0.1609 | 590.5608 | 1.0449 | 1.1197 | 1.558 | 0 | 37 | 11 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 358071.5585 | No Control Sample | 203843.5852 | off | off | 358071.5585 | 203843.5852 | off | 0.1 | off | 10 | off | 17.566 | off | 312.9738 | 3.2301 | 1.0109 | 268.9825 | 0.35987 | 0.096798 | 0 | 0 | 0 | 0 | 0 | 15.5527 | FALSE | 2.5 | 2.4848
4378 | 4379 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFHpA | 12.3349111292759 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 206.788 | 7.26 | 7.27 | Baseline | Baseline | 472156.6117 | 152975.3688 | 3.0865 | 114007.3181 | 35030.7245 | 3.2545 | 4.1415 | 4.3669 | 114057.1771 | 35046.1824 | 0.99611 | 0.99611 | 7.2883 | 7.2858 | 0.38951 | 0.028278 | 7.1436 | 7.1462 | 7.4747 | 7.4353 | 0.33112 | 0.28905 | 0.06351 | 0.068613 | 261.0955 | 198.9025 | 0.00081291 | 0.00090965 | FALSE | 1.0003 | TRUE | [M-H]- | 123.3491 | 7.2185 | 7.3593 | 0.14075 | 7.2283 | 7.3493 | 0.12108 | 279.8945 | 0.97703 | 0.94335 | 1.4785 | 0 | 39 | 9 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 472156.6117 | No Control Sample | 228328.7703 | off | off | 472156.6117 | 228328.7703 | off | 0.1 | off | 10 | off | 20.6788 | off | 362.9705 | 2.4347 | 0.88372 | 318.9795 | 1.1329 | 0.36137 | 0 | 0 | 0 | 0 | 0 | 0.60233 | FALSE | 2.5 | 3.0865
4380 | 4381 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFOA | 11.1785050773461 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 208.2402 | 7.96 | 7.94 | Baseline | Baseline | 398728.6918 | 147077.0943 | 2.711 | 96295.4627 | 34550.5785 | 2.7871 | 4.1407 | 4.2569 | 96346.8252 | 34552.2482 | 0.99479 | 0.99907 | 7.9727 | 7.9735 | 0.15975 | 0.012716 | 7.835 | 7.8279 | 8.1717 | 8.2267 | 0.33663 | 0.3988 | 0.063195 | 0.064331 | 160.8691 | 117.0109 | 1.3042e-05 | 0.00012726 | FALSE | 0.9999 | TRUE | [M-H]- | 111.7851 | 7.9078 | 8.054 | 0.14624 | 7.9168 | 8.0418 | 0.125 | 3.7307 | 1.1408 | 1.27 | 1.2812 | 0 | 25 | 7 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 398728.6918 | No Control Sample | 191475.4056 | off | off | 398728.6918 | 191475.4056 | off | 0.1 | off | 10 | off | 20.824 | off | 412.9679 | 3.5013 | 1.4459 | 368.9764 | 0.82576 | 0.30469 | 0 | 0 | 0 | 0 | 0 | 1.0107 | FALSE | 2.5 | 2.711
4382 | 4383 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFNA | 10.7263956670816 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 204.1995 | 8.46 | 8.47 | Baseline | Baseline | 324970.0829 | 130414.6117 | 2.4918 | 66127.7039 | 26546.7083 | 2.491 | 4.9143 | 4.9126 | 66143.0633 | 26551.4646 | 0.9977 | 0.99835 | 8.4867 | 8.484 | 0.31505 | 0.026654 | 8.2999 | 8.3242 | 8.6989 | 8.7441 | 0.39905 | 0.41997 | 0.077348 | 0.077406 | 175.8261 | 190.0561 | 0.00029561 | 0.00014128 | FALSE | 1.0003 | TRUE | [M-H]- | 107.264 | 8.4039 | 8.5626 | 0.15879 | 8.4129 | 8.5543 | 0.14137 | 48.987 | 0.89504 | 0.77489 | 1.3131 | 0 | 23 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 324970.0829 | No Control Sample | 159143.3885 | off | off | 324970.0829 | 159143.3885 | off | 0.1 | off | 10 | off | 20.42 | off | 462.9646 | 3.0185 | 1.3975 | 418.973 | -0.48609 | -0.20366 | 0 | 0 | 0 | 0 | 0 | 0.93622 | FALSE | 2.5 | 2.4918
4384 | 4385 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFDA | 10.9425780779616 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.87 | 8.86 | Baseline | Baseline | 291911.4955 | 107461.2291 | 2.7164 | 59976.4547 | 22613.2364 | 2.6523 | 4.8671 | 4.7521 | 60015.6887 | 22613.2388 | 0.99527 | 0.9996 | 8.8971 | 8.8929 | 0.30538 | 0.027088 | 8.7372 | 8.6711 | 9.0449 | 9.169 | 0.3077 | 0.49785 | 0.076158 | 0.073631 | 123.075 | 138.4874 | 0.0012144 | 2.3543e-07 | FALSE | 1.0005 | TRUE | [M-H]- | 109.4258 | 8.8163 | 8.9798 | 0.16343 | 8.8244 | 8.9668 | 0.1424 | 236.7061 | 0.96509 | 0.85828 | 0.75881 | 0 | 16 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 291911.4955 | No Control Sample | 0 | off | off | 291911.4955 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9623 | 4.4052 | 2.2597 | 468.9711 | 2.4187 | 1.1343 | 0 | 0 | 0 | 0 | 0 | 9.5536 | FALSE | 2.5 | 2.7164
4386 | 4387 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFUdA | 10.7552443758107 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.21 | 9.2 | Baseline | Baseline | 211082.3603 | 88498.4064 | 2.3852 | 45338.5444 | 19809.9138 | 2.2887 | 4.6557 | 4.4674 | 45348.358 | 19824.2698 | 0.99749 | 0.99508 | 9.2279 | 9.2281 | 0.19381 | 0.01785 | 9.027 | 9.0916 | 9.4696 | 9.3602 | 0.4426 | 0.26855 | 0.073046 | 0.069695 | 239.9084 | 174.0596 | 0.00048371 | 0.0013647 | FALSE | 0.99997 | TRUE | [M-H]- | 107.5524 | 9.1546 | 9.3091 | 0.15457 | 9.1649 | 9.2968 | 0.1319 | 49.5501 | 1.1803 | 1.3925 | 1.2954 | 0 | 25 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 211082.3603 | No Control Sample | 0 | off | off | 211082.3603 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9588 | 3.3909 | 1.9089 | 518.9677 | 0.7245 | 0.37599 | 0 | 0 | 0 | 0 | 0 | 0.81088 | FALSE | 2.5 | 2.3852
4388 | 4389 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFDoA | 11.3479492850116 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.47 | 9.47 | Baseline | Baseline | 213780.6227 | 72798.6808 | 2.9366 | 52561.5847 | 17371.2534 | 3.0258 | 4.0672 | 4.1908 | 52620.4361 | 17377.7636 | 0.99335 | 0.99511 | 9.4947 | 9.496 | 0.26042 | 0.024661 | 9.3713 | 9.3438 | 9.6041 | 9.6381 | 0.23277 | 0.29433 | 0.063687 | 0.066256 | 292.8855 | 445.3553 | 0.0020269 | 0.00072573 | FALSE | 0.99986 | TRUE | [M-H]- | 113.4795 | 9.4278 | 9.5605 | 0.13272 | 9.438 | 9.5536 | 0.11556 | 457.6883 | 1.1224 | 1.3633 | 1.9005 | 0 | 16 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Fail | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 213780.6227 | No Control Sample | 0 | off | off | 213780.6227 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9549 | 2.11 | 1.2934 | 568.9645 | 1.8506 | 1.0529 | 0 | 0 | 0 | 0 | 0 | 32.2642 | FALSE | 2.5 | 2.9366
4390 | 4391 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFTrDA | 12.4133370911157 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.68 | 9.47 | Baseline | Baseline | 170460.8939 | 72798.6808 | 2.3415 | 49056.2611 | 17371.2534 | 2.824 | 3.4748 | 4.1908 | 49081.9279 | 17377.7636 | 0.99528 | 0.99511 | 9.708 | 9.496 | 0.28946 | 0.02802 | 9.5758 | 9.3438 | 9.8471 | 9.6381 | 0.27128 | 0.29433 | 0.053638 | 0.066256 | 588.3177 | 445.3553 | 0.0010269 | 0.00072573 | FALSE | 1.0223 | TRUE | [M-H]- | 124.1334 | 9.651 | 9.7688 | 0.11776 | 9.658 | 9.7604 | 0.10241 | 185.688 | 0.97324 | 0.91243 | 1.1022 | 0 | 24 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 170460.8939 | No Control Sample | 0 | off | off | 170460.8939 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9528 | 3.5954 | 2.3836 | 618.9614 | -0.52235 | -0.32331 | 0 | 0 | 0 | 0 | 0 | 1.7545 | FALSE | 2.5 | 2.3415
4392 | 4393 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFTeDA | 10.0066377170487 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.87 | 9.86 | Baseline | Baseline | 85406.4706 | 24072.9319 | 3.5478 | 25959.0669 | 7321.9776 | 3.5454 | 3.29 | 3.2878 | 25967.6914 | 7335.8713 | 0.99084 | 0.99046 | 9.883 | 9.886 | 0.13189 | 0.013017 | 9.7575 | 9.7898 | 9.997 | 9.9736 | 0.23953 | 0.18382 | 0.050341 | 0.050522 | 914.8869 | 919.3951 | 0.00065237 | 0.0034211 | FALSE | 0.9997 | TRUE | [M-H]- | 100.0664 | 9.8323 | 9.9453 | 0.11296 | 9.8393 | 9.9372 | 0.097956 | 70.6994 | 1.2318 | 1.5161 | 1.6366 | 0 | 26 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 85406.4706 | No Control Sample | 0 | off | off | 85406.4706 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9492 | 2.7573 | 1.9658 | 668.9595 | 2.4645 | 1.6487 | 0 | 0 | 0 | 0 | 0 | 36.0288 | FALSE | 2.5 | 3.5478
4399 | 4400 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFBS | 9.7013746508089 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 3.5 | 3.5 | Baseline | Manual | 103790.5256 | 40419.4586 | 2.5678 | 9896.6685 | 3693.0607 | 2.6798 | 10.4874 | 10.9447 | 10210.9469 | 3706.765 | 0.91154 | 0.9871 | 3.8454 | 3.8456 | 9.8689 | 0.34541 | 3.6006 | 3.6187 | 4.0471 | 4.1732 | 0.44653 | 0.55449 | 0.16081 | 0.16666 | 444.6432 | 893.9515 | 0.04579 | 0.0022811 | FALSE | 0.99994 | TRUE | [M-H]- | 97.0137 | 3.6898 | 4.0317 | 0.34193 | 3.7034 | 4.0101 | 0.30663 | 1014.8663 | 1.1614 | 1.296 | 13.2349 | 0 | 69 | 26 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 103790.5256 | No Control Sample | 0 | off | off | 103790.5256 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9438 | 2.817 | 0.84211 | 79.9568 | -5.5492 | -0.4437 | 0 | 0 | 0 | 0 | 0 | 0.88994 | FALSE | 2.5 | 2.5678
4401 | 4402 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFPeS | 12.1155866865376 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 6.38 | 3.5 | Baseline | Manual | 130511.59 | 40419.4586 | 3.2289 | 26284.3637 | 3693.0607 | 7.1172 | 4.9654 | 10.9447 | 26327.6257 | 3706.765 | 0.98871 | 0.9871 | 6.3963 | 3.8456 | 0.25553 | 0.016303 | 6.2563 | 3.6187 | 6.5483 | 4.1732 | 0.29202 | 0.55449 | 0.076872 | 0.16666 | 279.0946 | 893.9515 | 0.0027945 | 0.0022811 | FALSE | 1.6633 | TRUE | [M-H]- | 121.1559 | 6.3145 | 6.4818 | 0.1673 | 6.3262 | 6.4706 | 0.14441 | 251.5286 | 0.98766 | 0.97805 | 8.3387 | 0 | 32 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 130511.59 | No Control Sample | 0 | off | off | 130511.59 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9408 | 2.7781 | 0.96939 | 79.9568 | -5.2403 | -0.419 | 0 | 0 | 0 | 0 | 0 | 2.0189 | FALSE | 2.5 | 3.2289
4403 | 4404 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | L-PFHxS | 8.73108921632125 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 8.11 | 4 | 2.0275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 93814.8681 | 57522.7518 | 1.6309 | 23587.7044 | 14089.0354 | 1.6742 | 3.9773 | 4.0828 | 23843.0342 | 14093.8918 | 0.77119 | 0.99334 | 7.3614 | 7.3616 | 0.29217 | 0.021445 | 7.2755 | 7.242 | 7.5189 | 7.5671 | 0.24342 | 0.32517 | 0.062082 | 0.06323 | 116.0204 | 187.7241 | 0.014772 | 0.00069248 | FALSE | 0.99998 | TRUE | [M-H]- | 107.6583 | 7.2999 | 7.4315 | 0.13165 | 7.3076 | 7.4222 | 0.1146 | -1431.4473 | 1.1007 | 1.1979 | 12.8653 | 0 | 29 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 93814.8681 | No Control Sample | 0 | off | off | 93814.8681 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9377 | 2.6627 | 1.0623 | 79.9569 | -3.7841 | -0.30256 | 0 | 0 | 0 | 0 | 0 | 0.98867 | FALSE | 2.0275 | 1.6309
4405 | 4406 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | Br-PFHxS | 1.62019851828491 | Qualifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 1.89 | 4 | 0.4725 | 0 | 7.16 | 7.35 | Manual | Baseline | 23652.4033 | 57522.7518 | 0.41118 | 4502.5893 | 14089.0354 | 0.31958 | 5.2531 | 4.0828 | 23843.0342 | 14093.8918 | 0.19458 | 0.99334 | 7.1775 | 7.3616 | 0.24485 | 0.017531 | 7.0796 | 7.242 | 7.2743 | 7.5671 | 0.19462 | 0.32517 | 0.083736 | 0.06323 | 22.1468 | 187.7241 | 0 | 0.00069248 | TRUE | 0.975 | TRUE | [M-H]- | 85.7248 | 7.0934 | 7.2755 | 0.18205 | 7.1011 | 7.2599 | 0.15878 | 0 | 1.0159 | 0.93886 | 0.25212 | 0 | 23 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 23652.4033 | No Control Sample | 0 | off | off | 93814.8681 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9377 | 2.6627 | 1.0623 | 79.9518 | -68.0301 | -5.4395 | 0 | 0 | 0 | 0 | 0 | 0.98867 | FALSE | 0.4725 | 0.41118
4407 | 4408 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFHpS | 12.0789154404374 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7.98 | 7.35 | Baseline | Baseline | 105746.202 | 57522.7518 | 1.8383 | 24658.2791 | 14089.0354 | 1.7502 | 4.2885 | 4.0828 | 24675.4412 | 14093.8918 | 0.99342 | 0.99334 | 7.9964 | 7.3616 | 0.20546 | 0.016395 | 7.8786 | 7.242 | 8.1911 | 7.5671 | 0.31245 | 0.32517 | 0.066884 | 0.06323 | 132.1444 | 187.7241 | 0.00046728 | 0.00069248 | FALSE | 1.0862 | TRUE | [M-H]- | 120.7892 | 7.933 | 8.0758 | 0.14274 | 7.9416 | 8.0638 | 0.1222 | 36.8776 | 1.0759 | 1.116 | 11.3987 | 0 | 23 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 105746.202 | No Control Sample | 0 | off | off | 105746.202 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9346 | 2.7495 | 1.2343 | 79.957 | -2.6407 | -0.21114 | 0 | 0 | 0 | 0 | 0 | 1.0079 | FALSE | 2.5 | 1.8383
4409 | 4410 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | L-PFOS | 7.70420385077814 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 7.88 | 4 | 1.97 | 0 | 8.46 | 8.45 | Manual | Baseline | 75671.6915 | 37787.6906 | 2.0025 | 15686.1133 | 7874.098 | 1.9921 | 4.8241 | 4.799 | 15891.1681 | 7875.096 | 0.67249 | 0.99489 | 8.4817 | 8.4831 | 0.25592 | 0.021651 | 8.4035 | 8.2934 | 8.5809 | 8.7106 | 0.1774 | 0.41718 | 0.076801 | 0.075172 | 48.8415 | 151.6823 | 0.014136 | 0.00027461 | TRUE | 0.99982 | TRUE | [M-H]- | 97.7691 | 8.3933 | 8.5592 | 0.16583 | 8.3933 | 8.5489 | 0.15559 | -1249.901 | 1.003 | 0.88216 | 14.1875 | 0 | 11 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 75671.6915 | No Control Sample | 0 | off | off | 75671.6915 | 0 | off | 0 | off | 0 | off | 0 | off | 498.931 | 1.6322 | 0.81437 | 79.9567 | -7.049 | -0.56362 | 0 | 0 | 0 | 0 | 0 | 1.0937 | FALSE | 1.97 | 2.0025
4411 | 4412 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | Br-PFOS | 1.50669866018931 | Qualifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 2.11 | 4 | 0.5275 | 0 | 8.3 | 8.45 | Manual | Baseline | 31682.1113 | 37787.6906 | 0.83842 | 5312.7906 | 7874.098 | 0.67472 | 5.9634 | 4.799 | 15891.1681 | 7875.096 | 0.28167 | 0.99489 | 8.3348 | 8.4831 | 0.41893 | 0.034771 | 8.216 | 8.2934 | 8.4077 | 8.7106 | 0.1917 | 0.41718 | 0.099077 | 0.075172 | 16.5423 | 151.6823 | 0.011556 | 0.00027461 | TRUE | 0.98251 | TRUE | [M-H]- | 71.4075 | 8.247 | 8.4122 | 0.16517 | 8.2569 | 8.4122 | 0.15527 | 320.277 | 0.91538 | 0.93316 | 0.41868 | 0 | 13 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 31682.1113 | No Control Sample | 0 | off | off | 75671.6915 | 0 | off | 0 | off | 0 | off | 0 | off | 498.931 | 1.6322 | 0.81437 | 79.962 | 58.799 | 4.7014 | 0 | 0 | 0 | 0 | 0 | 1.0937 | FALSE | 0.5275 | 0.83842
4413 | 4414 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFNS | 10.9377622593607 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.85 | 8.45 | Baseline | Baseline | 91609.1723 | 37787.6906 | 2.4243 | 18978.9048 | 7874.098 | 2.4103 | 4.8269 | 4.799 | 19006.5851 | 7875.096 | 0.9886 | 0.99489 | 8.8814 | 8.4831 | 0.35467 | 0.031388 | 8.7415 | 8.2934 | 9.0119 | 8.7106 | 0.27038 | 0.41718 | 0.075374 | 0.075172 | 118.7719 | 151.6823 | 0.0014905 | 0.00027461 | FALSE | 1.0469 | TRUE | [M-H]- | 109.3776 | 8.803 | 8.9643 | 0.16136 | 8.8103 | 8.9528 | 0.14253 | 104.6228 | 0.98423 | 0.90958 | 11.0168 | 0 | 14 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 91609.1723 | No Control Sample | 0 | off | off | 91609.1723 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9276 | 0.99513 | 0.54625 | 79.9566 | -5.6557 | -0.45222 | 0 | 0 | 0 | 0 | 0 | 1.552 | FALSE | 2.5 | 2.4243
4415 | 4416 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | PFDS | 10.5932475088292 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.19 | 8.45 | Baseline | Baseline | 91159.8705 | 37787.6906 | 2.4124 | 20187.8011 | 7874.098 | 2.5638 | 4.5156 | 4.799 | 20194.6451 | 7875.096 | 0.99374 | 0.99489 | 9.206 | 8.4831 | 0.17434 | 0.016022 | 9.0688 | 8.2934 | 9.4085 | 8.7106 | 0.33965 | 0.41718 | 0.070562 | 0.075172 | 265.1976 | 151.6823 | 0.00082677 | 0.00027461 | FALSE | 1.0852 | TRUE | [M-H]- | 105.9325 | 9.1328 | 9.2833 | 0.15048 | 9.1436 | 9.2756 | 0.13206 | 49.1409 | 1.0474 | 1.161 | 11.9787 | 0 | 19 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 91159.8705 | No Control Sample | 0 | off | off | 91159.8705 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9242 | 0.6955 | 0.41655 | 79.9568 | -0.82291 | -0.065798 | 0 | 0 | 0 | 0 | 0 | 1.5158 | FALSE | 2.5 | 2.4124
4419 | 4420 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FBSA | 8.99051866073011 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Total Width | 0 | 0 | 10 | 4 | 2.5 | 0 | 4.6 | 8.86 | Baseline | Baseline | 155281.151 | 65502.9694 | 2.3706 | 14208.5545 | 14196.1839 | 1.0009 | 10.9287 | 4.6141 | 14540.6711 | 14200.5271 | 0.93519 | 0.99697 | 4.9689 | 9.0213 | 8.019 | 0.36887 | 4.6689 | 8.8099 | 5.18 | 9.2103 | 0.5111 | 0.40037 | 0.17175 | 0.07176 | 345.8938 | 199.408 | 0.035328 | 0.00057553 | FALSE | 0.55079 | TRUE | [M-H]- | 89.9052 | 4.7892 | 5.154 | 0.36479 | 4.8099 | 5.1281 | 0.31822 | 982.1075 | 1.0122 | 0.99477 | 5.9444 | 0 | 91 | 32 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 155281.151 | No Control Sample | 0 | off | off | 155281.151 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9602 | 4.1308 | 1.2308 | 77.965 | -4.3376 | -0.33818 | 0 | 0 | 0 | 0 | 0 | 1.3175 | FALSE | 2.5 | 2.3706
4421 | 4422 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FPeSA | 11.7884498595943 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7 | 8.86 | Baseline | Baseline | 316181.9179 | 65502.9694 | 4.827 | 65530.2391 | 14196.1839 | 4.616 | 4.825 | 4.6141 | 65592.3204 | 14200.5271 | 0.99213 | 0.99697 | 7.1128 | 9.0213 | 1.6115 | 0.11281 | 6.953 | 8.8099 | 7.2914 | 9.2103 | 0.33843 | 0.40037 | 0.074408 | 0.07176 | 224.2641 | 199.408 | 0.0020556 | 0.00057553 | FALSE | 0.78845 | TRUE | [M-H]- | 117.8845 | 7.034 | 7.195 | 0.16091 | 7.0442 | 7.1839 | 0.13963 | 398.0307 | 1.0938 | 1.2038 | 3.1106 | 0 | 37 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 316181.9179 | No Control Sample | 0 | off | off | 316181.9179 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9566 | 2.226 | 0.77455 | 77.9645 | -9.6971 | -0.75603 | 0 | 0 | 0 | 0 | 0 | 1.5579 | FALSE | 2.5 | 4.827
4423 | 4424 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FHxSA | 12.2110975651518 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 7.86 | 8.86 | Baseline | Baseline | 163438.5186 | 65502.9694 | 2.4951 | 35935.9992 | 14196.1839 | 2.5314 | 4.548 | 4.6141 | 35954.6696 | 14200.5271 | 0.99457 | 0.99697 | 8.0013 | 9.0213 | 1.7973 | 0.14126 | 7.8232 | 8.8099 | 8.2052 | 9.2103 | 0.38202 | 0.40037 | 0.070779 | 0.07176 | 166.0385 | 199.408 | 0.0010055 | 0.00057553 | FALSE | 0.88693 | TRUE | [M-H]- | 122.111 | 7.933 | 8.0841 | 0.15114 | 7.942 | 8.0733 | 0.13129 | 94.5909 | 1.185 | 1.3968 | 5.7008 | 0 | 28 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 163438.5186 | No Control Sample | 0 | off | off | 163438.5186 | 0 | off | 0 | off | 0 | off | 0 | off | 397.954 | 3.567 | 1.4195 | 77.9648 | -8.0593 | -0.62834 | 0 | 0 | 0 | 0 | 0 | 0.69328 | FALSE | 2.5 | 2.4951
4425 | 4426 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FHpSA | 11.3384308978904 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.58 | 8.86 | Baseline | Baseline | 314333.1911 | 65502.9694 | 4.7988 | 71075.5636 | 14196.1839 | 5.0067 | 4.4225 | 4.6141 | 71100.7651 | 14200.5271 | 0.99572 | 0.99697 | 8.5837 | 9.0213 | 0.043081 | 0.0036963 | 8.3355 | 8.8099 | 8.7991 | 9.2103 | 0.46358 | 0.40037 | 0.068727 | 0.07176 | 118.732 | 199.408 | 0.00063839 | 0.00057553 | FALSE | 0.95149 | TRUE | [M-H]- | 113.3843 | 8.5119 | 8.6603 | 0.14838 | 8.5197 | 8.6481 | 0.12847 | 97.8762 | 1.0604 | 1.0663 | 2.7328 | 0 | 26 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 314333.1911 | No Control Sample | 0 | off | off | 314333.1911 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9501 | 1.6407 | 0.73497 | 77.9647 | -9.4164 | -0.73416 | 0 | 0 | 0 | 0 | 0 | 1.2088 | FALSE | 2.5 | 4.7988
4427 | 4428 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | FOSA | 11.8812408288588 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.88 | 8.86 | Baseline | Baseline | 177079.4884 | 65502.9694 | 2.7034 | 36958.6136 | 14196.1839 | 2.6034 | 4.7913 | 4.6141 | 36971.5799 | 14200.5271 | 0.99233 | 0.99697 | 9.0234 | 9.0213 | 1.6148 | 0.14339 | 8.7999 | 8.8099 | 9.2423 | 9.2103 | 0.44235 | 0.40037 | 0.074713 | 0.07176 | 114.1778 | 199.408 | 0.00062152 | 0.00057553 | FALSE | 1.0002 | TRUE | [M-H]- | 118.8124 | 8.945 | 9.1013 | 0.15639 | 8.9516 | 9.0915 | 0.13991 | 51.9283 | 0.93275 | 0.81176 | 5.9526 | 0 | 23 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 177079.4884 | No Control Sample | 0 | off | off | 177079.4884 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9467 | 0.99591 | 0.49591 | 77.9648 | -7.2196 | -0.56288 | 0 | 0 | 0 | 0 | 0 | 1.5616 | FALSE | 2.5 | 2.7034
4446 | 4447 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 4:2 FTS | 10.3189000269764 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 5.94 | 5.92 | Baseline | Baseline | 87082.8359 | 16620.7224 | 5.2394 | 14104.203 | 2729.8472 | 5.1667 | 6.1742 | 6.0885 | 14142.493 | 2749.0687 | 0.98779 | 0.97476 | 5.9531 | 5.9481 | 0.22003 | 0.01307 | 5.7852 | 5.8011 | 6.1055 | 6.0799 | 0.32022 | 0.27875 | 0.095516 | 0.096704 | 178.2887 | 151.8747 | 0.0024616 | 0.008449 | FALSE | 1.0008 | TRUE | [M-H]- | 103.189 | 5.8503 | 6.0554 | 0.20512 | 5.8628 | 6.0396 | 0.1768 | 108.4249 | 1.0259 | 1.0208 | 5.0646 | 0 | 38 | 13 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 87082.8359 | No Control Sample | 0 | off | off | 87082.8359 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9756 | 4.0772 | 1.3331 | 306.9689 | 4.1248 | 1.2662 | 0 | 0 | 0 | 0 | 0 | 5.3291 | FALSE | 2.5 | 5.2394
4448 | 4449 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 6:2 FTS | 10.4905944630528 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 7.92 | 7.92 | Baseline | Baseline | 87387.3963 | 13729.5463 | 6.3649 | 20774.5247 | 3495.9877 | 5.9424 | 4.2065 | 3.9272 | 20774.5607 | 3498.5645 | 0.99846 | 0.9949 | 7.9345 | 7.9318 | 0.1837 | 0.014549 | 7.7577 | 7.8178 | 8.2246 | 8.1034 | 0.46685 | 0.28553 | 0.063822 | 0.058085 | 115.3343 | 112.9938 | 4.3464e-06 | 0.0016869 | FALSE | 1.0004 | TRUE | [M-H]- | 104.9059 | 7.8666 | 8.01 | 0.14344 | 7.8749 | 7.9966 | 0.12171 | 0.19341 | 1.1341 | 1.2163 | 3.6344 | 0 | 34 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 87387.3963 | No Control Sample | 0 | off | off | 87387.3963 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9701 | 5.0277 | 2.1467 | 406.9635 | 4.8232 | 1.9629 | 0 | 0 | 0 | 0 | 0 | 5.5018 | FALSE | 2.5 | 6.3649
4450 | 4451 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 8:2 FTS | 11.4804278220497 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 8.88 | 8.89 | Baseline | Baseline | 78631.1159 | 13959.2719 | 5.6329 | 15319.8658 | 2731.7814 | 5.608 | 5.1326 | 5.11 | 15321.0891 | 2731.7859 | 0.99881 | 0.99396 | 8.892 | 8.8928 | 0.13496 | 0.011985 | 8.705 | 8.7132 | 9.0697 | 9.0592 | 0.36467 | 0.34597 | 0.080654 | 0.079424 | 148.7949 | 110.134 | 0.00015924 | 3.0606e-06 | FALSE | 0.9999 | TRUE | [M-H]- | 114.8043 | 8.8089 | 8.9816 | 0.17263 | 8.8202 | 8.9676 | 0.14735 | 6.6896 | 1.1106 | 1.2176 | 2.7384 | 0 | 19 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 78631.1159 | No Control Sample | 0 | off | off | 78631.1159 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9636 | 3.9853 | 2.1001 | 506.9571 | 2.7 | 1.3688 | 0 | 0 | 0 | 0 | 0 | 5.7212 | FALSE | 2.5 | 5.6329
4452 | 4453 | PFAS CS7 10ng/ml | 32 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 33) | CS7 | 0 | 0 | Standard | 0 | 0 | 0 | 8 | 1 | 10:2 FTS | 10.5650410837034 | Quantifiers | 10/27/2023 03:42:18 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 10 | 4 | 2.5 | 0 | 9.5 | 8.89 | Baseline | Baseline | 75329.8316 | 13959.2719 | 5.3964 | 18419.1748 | 2731.7814 | 6.7426 | 4.0898 | 5.11 | 18419.5145 | 2731.7859 | 0.99717 | 0.99396 | 9.5097 | 8.8928 | 0.10168 | 0.0096593 | 9.3761 | 8.7132 | 9.651 | 9.0592 | 0.27488 | 0.34597 | 0.064195 | 0.079424 | 388.8939 | 110.134 | 3.5669e-05 | 3.0606e-06 | FALSE | 1.0694 | TRUE | [M-H]- | 105.6504 | 9.4421 | 9.5767 | 0.13464 | 9.4485 | 9.5677 | 0.1192 | 2.3901 | 1.05 | 1.0668 | 4.6564 | 0 | 19 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 75329.8316 | No Control Sample | 0 | off | off | 75329.8316 | 0 | off | 0 | off | 0 | off | 0 | off | 626.957 | 2.9885 | 1.8736 | 606.9511 | 5.3096 | 3.2226 | 0 | 0 | 0 | 0 | 0 | 1.2974 | FALSE | 2.5 | 5.3964
4513 | 4514 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFBA | 38.9008758768512 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 1.49 | 1.48 | Baseline | Baseline | 728814.7011 | 104447.0586 | 6.9778 | 99530.476 | 13754.2017 | 7.2364 | 7.3225 | 7.5938 | 99571.1977 | 13771.5844 | 0.99576 | 0.99125 | 1.4604 | 1.4617 | 1.9885 | 0.029628 | 1.2926 | 1.3091 | 1.7672 | 1.6854 | 0.47455 | 0.37633 | 0.11707 | 0.11222 | 298.7687 | 210.4811 | 0.00093485 | 0.0033454 | FALSE | 0.99909 | TRUE | [M-H]- | 97.2522 | 1.3687 | 1.6218 | 0.25311 | 1.3782 | 1.5998 | 0.22159 | 196.0714 | 1.4818 | 1.9183 | 1.4302 | 0 | 96 | 26 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 728814.7011 | No Control Sample | 0 | off | off | 728814.7011 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9803 | 5.1909 | 1.1055 | 168.9885 | -2.7113 | -0.45818 | 0 | 0 | 0 | 0 | 0 | 0.17775 | FALSE | 10 | 6.9778
4515 | 4516 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFPeA | 37.8119772381841 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 3.18 | 3.21 | Baseline | Baseline | 644784.7453 | 76249.5587 | 8.4562 | 66899.983 | 8408.2387 | 7.9565 | 9.638 | 9.0684 | 67055.4091 | 8515.3806 | 0.98788 | 0.95662 | 3.2167 | 3.2192 | 1.1529 | 0.036663 | 2.9726 | 3.0398 | 3.528 | 3.4904 | 0.55542 | 0.45053 | 0.14569 | 0.14224 | 1000.1252 | 643.9141 | 0.003921 | 0.0094855 | FALSE | 0.99921 | TRUE | [M-H]- | 94.5299 | 3.0677 | 3.4077 | 0.33998 | 3.0851 | 3.377 | 0.29192 | 472.283 | 1.2004 | 1.3495 | 1.7229 | 0 | 97 | 28 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 644784.7453 | No Control Sample | 0 | off | off | 644784.7453 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9772 | 4.6644 | 1.2266 | 218.9855 | -1.3462 | -0.29479 | 0 | 0 | 0 | 0 | 0 | 0.16435 | FALSE | 10 | 8.4562
4517 | 4518 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFHxA | 39.8335072706634 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 642.0117 | 6.13 | 6.12 | Baseline | Baseline | 1308699.6648 | 141170.3218 | 9.2704 | 212960.847 | 24773.6619 | 8.5963 | 6.1453 | 5.6984 | 213028.1517 | 24808.1141 | 0.99638 | 0.99113 | 6.1573 | 6.1555 | 0.44493 | 0.027274 | 5.9492 | 5.9849 | 6.4383 | 6.331 | 0.48903 | 0.3461 | 0.096422 | 0.087528 | 261.138 | 277.136 | 0.00075094 | 0.002166 | FALSE | 1.0003 | TRUE | [M-H]- | 99.5838 | 6.0618 | 6.2548 | 0.19298 | 6.0747 | 6.2415 | 0.16686 | 327.0128 | 1.1252 | 1.2887 | 1.1265 | 0 | 54 | 13 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1308699.6648 | No Control Sample | 203843.5852 | off | off | 1308699.6648 | 203843.5852 | off | 0.025 | off | 40 | off | 256.8047 | off | 312.9736 | 2.4911 | 0.77965 | 268.9817 | -2.7711 | -0.74537 | 0 | 0 | 0 | 0 | 0 | 3.9085 | FALSE | 10 | 9.2704
4519 | 4520 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFHpA | 36.3134613798335 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 573.987 | 7.26 | 7.27 | Baseline | Baseline | 1310577.5175 | 144233.6267 | 9.0865 | 257367.0344 | 32799.7631 | 7.8466 | 5.0923 | 4.3974 | 257400.6559 | 32825.6794 | 0.99813 | 0.99447 | 7.2914 | 7.2904 | 0.43218 | 0.031376 | 7.1426 | 7.1635 | 7.6166 | 7.4198 | 0.47397 | 0.25628 | 0.08116 | 0.068347 | 171.8478 | 249.8061 | 0.00038521 | 0.0015976 | FALSE | 1.0001 | TRUE | [M-H]- | 90.7837 | 7.218 | 7.3685 | 0.15051 | 7.2262 | 7.3593 | 0.13318 | 209.1716 | 1.1126 | 1.2404 | 1.1645 | 0 | 51 | 12 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Pass | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1310577.5175 | No Control Sample | 228328.7703 | off | off | 1310577.5175 | 228328.7703 | off | 0.025 | off | 40 | off | 229.5948 | off | 362.9704 | 2.1687 | 0.78716 | 318.9783 | -2.4215 | -0.77242 | 0 | 0 | 0 | 0 | 0 | 1.262 | FALSE | 10 | 9.0865
4521 | 4522 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFOA | 38.334971378076 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 567.0275 | 7.96 | 7.94 | Baseline | Baseline | 1085718.1504 | 116781.2815 | 9.297 | 231122.6224 | 26917.7276 | 8.5863 | 4.6976 | 4.3385 | 231232.9627 | 26928.3874 | 0.99597 | 0.99578 | 7.9737 | 7.9731 | 0.17201 | 0.013692 | 7.8469 | 7.8262 | 8.322 | 8.169 | 0.47505 | 0.34283 | 0.073146 | 0.065479 | 127.8742 | 123.7871 | 8.1815e-05 | 0.00088972 | FALSE | 1.0001 | TRUE | [M-H]- | 95.8374 | 7.9046 | 8.058 | 0.15341 | 7.9134 | 8.0466 | 0.13328 | -39.8051 | 1.0212 | 1.0083 | 0.7282 | 0 | 34 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1085718.1504 | No Control Sample | 191475.4056 | off | off | 1085718.1504 | 191475.4056 | off | 0.025 | off | 40 | off | 226.811 | off | 412.9678 | 3.2727 | 1.3515 | 368.9753 | -2.2832 | -0.84245 | 0 | 0 | 0 | 0 | 0 | 1.7506 | FALSE | 10 | 9.297
4523 | 4524 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFNA | 37.7750574779428 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 581.9157 | 8.46 | 8.47 | Baseline | Baseline | 926080.3713 | 105531.0428 | 8.7754 | 172479.9381 | 21226.7749 | 8.1256 | 5.3692 | 4.9716 | 172513.7002 | 21261.3907 | 0.99772 | 0.99129 | 8.4878 | 8.4853 | 0.32823 | 0.027768 | 8.3229 | 8.3138 | 8.7375 | 8.6229 | 0.41465 | 0.30915 | 0.084368 | 0.077969 | 142.6443 | 150.1672 | 0.00046481 | 0.0027522 | FALSE | 1.0003 | TRUE | [M-H]- | 94.4376 | 8.398 | 8.5665 | 0.16847 | 8.407 | 8.5582 | 0.15122 | 193.3458 | 1.0285 | 1.0744 | 0.7118 | 0 | 24 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 926080.3713 | No Control Sample | 159143.3885 | off | off | 926080.3713 | 159143.3885 | off | 0.025 | off | 40 | off | 232.7663 | off | 462.9645 | 2.6501 | 1.2269 | 418.9728 | -1.072 | -0.44912 | 0 | 0 | 0 | 0 | 0 | 1.1682 | FALSE | 10 | 8.7754
4525 | 4526 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFDA | 35.7349347419302 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.87 | 8.86 | Baseline | Baseline | 755517.6861 | 85167.1083 | 8.871 | 134951.0109 | 16685.7497 | 8.0878 | 5.5985 | 5.1042 | 134970.738 | 16686.2827 | 0.99842 | 0.9986 | 8.8936 | 8.8954 | 0.26621 | 0.023613 | 8.7198 | 8.7099 | 9.1969 | 9.1073 | 0.47717 | 0.39735 | 0.088944 | 0.079854 | 120.2236 | 119.4445 | 0.00024288 | 7.0182e-05 | FALSE | 0.9998 | TRUE | [M-H]- | 89.3373 | 8.8091 | 8.9847 | 0.17569 | 8.8204 | 8.9763 | 0.15585 | 68.6903 | 0.95945 | 0.94404 | 0.66933 | 0 | 25 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 755517.6861 | No Control Sample | 0 | off | off | 755517.6861 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9623 | 4.4926 | 2.3045 | 468.9699 | -0.10735 | -0.050342 | 0 | 0 | 0 | 0 | 0 | 3.8973 | FALSE | 10 | 8.871
4527 | 4528 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFUdA | 39.6699627885479 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.21 | 9.2 | Baseline | Baseline | 635666.9058 | 72255.6146 | 8.7975 | 129649.0192 | 16250.1624 | 7.9783 | 4.903 | 4.4465 | 129664.0196 | 16254.0087 | 0.99794 | 0.9974 | 9.2323 | 9.2291 | 0.24246 | 0.022331 | 9.0813 | 9.0895 | 9.4977 | 9.3735 | 0.41647 | 0.28398 | 0.077445 | 0.068766 | 241.2178 | 150.1329 | 0.00028058 | 0.00046259 | FALSE | 1.0003 | TRUE | [M-H]- | 99.1749 | 9.1538 | 9.3135 | 0.15976 | 9.1635 | 9.3048 | 0.14132 | 87.348 | 0.96206 | 0.92676 | 0.6914 | 0 | 24 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 635666.9058 | No Control Sample | 0 | off | off | 635666.9058 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9589 | 3.6593 | 2.06 | 518.9667 | -1.1207 | -0.5816 | 0 | 0 | 0 | 0 | 0 | 3.1113 | FALSE | 10 | 8.7975
4529 | 4530 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFDoA | 37.7502730475106 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.47 | 9.47 | Baseline | Baseline | 662615.3572 | 67828.7725 | 9.7689 | 156688.2196 | 16320.4289 | 9.6007 | 4.2289 | 4.1561 | 156761.8526 | 16337.9883 | 0.99615 | 0.9904 | 9.4971 | 9.4966 | 0.28634 | 0.027116 | 9.3679 | 9.3576 | 9.6689 | 9.617 | 0.30102 | 0.25932 | 0.067433 | 0.06498 | 335.4461 | 248.4206 | 0.00089236 | 0.0018943 | FALSE | 1.0001 | TRUE | [M-H]- | 94.3757 | 9.4284 | 9.5657 | 0.13727 | 9.4365 | 9.5551 | 0.11855 | 464.4995 | 0.97742 | 0.90918 | 1.0814 | 0 | 21 | 7 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 662615.3572 | No Control Sample | 0 | off | off | 662615.3572 | 0 | off | 0 | off | 0 | off | 0 | off | 612.954 | 0.55353 | 0.33929 | 568.9635 | 0.19173 | 0.10909 | 0 | 0 | 0 | 0 | 0 | 8.421 | FALSE | 10 | 9.7689
4531 | 4532 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFTrDA | 40.637829247881 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.68 | 9.47 | Baseline | Baseline | 519944.7076 | 67828.7725 | 7.6655 | 134400.9304 | 16320.4289 | 8.2351 | 3.8686 | 4.1561 | 134482.7202 | 16337.9883 | 0.99497 | 0.9904 | 9.7111 | 9.4966 | 0.32085 | 0.031058 | 9.556 | 9.3576 | 9.8318 | 9.617 | 0.2758 | 0.25932 | 0.061599 | 0.06498 | 737.8393 | 248.4206 | 0.0010312 | 0.0018943 | FALSE | 1.0226 | TRUE | [M-H]- | 101.5946 | 9.6499 | 9.7741 | 0.12415 | 9.6585 | 9.7664 | 0.10793 | 502.5409 | 1.0949 | 1.2427 | 0.78046 | 0 | 24 | 8 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 519944.7076 | No Control Sample | 0 | off | off | 519944.7076 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9515 | 1.548 | 1.0262 | 618.9617 | 0.0078854 | 0.0048807 | 0 | 0 | 0 | 0 | 0 | 3.9162 | FALSE | 10 | 7.6655
4533 | 4534 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFTeDA | 34.9997835585167 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 9.87 | 9.86 | Baseline | Baseline | 227941.159 | 18368.9268 | 12.4091 | 60302.5453 | 5011.6259 | 12.0325 | 3.78 | 3.6653 | 60423.1841 | 5012.1298 | 0.98918 | 0.99725 | 9.8902 | 9.8897 | 0.20505 | 0.020239 | 9.7925 | 9.7946 | 9.984 | 10.0081 | 0.19152 | 0.21352 | 0.059727 | 0.058372 | 984.6586 | 816.3861 | 0.0032734 | 0.0002308 | FALSE | 1.0001 | TRUE | [M-H]- | 87.4995 | 9.8356 | 9.9504 | 0.11486 | 9.843 | 9.9427 | 0.099705 | 1030.7056 | 1.0096 | 1.0163 | 1.4455 | 0 | 21 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 227941.159 | No Control Sample | 0 | off | off | 227941.159 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9491 | 2.5379 | 1.8094 | 668.9591 | 1.7768 | 1.1886 | 0 | 0 | 0 | 0 | 0 | 8.5403 | FALSE | 10 | 12.4091
4540 | 4541 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFBS | 37.1050350452568 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 3.5 | 3.5 | Baseline | Manual | 416247.2299 | 42382.3045 | 9.8213 | 36785.6564 | 3814.2242 | 9.6443 | 11.3155 | 11.1116 | 36953.3668 | 3826.108 | 0.98305 | 0.98523 | 3.8699 | 3.8753 | 10.5679 | 0.36988 | 3.5748 | 3.6363 | 4.1855 | 4.1305 | 0.61072 | 0.49422 | 0.17684 | 0.16285 | 688.556 | 765.8939 | 0.0076905 | 0 | FALSE | 0.9986 | TRUE | [M-H]- | 92.7626 | 3.6952 | 4.077 | 0.38177 | 3.7134 | 4.0515 | 0.33806 | 463.2234 | 1.0328 | 1.0291 | 8.0528 | 0 | 95 | 30 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 416247.2299 | No Control Sample | 0 | off | off | 416247.2299 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9437 | 2.2916 | 0.68506 | 79.9567 | -6.6695 | -0.53328 | 0 | 0 | 0 | 0 | 0 | 1.4598 | FALSE | 10 | 9.8213
4542 | 4543 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFPeS | 39.3243964262288 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 6.38 | 3.5 | Baseline | Manual | 444181.813 | 42382.3045 | 10.4804 | 89385.2414 | 3814.2242 | 23.4347 | 4.9693 | 11.1116 | 89491.498 | 3826.108 | 0.99036 | 0.98523 | 6.4088 | 3.8753 | 0.45143 | 0.028801 | 6.2411 | 3.6363 | 6.6045 | 4.1305 | 0.3635 | 0.49422 | 0.076761 | 0.16285 | 346.8114 | 765.8939 | 0.0025878 | 0 | FALSE | 1.6538 | TRUE | [M-H]- | 98.311 | 6.3241 | 6.4925 | 0.16838 | 6.3356 | 6.4804 | 0.14485 | 636.3485 | 1.0721 | 1.1605 | 5.191 | 0 | 40 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 444181.813 | No Control Sample | 0 | off | off | 444181.813 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9405 | 1.9602 | 0.68399 | 79.9566 | -7.4855 | -0.59852 | 0 | 0 | 0 | 0 | 0 | 1.2799 | FALSE | 10 | 10.4804
4544 | 4545 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | L-PFHxS | 31.304634061445 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 32.44 | 4 | 8.11 | 0 | 7.34 | 7.35 | Baseline | Baseline | 332665.6467 | 56889.9734 | 5.8475 | 82108.2175 | 13980.6868 | 5.873 | 4.0516 | 4.0692 | 83125.4154 | 14001.6543 | 0.75132 | 0.98765 | 7.3666 | 7.3664 | 0.36195 | 0.026567 | 7.2791 | 7.2475 | 7.4932 | 7.4865 | 0.21413 | 0.23903 | 0.063803 | 0.062429 | 113.4381 | 161.5339 | 0.014768 | 0.0021829 | FALSE | 1 | TRUE | [M-H]- | 96.5001 | 7.304 | 7.4344 | 0.13037 | 7.3108 | 7.4256 | 0.11479 | -5662.7857 | 0.99341 | 0.95133 | 7.2268 | 0 | 27 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 332665.6467 | No Control Sample | 0 | off | off | 332665.6467 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9373 | 1.7924 | 0.71506 | 79.9564 | -10.5764 | -0.84566 | 0 | 0 | 0 | 0 | 0 | 0.90264 | FALSE | 8.11 | 5.8475
4546 | 4547 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | Br-PFHxS | 6.37405829607918 | Qualifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 7.56 | 4 | 1.89 | 0 | 7.16 | 7.35 | Manual | Baseline | 92027.8241 | 56889.9734 | 1.6176 | 17373.1788 | 13980.6868 | 1.2427 | 5.2971 | 4.0692 | 83125.4154 | 14001.6543 | 0.20793 | 0.98765 | 7.183 | 7.3664 | 0.32184 | 0.023044 | 7.0848 | 7.2475 | 7.2797 | 7.4865 | 0.19492 | 0.23903 | 0.084395 | 0.062429 | 24.0022 | 161.5339 | 0.0091825 | 0.0021829 | TRUE | 0.97511 | TRUE | [M-H]- | 84.3129 | 7.1001 | 7.2873 | 0.18721 | 7.1085 | 7.2663 | 0.1578 | 818.4562 | 1.1598 | 1.1801 | 0.27664 | 0 | 24 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 92027.8241 | No Control Sample | 0 | off | off | 332665.6467 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9373 | 1.7924 | 0.71506 | 79.9751 | 223.6048 | 17.8788 | 0 | 0 | 0 | 0 | 0 | 0.90264 | FALSE | 1.89 | 1.6176
4548 | 4549 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFHpS | 38.209643477797 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 7.98 | 7.35 | Baseline | Baseline | 330830.7732 | 56889.9734 | 5.8153 | 74619.018 | 13980.6868 | 5.3373 | 4.4336 | 4.0692 | 74654.6068 | 14001.6543 | 0.99617 | 0.98765 | 7.9981 | 7.3664 | 0.22619 | 0.01805 | 7.8765 | 7.2475 | 8.3805 | 7.4865 | 0.50408 | 0.23903 | 0.068761 | 0.062429 | 97.9656 | 161.5339 | 0.00052239 | 0.0021829 | FALSE | 1.0857 | TRUE | [M-H]- | 95.5241 | 7.9318 | 8.0794 | 0.14752 | 7.9405 | 8.0683 | 0.12778 | -77.3285 | 1.1931 | 1.4044 | 4.9232 | 0 | 35 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 330830.7732 | No Control Sample | 0 | off | off | 330830.7732 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9339 | 1.0926 | 0.49052 | 79.9565 | -8.9863 | -0.71852 | 0 | 0 | 0 | 0 | 0 | 6.5537 | FALSE | 10 | 5.8153
4550 | 4551 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | L-PFOS | 30.90326843373 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 31.52 | 4 | 7.88 | 0 | 8.46 | 8.45 | Manual | Baseline | 255386.9674 | 31793.5476 | 8.0327 | 51770.3935 | 6252.8255 | 8.2795 | 4.9331 | 5.0847 | 52295.9812 | 6258.5066 | 0.64087 | 0.98747 | 8.483 | 8.4823 | 0.27226 | 0.023033 | 8.4004 | 8.3341 | 8.5829 | 8.6417 | 0.18248 | 0.3076 | 0.076899 | 0.080147 | 58.1752 | 161.8547 | 0 | 0.0015593 | TRUE | 1.0001 | TRUE | [M-H]- | 98.0434 | 8.3999 | 8.5598 | 0.15997 | 8.3999 | 8.5504 | 0.15051 | 0 | 0.94785 | 0.78367 | 7.9392 | 0 | 11 | 5 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 255386.9674 | No Control Sample | 0 | off | off | 255386.9674 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9307 | 0.96032 | 0.47913 | 79.9566 | -8.7604 | -0.70046 | 0 | 0 | 0 | 0 | 0 | 1.3646 | FALSE | 7.88 | 8.0327
4552 | 4553 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | Br-PFOS | 7.11287274917257 | Qualifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 8.44 | 4 | 2.11 | 0 | 8.3 | 8.45 | Manual | Baseline | 125840.7774 | 31793.5476 | 3.9581 | 21978.5377 | 6252.8255 | 3.515 | 5.7256 | 5.0847 | 52295.9812 | 6258.5066 | 0.31592 | 0.98747 | 8.3342 | 8.4823 | 0.41258 | 0.034244 | 8.2231 | 8.3341 | 8.4013 | 8.6417 | 0.17822 | 0.3076 | 0.095325 | 0.080147 | 24.6976 | 161.8547 | 0 | 0.0015593 | TRUE | 0.98255 | TRUE | [M-H]- | 84.2757 | 8.2456 | 8.4212 | 0.17551 | 8.2534 | 8.4212 | 0.1678 | 0 | 0.88512 | 0.83526 | 0.49275 | 0 | 13 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 125840.7774 | No Control Sample | 0 | off | off | 255386.9674 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9307 | 0.96032 | 0.47913 | 79.9587 | 17.0208 | 1.3609 | 0 | 0 | 0 | 0 | 0 | 1.3646 | FALSE | 2.11 | 3.9581
4554 | 4555 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFNS | 41.8368010247848 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.85 | 8.45 | Baseline | Baseline | 294820.4398 | 31793.5476 | 9.273 | 59691.0387 | 6252.8255 | 9.5463 | 4.9391 | 5.0847 | 59731.8239 | 6258.5066 | 0.99303 | 0.98747 | 8.8817 | 8.4823 | 0.35873 | 0.031748 | 8.7427 | 8.3341 | 9.123 | 8.6417 | 0.3803 | 0.3076 | 0.07722 | 0.080147 | 134.3773 | 161.8547 | 0.00040458 | 0.0015593 | FALSE | 1.0471 | TRUE | [M-H]- | 104.592 | 8.8027 | 8.9653 | 0.16257 | 8.8097 | 8.9549 | 0.14528 | -63.5014 | 0.99311 | 0.93922 | 7.4214 | 0 | 20 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 294820.4398 | No Control Sample | 0 | off | off | 294820.4398 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9276 | 0.97697 | 0.53628 | 79.9567 | -5.5659 | -0.44503 | 0 | 0 | 0 | 0 | 0 | 1.6275 | FALSE | 10 | 9.273
4556 | 4557 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFDS | 41.9982122004606 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.19 | 8.45 | Baseline | Baseline | 304084.3183 | 31793.5476 | 9.5643 | 66138.085 | 6252.8255 | 10.5773 | 4.5977 | 5.0847 | 66157.4818 | 6258.5066 | 0.99366 | 0.98747 | 9.207 | 8.4823 | 0.18502 | 0.017003 | 9.03 | 8.3341 | 9.3882 | 8.6417 | 0.35818 | 0.3076 | 0.072248 | 0.080147 | 251.0307 | 161.8547 | 0.00057663 | 0.0015593 | FALSE | 1.0854 | TRUE | [M-H]- | 104.9955 | 9.1333 | 9.2884 | 0.15514 | 9.1447 | 9.277 | 0.13229 | 106.474 | 1.1262 | 1.3027 | 8.0716 | 0 | 20 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 304084.3183 | No Control Sample | 0 | off | off | 304084.3183 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9244 | 0.90175 | 0.54008 | 79.9565 | -4.5023 | -0.35999 | 0 | 0 | 0 | 0 | 0 | 0.94485 | FALSE | 10 | 9.5643
4560 | 4561 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FBSA | 32.8685440081386 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy, Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 4.6 | 8.86 | Manual | Baseline | 586280.9048 | 67647.5818 | 8.6667 | 51926.2645 | 14471.3414 | 3.5882 | 11.2906 | 4.6746 | 52188.5184 | 14473.4809 | 0.97942 | 0.9968 | 4.9863 | 9.0227 | 8.3984 | 0.38632 | 4.7295 | 8.8724 | 5.2394 | 9.2273 | 0.50989 | 0.35487 | 0.17801 | 0.072737 | 454.1529 | 129.507 | 0 | 0.00035723 | TRUE | 0.55264 | TRUE | [M-H]- | 82.1714 | 4.7862 | 5.1879 | 0.40173 | 4.8199 | 5.1597 | 0.33974 | 0 | 1.0198 | 1.0817 | 4.5 | 0 | 93 | 34 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 586280.9048 | No Control Sample | 0 | off | off | 586280.9048 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9597 | 2.5033 | 0.74589 | 77.9649 | -5.6759 | -0.44253 | 0 | 0 | 0 | 0 | 0 | 0.99162 | FALSE | 10 | 8.6667
4562 | 4563 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FPeSA | 41.6942376513238 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7 | 8.86 | Baseline | Baseline | 1154908.7045 | 67647.5818 | 17.0724 | 238360.0113 | 14471.3414 | 16.4712 | 4.8452 | 4.6746 | 238508.2478 | 14473.4809 | 0.99409 | 0.9968 | 7.1186 | 9.0227 | 1.694 | 0.11858 | 6.9525 | 8.8724 | 7.3195 | 9.2273 | 0.36693 | 0.35487 | 0.075725 | 0.072737 | 256.7054 | 129.507 | 0.0013334 | 0.00035723 | FALSE | 0.78897 | TRUE | [M-H]- | 104.2356 | 7.0402 | 7.2016 | 0.16136 | 7.051 | 7.1907 | 0.13977 | 866.1728 | 1.0627 | 1.1453 | 1.7946 | 0 | 41 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1154908.7045 | No Control Sample | 0 | off | off | 1154908.7045 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9565 | 2.1235 | 0.7389 | 77.9643 | -12.2423 | -0.95448 | 0 | 0 | 0 | 0 | 0 | 1.3849 | FALSE | 10 | 17.0724
4564 | 4565 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FHxSA | 38.7779481848076 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7.86 | 8.86 | Baseline | Baseline | 536013.6261 | 67647.5818 | 7.9236 | 116069.2801 | 14471.3414 | 8.0206 | 4.618 | 4.6746 | 116155.908 | 14473.4809 | 0.99407 | 0.9968 | 8.0041 | 9.0227 | 1.8334 | 0.14411 | 7.8478 | 8.8724 | 8.1951 | 9.2273 | 0.34728 | 0.35487 | 0.072046 | 0.072737 | 124.9607 | 129.507 | 0.0014007 | 0.00035723 | FALSE | 0.88711 | TRUE | [M-H]- | 96.9449 | 7.9343 | 8.0857 | 0.15136 | 7.9427 | 8.0768 | 0.13407 | 468.1433 | 1.0798 | 1.1738 | 2.3612 | 0 | 26 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 536013.6261 | No Control Sample | 0 | off | off | 536013.6261 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9532 | 1.6102 | 0.6408 | 77.9644 | -13.2565 | -1.0336 | 0 | 0 | 0 | 0 | 0 | 0.60938 | FALSE | 10 | 7.9236
4566 | 4567 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FHpSA | 38.0222803560105 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.58 | 8.86 | Baseline | Baseline | 1088595.9942 | 67647.5818 | 16.0922 | 237780.8649 | 14471.3414 | 16.4312 | 4.5781 | 4.6746 | 238162.2837 | 14473.4809 | 0.98984 | 0.9968 | 8.5805 | 9.0227 | 0.0060981 | 0.00052322 | 8.4617 | 8.8724 | 8.9024 | 9.2273 | 0.44065 | 0.35487 | 0.071793 | 0.072737 | 131.4428 | 129.507 | 0.0015715 | 0.00035723 | FALSE | 0.951 | TRUE | [M-H]- | 95.0557 | 8.5089 | 8.6603 | 0.15134 | 8.5193 | 8.6483 | 0.12897 | -848.0242 | 0.99172 | 0.95673 | 1.3822 | 0 | 25 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1088595.9942 | No Control Sample | 0 | off | off | 1088595.9942 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9501 | 1.4629 | 0.65531 | 77.9642 | -15.7062 | -1.2245 | 0 | 0 | 0 | 0 | 0 | 1.5324 | FALSE | 10 | 16.0922
4568 | 4569 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FOSA | 40.1994750175858 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.88 | 8.86 | Baseline | Baseline | 618754.1564 | 67647.5818 | 9.1467 | 131967.5968 | 14471.3414 | 9.1192 | 4.6887 | 4.6746 | 132012.2867 | 14473.4809 | 0.99328 | 0.9968 | 9.0246 | 9.0227 | 1.6279 | 0.14455 | 8.8211 | 8.8724 | 9.2394 | 9.2273 | 0.41827 | 0.35487 | 0.072429 | 0.072737 | 145.6658 | 129.507 | 0.000666 | 0.00035723 | FALSE | 1.0002 | TRUE | [M-H]- | 100.4987 | 8.9453 | 9.1008 | 0.15556 | 8.9527 | 9.0923 | 0.1396 | 210.1307 | 0.9473 | 0.86832 | 2.8598 | 0 | 22 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 618754.1564 | No Control Sample | 0 | off | off | 618754.1564 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9466 | 0.8824 | 0.43939 | 77.9644 | -12.8257 | -0.99996 | 0 | 0 | 0 | 0 | 0 | 0.99025 | FALSE | 10 | 9.1467
4587 | 4588 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 4:2 FTS | 38.7196034955492 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 5.94 | 5.92 | Baseline | Baseline | 315930.4195 | 16069.8299 | 19.6598 | 53459.1725 | 2748.1447 | 19.4528 | 5.9098 | 5.8475 | 53489.3915 | 2749.9738 | 0.99528 | 0.99592 | 5.9691 | 5.9692 | 0.48931 | 0.029065 | 5.7816 | 5.821 | 6.1685 | 6.1528 | 0.38687 | 0.33178 | 0.090432 | 0.088095 | 197.143 | 165.4797 | 0.00062098 | 0.0011504 | FALSE | 0.99997 | TRUE | [M-H]- | 96.799 | 5.8686 | 6.0683 | 0.1996 | 5.882 | 6.0545 | 0.1725 | 85.8096 | 0.93155 | 0.83981 | 3.2377 | 0 | 46 | 13 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 315930.4195 | No Control Sample | 0 | off | off | 315930.4195 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9754 | 3.4713 | 1.135 | 306.968 | 1.3263 | 0.40712 | 0 | 0 | 0 | 0 | 0 | 4.1263 | FALSE | 10 | 19.6598
4589 | 4590 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 6:2 FTS | 39.595034332565 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 7.92 | 7.92 | Baseline | Baseline | 255213.2065 | 10623.5559 | 24.0233 | 61335.0821 | 2601.5631 | 23.5762 | 4.161 | 4.0835 | 61339.4161 | 2604.1459 | 0.99866 | 0.99149 | 7.9345 | 7.9349 | 0.1835 | 0.014533 | 7.756 | 7.8034 | 8.2621 | 8.0948 | 0.5061 | 0.2914 | 0.062598 | 0.06125 | 124.4798 | 87.0984 | 0.0001917 | 0.0022711 | FALSE | 0.99996 | TRUE | [M-H]- | 98.9876 | 7.8664 | 8.0081 | 0.14173 | 7.8757 | 7.9966 | 0.12087 | 23.2323 | 0.98156 | 0.92102 | 2.0718 | 0 | 37 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 255213.2065 | No Control Sample | 0 | off | off | 255213.2065 | 0 | off | 0 | off | 0 | off | 0 | off | 426.97 | 4.9067 | 2.095 | 406.9617 | 0.54326 | 0.22109 | 0 | 0 | 0 | 0 | 0 | 7.6323 | FALSE | 10 | 24.0233
4591 | 4592 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 8:2 FTS | 41.429918735981 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.88 | 8.89 | Baseline | Baseline | 207609.6417 | 10213.153 | 20.3277 | 38636.3019 | 1877.8981 | 20.5742 | 5.3734 | 5.4386 | 38648.9859 | 1877.8981 | 0.99715 | 0.99795 | 8.8925 | 8.8961 | 0.14105 | 0.012525 | 8.725 | 8.6434 | 9.1052 | 9.1927 | 0.38017 | 0.54923 | 0.084446 | 0.084442 | 141.2249 | 91.316 | 0.0007714 | 9.7287e-14 | FALSE | 0.99959 | TRUE | [M-H]- | 103.5748 | 8.8063 | 8.9847 | 0.17835 | 8.8157 | 8.9697 | 0.15406 | 78.3969 | 1.1149 | 1.1813 | 1.951 | 0 | 20 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 207609.6417 | No Control Sample | 0 | off | off | 207609.6417 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9645 | 5.6294 | 2.9665 | 506.9571 | 2.7616 | 1.4 | 0 | 0 | 0 | 0 | 0 | 10.6509 | FALSE | 10 | 20.3277
4593 | 4594 | PFAS CS8 40ng/ml | 33 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 34) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 10:2 FTS | 41.8025897447866 | Quantifiers | 10/27/2023 03:58:55 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.5 | 8.89 | Baseline | Baseline | 218070.0709 | 10213.153 | 21.3519 | 51594.5781 | 1877.8981 | 27.4746 | 4.2266 | 5.4386 | 51601.7558 | 1877.8981 | 0.99727 | 0.99795 | 9.5122 | 8.8961 | 0.12877 | 0.012233 | 9.3894 | 8.6434 | 9.7088 | 9.1927 | 0.31932 | 0.54923 | 0.066289 | 0.084442 | 302.5617 | 91.316 | 0.00031183 | 9.7287e-14 | FALSE | 1.0693 | TRUE | [M-H]- | 104.5065 | 9.4415 | 9.5822 | 0.1407 | 9.4508 | 9.572 | 0.12117 | 50.3847 | 0.94768 | 0.86726 | 2.5166 | 0 | 23 | 7 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 218070.0709 | No Control Sample | 0 | off | off | 218070.0709 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9572 | 3.3623 | 2.108 | 606.9497 | 2.9825 | 1.8102 | 0 | 0 | 0 | 0 | 0 | 0.78444 | FALSE | 10 | 21.3519
4654 | 4655 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFBA | 42.3248392108194 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 1.49 | 1.48 | Baseline | Baseline | 747658.375 | 98479.618 | 7.592 | 97158.0654 | 13316.985 | 7.2958 | 7.6953 | 7.395 | 97243.609 | 13471.1982 | 0.99364 | 0.95167 | 1.4504 | 1.4465 | 2.6591 | 0.03962 | 1.286 | 1.3075 | 1.7214 | 1.6247 | 0.43542 | 0.31723 | 0.11951 | 0.11588 | 292.2882 | 179.1597 | 0.0017886 | 0.028398 | FALSE | 1.0026 | TRUE | [M-H]- | 105.8121 | 1.3572 | 1.617 | 0.25977 | 1.3664 | 1.593 | 0.2266 | 399.1101 | 1.5252 | 1.9837 | 1.3975 | 0 | 88 | 26 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 747658.375 | No Control Sample | 0 | off | off | 747658.375 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9804 | 5.4156 | 1.1534 | 168.9884 | -3.5989 | -0.60817 | 0 | 0 | 0 | 0 | 0 | 0.2754 | FALSE | 10 | 7.592
4656 | 4657 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFPeA | 37.3384018499951 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 3.18 | 3.21 | Baseline | Baseline | 643248.5158 | 77032.685 | 8.3503 | 64013.9256 | 7751.8827 | 8.2579 | 10.0486 | 9.9373 | 64128.6033 | 7768.2085 | 0.99023 | 0.9886 | 3.1691 | 3.1653 | 0.34366 | 0.010928 | 2.9112 | 2.9358 | 3.504 | 3.4402 | 0.59273 | 0.50432 | 0.15538 | 0.1561 | 796.5097 | 580.2809 | 0.0024372 | 0.0009368 | FALSE | 1.0012 | TRUE | [M-H]- | 93.346 | 3.0186 | 3.3624 | 0.3438 | 3.0376 | 3.3385 | 0.3009 | 263.2095 | 1.1787 | 1.3707 | 1.8467 | 0 | 106 | 29 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 643248.5158 | No Control Sample | 0 | off | off | 643248.5158 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9772 | 4.3705 | 1.1493 | 218.9858 | -0.033324 | -0.0072975 | 0 | 0 | 0 | 0 | 0 | 0.44744 | FALSE | 10 | 8.3503
4658 | 4659 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFHxA | 40.9220894048062 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 655.3414 | 6.13 | 6.12 | Baseline | Baseline | 1335871.3094 | 140268.0577 | 9.5237 | 224257.693 | 23466.1536 | 9.5566 | 5.9569 | 5.9775 | 224311.7939 | 23470.3486 | 0.99701 | 0.99741 | 6.1209 | 6.1198 | 0.14817 | 0.0090831 | 5.9196 | 5.9305 | 6.4164 | 6.3553 | 0.4968 | 0.42483 | 0.093814 | 0.091322 | 263.9726 | 218.601 | 0.00058067 | 0.00040151 | FALSE | 1.0002 | TRUE | [M-H]- | 102.3052 | 6.0238 | 6.2187 | 0.19485 | 6.0365 | 6.206 | 0.16945 | 262.1153 | 0.96307 | 0.91519 | 1.1514 | 0 | 55 | 12 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1335871.3094 | No Control Sample | 203843.5852 | off | off | 1335871.3094 | 203843.5852 | off | 0.025 | off | 40 | off | 262.1365 | off | 312.9737 | 2.7655 | 0.86554 | 268.9821 | -1.2104 | -0.32557 | 0 | 0 | 0 | 0 | 0 | 4.2469 | FALSE | 10 | 9.5237
4660 | 4661 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFHpA | 37.6070792168427 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 583.8727 | 7.26 | 7.27 | Baseline | Baseline | 1333149.3482 | 141670.9016 | 9.4102 | 263002.6607 | 31136.8871 | 8.4467 | 5.069 | 4.5499 | 263210.9771 | 31159.1922 | 0.99475 | 0.99361 | 7.2752 | 7.275 | 0.209 | 0.015173 | 7.123 | 7.1347 | 7.4412 | 7.4356 | 0.31823 | 0.30095 | 0.081052 | 0.0706 | 192.2345 | 238.5663 | 0.0015641 | 0.0015029 | FALSE | 1 | TRUE | [M-H]- | 94.0177 | 7.1998 | 7.3504 | 0.15057 | 7.2085 | 7.3414 | 0.13292 | 1292.6051 | 0.97956 | 0.95009 | 1.2532 | 0 | 38 | 11 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Pass | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1333149.3482 | No Control Sample | 228328.7703 | off | off | 1333149.3482 | 228328.7703 | off | 0.025 | off | 40 | off | 233.5491 | off | 362.9703 | 2.0026 | 0.72687 | 318.9779 | -3.607 | -1.1506 | 0 | 0 | 0 | 0 | 0 | 0.73367 | FALSE | 10 | 9.4102
4662 | 4663 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFOA | 37.5947305768353 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 566.7738 | 7.96 | 7.94 | Baseline | Baseline | 1085232.4579 | 119027.4363 | 9.1175 | 227711.8265 | 27517.8688 | 8.2751 | 4.7658 | 4.3255 | 227889.2837 | 27520.6948 | 0.9942 | 0.99854 | 7.9637 | 7.962 | 0.046626 | 0.0037115 | 7.8339 | 7.8268 | 8.1942 | 8.1856 | 0.36033 | 0.35885 | 0.07333 | 0.065313 | 124.0591 | 121.5942 | 0.00010426 | 0.00025079 | FALSE | 1.0002 | TRUE | [M-H]- | 93.9868 | 7.8944 | 8.0485 | 0.15415 | 7.9033 | 8.0365 | 0.1332 | 65.8897 | 1.0376 | 1.0376 | 1.0123 | 0 | 27 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1085232.4579 | No Control Sample | 191475.4056 | off | off | 1085232.4579 | 191475.4056 | off | 0.025 | off | 40 | off | 226.7095 | off | 412.9676 | 2.8768 | 1.188 | 368.9754 | -1.825 | -0.67339 | 0 | 0 | 0 | 0 | 0 | 0.77171 | FALSE | 10 | 9.1175
4664 | 4665 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFNA | 37.5009272558315 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 578.4905 | 8.46 | 8.47 | Baseline | Baseline | 920629.3751 | 105676.764 | 8.7117 | 172531.0104 | 21496.0998 | 8.0262 | 5.336 | 4.9161 | 172570.1097 | 21497.5543 | 0.99743 | 0.99892 | 8.4785 | 8.478 | 0.21817 | 0.018457 | 8.3061 | 8.2828 | 8.7575 | 8.7281 | 0.45143 | 0.44533 | 0.085106 | 0.076664 | 140.8414 | 190.411 | 0.00056723 | 0.0001507 | FALSE | 1.0001 | TRUE | [M-H]- | 93.7523 | 8.3909 | 8.5618 | 0.17087 | 8.4005 | 8.5499 | 0.14932 | 216.7884 | 0.95387 | 0.86823 | 0.84845 | 0 | 26 | 7 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 920629.3751 | No Control Sample | 159143.3885 | off | off | 920629.3751 | 159143.3885 | off | 0.025 | off | 40 | off | 231.3962 | off | 462.9645 | 2.8444 | 1.3169 | 418.9729 | -0.71976 | -0.30156 | 0 | 0 | 0 | 0 | 0 | 1.6795 | FALSE | 10 | 8.7117
4666 | 4667 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFDA | 43.7493629084857 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 8.87 | 8.86 | Baseline | Baseline | 845779.6737 | 77876.3827 | 10.8605 | 157428.8677 | 14567.2597 | 10.807 | 5.3725 | 5.346 | 157445.5502 | 14578.0334 | 0.99891 | 0.9955 | 8.8912 | 8.8872 | 0.23943 | 0.021238 | 8.6848 | 8.7113 | 9.2177 | 9.0527 | 0.53293 | 0.3414 | 0.084228 | 0.084704 | 152.9085 | 109.9575 | 0.00026486 | 0.0012604 | FALSE | 1.0005 | TRUE | [M-H]- | 109.3734 | 8.8044 | 8.9783 | 0.17391 | 8.8154 | 8.9636 | 0.14822 | 78.241 | 1.1216 | 1.2289 | 0.65637 | 0 | 28 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Fail | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 845779.6737 | No Control Sample | 0 | off | off | 845779.6737 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9627 | 5.2252 | 2.6803 | 468.9697 | -0.69093 | -0.32403 | 0 | 0 | 0 | 0 | 0 | 4.4582 | FALSE | 10 | 10.8605
4668 | 4669 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFUdA | 45.4523670922532 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.21 | 9.2 | Baseline | Baseline | 657037.2673 | 65183.4419 | 10.0798 | 135750.8297 | 14294.812 | 9.4965 | 4.84 | 4.5599 | 135767.1706 | 14304.4174 | 0.99788 | 0.99455 | 9.2247 | 9.2272 | 0.15927 | 0.014668 | 9.0454 | 9.0908 | 9.4978 | 9.3741 | 0.45247 | 0.28328 | 0.075981 | 0.071001 | 198.799 | 132.7628 | 0.00029224 | 0.00042104 | FALSE | 0.99972 | TRUE | [M-H]- | 113.6309 | 9.1462 | 9.3073 | 0.16109 | 9.1574 | 9.2951 | 0.13777 | 87.6793 | 1.1125 | 1.2494 | 0.61561 | 0 | 26 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 657037.2673 | No Control Sample | 0 | off | off | 657037.2673 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9595 | 4.7454 | 2.6714 | 518.9667 | -1.2187 | -0.63245 | 0 | 0 | 0 | 0 | 0 | 0.14713 | FALSE | 10 | 10.0798
4670 | 4671 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFDoA | 35.7348348101178 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.47 | 9.47 | Baseline | Baseline | 626157.0065 | 67711.7477 | 9.2474 | 142149.9338 | 16799.4618 | 8.4616 | 4.4049 | 4.0306 | 142192.989 | 16799.9898 | 0.99645 | 0.99318 | 9.4894 | 9.488 | 0.20515 | 0.019428 | 9.3517 | 9.3413 | 9.717 | 9.698 | 0.36533 | 0.35668 | 0.069855 | 0.062164 | 307.3549 | 259.7375 | 0.00080884 | 7.5693e-05 | FALSE | 1.0001 | TRUE | [M-H]- | 89.3371 | 9.4204 | 9.5601 | 0.13968 | 9.4277 | 9.5514 | 0.12364 | 314.7163 | 1.1171 | 1.2417 | 0.77342 | 0 | 26 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 626157.0065 | No Control Sample | 0 | off | off | 626157.0065 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9545 | 1.3206 | 0.80946 | 568.963 | -0.73175 | -0.41634 | 0 | 0 | 0 | 0 | 0 | 9.4042 | FALSE | 10 | 9.2474
4672 | 4673 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFTrDA | 40.8349233572104 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.68 | 9.47 | Baseline | Baseline | 521565.0376 | 67711.7477 | 7.7027 | 134546.0636 | 16799.4618 | 8.009 | 3.8765 | 4.0306 | 134597.3952 | 16799.9898 | 0.99647 | 0.99318 | 9.7036 | 9.488 | 0.24378 | 0.023598 | 9.5702 | 9.3413 | 9.8427 | 9.698 | 0.27253 | 0.35668 | 0.061548 | 0.062164 | 800.4089 | 259.7375 | 0.00069809 | 7.5693e-05 | FALSE | 1.0227 | TRUE | [M-H]- | 102.0873 | 9.6427 | 9.7645 | 0.12187 | 9.6483 | 9.7575 | 0.10917 | 344.6369 | 0.95521 | 0.87802 | 0.84475 | 0 | 24 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 521565.0376 | No Control Sample | 0 | off | off | 521565.0376 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9513 | 1.2947 | 0.85833 | 618.9608 | -1.4303 | -0.8853 | 0 | 0 | 0 | 0 | 0 | 2.4231 | FALSE | 10 | 7.7027
4674 | 4675 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFTeDA | 34.9890247599149 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.87 | 9.86 | Baseline | Baseline | 245590.9316 | 19797.3417 | 12.4052 | 69801.0761 | 5817.8101 | 11.9978 | 3.5184 | 3.4029 | 69820.3601 | 5824.0907 | 0.99608 | 0.99284 | 9.8797 | 9.8817 | 0.097834 | 0.0096562 | 9.7417 | 9.7752 | 10.0152 | 9.9864 | 0.2735 | 0.21112 | 0.055209 | 0.052914 | 1329.3085 | 754.7434 | 0.00056759 | 0.0020335 | FALSE | 0.99979 | TRUE | [M-H]- | 87.4726 | 9.8273 | 9.9421 | 0.1148 | 9.8344 | 9.9341 | 0.099718 | 144.8572 | 1.2385 | 1.543 | 1.6411 | 0 | 30 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 245590.9316 | No Control Sample | 0 | off | off | 245590.9316 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9483 | 1.4647 | 1.0442 | 668.9589 | 1.5201 | 1.0169 | 0 | 0 | 0 | 0 | 0 | 8.0683 | FALSE | 10 | 12.4052
4681 | 4682 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFBS | 36.7259187704713 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 3.5 | 3.5 | Baseline | Baseline | 399440.5147 | 41090.8855 | 9.7209 | 34220.9907 | 3677.5825 | 9.3053 | 11.6724 | 11.1733 | 34453.4705 | 3727.4917 | 0.97252 | 0.95414 | 3.8026 | 3.8055 | 8.6455 | 0.30259 | 3.4917 | 3.546 | 4.0834 | 4.0391 | 0.59172 | 0.49313 | 0.1762 | 0.18073 | 591.8345 | 843.9381 | 0.013439 | 0.021602 | FALSE | 0.99924 | TRUE | [M-H]- | 91.8148 | 3.6229 | 4.0188 | 0.39587 | 3.6446 | 3.9875 | 0.3429 | 777.2108 | 1.3822 | 1.8228 | 7.8922 | 0 | 92 | 30 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 399440.5147 | No Control Sample | 0 | off | off | 399440.5147 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9437 | 2.3485 | 0.70207 | 79.9568 | -5.2344 | -0.41853 | 0 | 0 | 0 | 0 | 0 | 1.7322 | FALSE | 10 | 9.7209
4683 | 4684 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFPeS | 41.0681129334653 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 6.38 | 3.5 | Baseline | Baseline | 449742.9743 | 41090.8855 | 10.9451 | 83312.124 | 3677.5825 | 22.654 | 5.3983 | 11.1733 | 83360.6663 | 3727.4917 | 0.99372 | 0.95414 | 6.3762 | 3.8055 | 0.05966 | 0.0038063 | 6.2194 | 3.546 | 6.6031 | 4.0391 | 0.38373 | 0.49313 | 0.083808 | 0.18073 | 293.9077 | 843.9381 | 0.0013142 | 0.021602 | FALSE | 1.6755 | TRUE | [M-H]- | 102.6703 | 6.2922 | 6.4661 | 0.17386 | 6.3031 | 6.4537 | 0.15057 | 285.325 | 0.98035 | 0.93568 | 5.6833 | 0 | 42 | 12 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 449742.9743 | No Control Sample | 0 | off | off | 449742.9743 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9405 | 2.1092 | 0.73597 | 79.9564 | -10.6098 | -0.84833 | 0 | 0 | 0 | 0 | 0 | 1.1163 | FALSE | 10 | 10.9451
4685 | 4686 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | L-PFHxS | 31.0324557550696 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 32.44 | 4 | 8.11 | 0 | 7.34 | 7.35 | Baseline | Baseline | 329068.165 | 56768.3318 | 5.7967 | 80343.8639 | 13682.0649 | 5.8722 | 4.0957 | 4.1491 | 81252.2246 | 13702.9241 | 0.75237 | 0.98841 | 7.3498 | 7.3489 | 0.13329 | 0.0097833 | 7.2619 | 7.2276 | 7.4933 | 7.4864 | 0.23135 | 0.25885 | 0.064072 | 0.065019 | 112.8855 | 200.0405 | 0.013472 | 0.0026091 | FALSE | 1.0001 | TRUE | [M-H]- | 95.6611 | 7.2864 | 7.4187 | 0.13236 | 7.2934 | 7.4098 | 0.11636 | -4678.6719 | 0.99934 | 0.96596 | 6.9404 | 0 | 29 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 329068.165 | No Control Sample | 0 | off | off | 329068.165 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9373 | 1.8511 | 0.73846 | 79.9565 | -8.2588 | -0.66035 | 0 | 0 | 0 | 0 | 0 | 0.91291 | FALSE | 8.11 | 5.7967
4687 | 4688 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | Br-PFHxS | 6.25966402911969 | Qualifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 7.56 | 4 | 1.89 | 0 | 7.16 | 7.35 | Manual | Baseline | 90182.9729 | 56768.3318 | 1.5886 | 15885.3293 | 13682.0649 | 1.161 | 5.6771 | 4.1491 | 81252.2246 | 13702.9241 | 0.20625 | 0.98841 | 7.1646 | 7.3489 | 0.064118 | 0.0045909 | 7.0516 | 7.2276 | 7.2619 | 7.4864 | 0.2103 | 0.25885 | 0.091468 | 0.065019 | 22.3194 | 200.0405 | 0.028932 | 0.0026091 | TRUE | 0.97492 | TRUE | [M-H]- | 82.7998 | 7.0809 | 7.2619 | 0.18108 | 7.0882 | 7.246 | 0.15786 | 2185.4657 | 1.0103 | 0.91795 | 0.27406 | 0 | 25 | 12 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 90182.9729 | No Control Sample | 0 | off | off | 329068.165 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9373 | 1.8511 | 0.73846 | 79.9486 | -107.8043 | -8.6197 | 0 | 0 | 0 | 0 | 0 | 0.91291 | FALSE | 1.89 | 1.5886
4689 | 4690 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFHpS | 39.9112630753659 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7.98 | 7.35 | Baseline | Baseline | 344825.034 | 56768.3318 | 6.0742 | 77376.7298 | 13682.0649 | 5.6553 | 4.4564 | 4.1491 | 77496.9765 | 13702.9241 | 0.98741 | 0.98841 | 7.9883 | 7.3489 | 0.10348 | 0.0082575 | 7.8634 | 7.2276 | 8.1344 | 7.4864 | 0.27097 | 0.25885 | 0.069607 | 0.065019 | 97.1586 | 200.0405 | 0.0017796 | 0.0026091 | FALSE | 1.087 | TRUE | [M-H]- | 99.7782 | 7.9214 | 8.0687 | 0.14732 | 7.9298 | 8.0577 | 0.12796 | 508.1739 | 1.0153 | 0.99467 | 7.0809 | 0 | 21 | 8 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 344825.034 | No Control Sample | 0 | off | off | 344825.034 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9344 | 2.1772 | 0.97742 | 79.9566 | -7.4481 | -0.59553 | 0 | 0 | 0 | 0 | 0 | 1.3357 | FALSE | 10 | 6.0742
4691 | 4692 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | L-PFOS | 30.8313946334062 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 31.52 | 4 | 7.88 | 0 | 8.46 | 8.45 | Manual | Baseline | 256010.6557 | 31945.4894 | 8.014 | 50161.1595 | 6300.8737 | 7.961 | 5.1038 | 5.07 | 50576.9596 | 6302.853 | 0.64524 | 0.99753 | 8.4748 | 8.4766 | 0.17537 | 0.014836 | 8.3989 | 8.3348 | 8.5781 | 8.6586 | 0.17921 | 0.32388 | 0.080735 | 0.079739 | 46.1486 | 146.2005 | 0.0050668 | 0.00054735 | TRUE | 0.99979 | TRUE | [M-H]- | 97.8153 | 8.382 | 8.5535 | 0.17148 | 8.382 | 8.545 | 0.16301 | -1418.1887 | 1.045 | 0.98678 | 7.9294 | 0 | 12 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 256010.6557 | No Control Sample | 0 | off | off | 256010.6557 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9306 | 0.75914 | 0.37876 | 79.9566 | -9.0803 | -0.72604 | 0 | 0 | 0 | 0 | 0 | 1.775 | FALSE | 7.88 | 8.014
4693 | 4694 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | Br-PFOS | 7.02779110284592 | Qualifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 8.44 | 4 | 2.11 | 0 | 8.3 | 8.45 | Manual | Baseline | 124929.7158 | 31945.4894 | 3.9107 | 21030.9071 | 6300.8737 | 3.3378 | 5.9403 | 5.07 | 50576.9596 | 6302.853 | 0.315 | 0.99753 | 8.3292 | 8.4766 | 0.35173 | 0.029193 | 8.2103 | 8.3348 | 8.3989 | 8.6586 | 0.1886 | 0.32388 | 0.098571 | 0.079739 | 19.3486 | 146.2005 | 0 | 0.00054735 | TRUE | 0.98261 | TRUE | [M-H]- | 83.2677 | 8.2367 | 8.4006 | 0.16388 | 8.2481 | 8.4006 | 0.15251 | 0 | 0.9071 | 0.93149 | 0.48799 | 0 | 13 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 124929.7158 | No Control Sample | 0 | off | off | 256010.6557 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9306 | 0.75914 | 0.37876 | 79.9578 | 6.4919 | 0.51907 | 0 | 0 | 0 | 0 | 0 | 1.775 | FALSE | 2.11 | 3.9107
4695 | 4696 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFNS | 40.9034722825026 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.85 | 8.45 | Baseline | Baseline | 289620.8687 | 31945.4894 | 9.0661 | 57355.342 | 6300.8737 | 9.1028 | 5.0496 | 5.07 | 57372.1178 | 6302.853 | 0.99523 | 0.99753 | 8.8729 | 8.4766 | 0.25931 | 0.022949 | 8.6531 | 8.3348 | 9.0861 | 8.6586 | 0.43303 | 0.32388 | 0.079344 | 0.079739 | 127.3471 | 146.2005 | 0.0005569 | 0.00054735 | FALSE | 1.0468 | TRUE | [M-H]- | 102.2587 | 8.7922 | 8.961 | 0.16876 | 8.8046 | 8.9466 | 0.14209 | 73.7615 | 1.145 | 1.3156 | 9.2709 | 0 | 23 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 289620.8687 | No Control Sample | 0 | off | off | 289620.8687 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9275 | 0.87862 | 0.4823 | 79.9566 | -5.838 | -0.46679 | 0 | 0 | 0 | 0 | 0 | 1.5158 | FALSE | 10 | 9.0661
4697 | 4698 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFDS | 42.7795224003681 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.19 | 8.45 | Baseline | Baseline | 311221.5829 | 31945.4894 | 9.7423 | 68261.7489 | 6300.8737 | 10.8337 | 4.5592 | 5.07 | 68274.7075 | 6302.853 | 0.99496 | 0.99753 | 9.2013 | 8.4766 | 0.12346 | 0.011346 | 9.0497 | 8.3348 | 9.5013 | 8.6586 | 0.45158 | 0.32388 | 0.070382 | 0.079739 | 190.6968 | 146.2005 | 0.00028955 | 0.00054735 | FALSE | 1.0855 | TRUE | [M-H]- | 106.9488 | 9.1279 | 9.2802 | 0.15237 | 9.1358 | 9.2711 | 0.13523 | 43.769 | 1.0078 | 0.99962 | 6.7818 | 0 | 26 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 311221.5829 | No Control Sample | 0 | off | off | 311221.5829 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9243 | 0.73617 | 0.44091 | 79.9566 | -4.082 | -0.32639 | 0 | 0 | 0 | 0 | 0 | 1.7668 | FALSE | 10 | 9.7423
4701 | 4702 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FBSA | 34.6791880550224 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 4.6 | 8.86 | Manual | Baseline | 595149.6271 | 65085.4989 | 9.1441 | 51608.5783 | 13985.6684 | 3.6901 | 11.532 | 4.6537 | 51903.2819 | 13991.9141 | 0.97803 | 0.99381 | 4.9305 | 9.0171 | 7.1848 | 0.3305 | 4.6862 | 8.8329 | 5.2042 | 9.1896 | 0.51801 | 0.35673 | 0.18091 | 0.07213 | 402.8429 | 152.5411 | 0.010108 | 0.00084745 | TRUE | 0.5468 | TRUE | [M-H]- | 86.698 | 4.7343 | 5.1424 | 0.40816 | 4.7615 | 5.1116 | 0.35018 | -1007.0079 | 1.1514 | 1.3337 | 4.1831 | 0 | 95 | 35 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 595149.6271 | No Control Sample | 0 | off | off | 595149.6271 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9598 | 2.8648 | 0.85359 | 77.9647 | -7.7108 | -0.60117 | 0 | 0 | 0 | 0 | 0 | 1.6701 | FALSE | 10 | 9.1441
4703 | 4704 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FPeSA | 42.1599241026076 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7 | 8.86 | Baseline | Baseline | 1123578.4464 | 65085.4989 | 17.2631 | 235422.0211 | 13985.6684 | 16.8331 | 4.7726 | 4.6537 | 235520.7566 | 13991.9141 | 0.99493 | 0.99381 | 7.1027 | 9.0171 | 1.4676 | 0.10273 | 6.932 | 8.8329 | 7.3517 | 9.1896 | 0.41973 | 0.35673 | 0.074455 | 0.07213 | 248.3836 | 152.5411 | 0.00097194 | 0.00084745 | FALSE | 0.7877 | TRUE | [M-H]- | 105.3998 | 7.0184 | 7.1819 | 0.16344 | 7.0302 | 7.1703 | 0.14005 | 545.1448 | 0.93739 | 0.85793 | 1.3722 | 0 | 47 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1123578.4464 | No Control Sample | 0 | off | off | 1123578.4464 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9564 | 1.913 | 0.66562 | 77.9645 | -10.5406 | -0.8218 | 0 | 0 | 0 | 0 | 0 | 38.7299 | FALSE | 10 | 17.2631
4705 | 4706 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FHxSA | 40.5124333297845 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7.86 | 8.86 | Baseline | Baseline | 538779.7885 | 65085.4989 | 8.278 | 120502.203 | 13985.6684 | 8.6161 | 4.4711 | 4.6537 | 120559.8686 | 13991.9141 | 0.99486 | 0.99381 | 7.9922 | 9.0171 | 1.6822 | 0.13222 | 7.8484 | 8.8329 | 8.2277 | 9.1896 | 0.37933 | 0.35673 | 0.069343 | 0.07213 | 113.7597 | 152.5411 | 0.00088716 | 0.00084745 | FALSE | 0.88634 | TRUE | [M-H]- | 101.2811 | 7.924 | 8.0754 | 0.15133 | 7.9331 | 8.0646 | 0.13146 | 281.8236 | 1.1102 | 1.2262 | 3.7092 | 0 | 28 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 538779.7885 | No Control Sample | 0 | off | off | 538779.7885 | 0 | off | 0 | off | 0 | off | 0 | off | 397.954 | 3.6605 | 1.4567 | 77.9646 | -10.6107 | -0.82727 | 0 | 0 | 0 | 0 | 0 | 1.3273 | FALSE | 10 | 8.278
4707 | 4708 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FHpSA | 37.2248395066445 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.58 | 8.86 | Baseline | Baseline | 1025400.1299 | 65085.4989 | 15.7547 | 221688.0013 | 13985.6684 | 15.8511 | 4.6254 | 4.6537 | 222085.0924 | 13991.9141 | 0.98958 | 0.99381 | 8.5743 | 9.0171 | 0.066332 | 0.0056913 | 8.4414 | 8.8329 | 8.8429 | 9.1896 | 0.40147 | 0.35673 | 0.072653 | 0.07213 | 119.9706 | 152.5411 | 0.0018967 | 0.00084745 | FALSE | 0.9509 | TRUE | [M-H]- | 93.0621 | 8.5028 | 8.6529 | 0.15014 | 8.5096 | 8.6434 | 0.13386 | -1047.373 | 1.1304 | 1.2448 | 1.8594 | 0 | 23 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1025400.1299 | No Control Sample | 0 | off | off | 1025400.1299 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9498 | 0.99685 | 0.44654 | 77.9641 | -16.2614 | -1.2678 | 0 | 0 | 0 | 0 | 0 | 1.0256 | FALSE | 10 | 15.7547
4709 | 4710 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FOSA | 42.2952997372728 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.88 | 8.86 | Baseline | Baseline | 626356.8292 | 65085.4989 | 9.6236 | 129694.182 | 13985.6684 | 9.2734 | 4.8295 | 4.6537 | 129743.2965 | 13991.9141 | 0.9921 | 0.99381 | 9.0168 | 9.0171 | 1.541 | 0.13684 | 8.8438 | 8.8329 | 9.2407 | 9.1896 | 0.39692 | 0.35673 | 0.075892 | 0.07213 | 131.2238 | 152.5411 | 0.00088242 | 0.00084745 | FALSE | 0.99998 | TRUE | [M-H]- | 105.7382 | 8.9358 | 9.097 | 0.16118 | 8.9472 | 9.0842 | 0.13695 | 288.3339 | 1.0863 | 1.1829 | 3.7879 | 0 | 21 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 626356.8292 | No Control Sample | 0 | off | off | 626356.8292 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9467 | 1.0734 | 0.5345 | 77.9645 | -11.6156 | -0.90561 | 0 | 0 | 0 | 0 | 0 | 0.78274 | FALSE | 10 | 9.6236
4728 | 4729 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 4:2 FTS | 37.9526222918303 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 5.94 | 5.92 | Baseline | Baseline | 299349.5638 | 15534.153 | 19.2704 | 48506.6811 | 2619.1209 | 18.5202 | 6.1713 | 5.9311 | 48543.5631 | 2652.6835 | 0.99456 | 0.95782 | 5.9282 | 5.9286 | 0.1986 | 0.011797 | 5.7376 | 5.7781 | 6.1276 | 6.0466 | 0.38997 | 0.26848 | 0.098071 | 0.090386 | 211.5827 | 188.0575 | 0.00098694 | 0.0056951 | FALSE | 0.99994 | TRUE | [M-H]- | 94.8816 | 5.8236 | 6.0312 | 0.2076 | 5.8413 | 6.0178 | 0.17655 | 122.7628 | 0.97888 | 0.99808 | 3.8019 | 0 | 48 | 13 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 299349.5638 | No Control Sample | 0 | off | off | 299349.5638 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9754 | 3.4887 | 1.1407 | 306.9682 | 1.8005 | 0.5527 | 0 | 0 | 0 | 0 | 0 | 5.3291 | FALSE | 10 | 19.2704
4730 | 4731 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 6:2 FTS | 39.8620257874915 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7.92 | 7.92 | Baseline | Baseline | 256475.6642 | 10604.5998 | 24.1853 | 62391.1928 | 2555.73 | 24.4123 | 4.1108 | 4.1493 | 62418.9964 | 2557.3533 | 0.99515 | 0.99389 | 7.9251 | 7.9222 | 0.063825 | 0.0050549 | 7.7568 | 7.7914 | 8.1048 | 8.0477 | 0.34808 | 0.25625 | 0.062597 | 0.06323 | 113.8753 | 134.7405 | 0.00087068 | 0.0012709 | FALSE | 1.0004 | TRUE | [M-H]- | 99.6551 | 7.8542 | 7.9945 | 0.14026 | 7.8646 | 7.9842 | 0.11962 | 156.0633 | 0.95861 | 0.90406 | 2.4328 | 0 | 27 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 256475.6642 | No Control Sample | 0 | off | off | 256475.6642 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9698 | 4.4805 | 1.913 | 406.9616 | 0.34243 | 0.13935 | 0 | 0 | 0 | 0 | 0 | 9.0623 | FALSE | 10 | 24.1853
4732 | 4733 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 8:2 FTS | 38.8106014443764 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.88 | 8.89 | Baseline | Baseline | 207206.1222 | 10881.2456 | 19.0425 | 39155.7703 | 1999.9118 | 19.5787 | 5.2918 | 5.4409 | 39189.9256 | 1999.9118 | 0.99404 | 1 | 8.8854 | 8.886 | 0.061339 | 0.0054469 | 8.7264 | 8.6613 | 9.0692 | 9.1752 | 0.34282 | 0.51397 | 0.083019 | 0.085334 | 108.6972 | 122.7213 | 0.0018005 | 0 | FALSE | 0.99994 | TRUE | [M-H]- | 97.0265 | 8.7983 | 8.9743 | 0.17594 | 8.8095 | 8.9642 | 0.1547 | 205.6447 | 0.98075 | 0.96945 | 2.1129 | 0 | 18 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 207206.1222 | No Control Sample | 0 | off | off | 207206.1222 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9643 | 5.3287 | 2.808 | 506.9562 | 0.96496 | 0.48919 | 0 | 0 | 0 | 0 | 0 | 7.2836 | FALSE | 10 | 19.0425
4734 | 4735 | PFAS CS8 40ng/ml | 34 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 35) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 10:2 FTS | 38.5680728073732 | Quantifiers | 10/27/2023 04:15:31 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.5 | 8.89 | Baseline | Baseline | 214357.9465 | 10881.2456 | 19.6998 | 49728.7521 | 1999.9118 | 24.8655 | 4.3105 | 5.4409 | 49763.4906 | 1999.9118 | 0.99525 | 1 | 9.506 | 8.886 | 0.063215 | 0.0060054 | 9.3732 | 8.6613 | 9.6446 | 9.1752 | 0.27135 | 0.51397 | 0.067571 | 0.085334 | 329.9224 | 122.7213 | 0.001449 | 0 | FALSE | 1.0698 | TRUE | [M-H]- | 96.4202 | 9.4343 | 9.5764 | 0.14216 | 9.4424 | 9.5674 | 0.12506 | 265.5524 | 1.0543 | 1.1082 | 1.7037 | 0 | 19 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 214357.9465 | No Control Sample | 0 | off | off | 214357.9465 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9569 | 2.7714 | 1.7375 | 606.9498 | 3.1081 | 1.8864 | 0 | 0 | 0 | 0 | 0 | 0.80093 | FALSE | 10 | 19.6998
4795 | 4796 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFBA | 40.2349934993275 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 1.49 | 1.48 | Baseline | Baseline | 766033.4268 | 106140.7752 | 7.2171 | 98321.5233 | 14240.327 | 6.9044 | 7.7911 | 7.4535 | 98418.1847 | 14261.044 | 0.9932 | 0.99144 | 1.4548 | 1.4563 | 2.3623 | 0.035198 | 1.2886 | 1.305 | 1.743 | 1.6905 | 0.45445 | 0.38553 | 0.1142 | 0.1148 | 285.0341 | 179.7158 | 0.0017524 | 0.0028165 | FALSE | 0.99896 | TRUE | [M-H]- | 100.5875 | 1.3627 | 1.6256 | 0.26293 | 1.3719 | 1.6005 | 0.22857 | 379.1417 | 1.5969 | 2.1253 | 1.3663 | 0 | 92 | 25 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 766033.4268 | No Control Sample | 0 | off | off | 766033.4268 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9803 | 4.9956 | 1.064 | 168.9885 | -2.9227 | -0.4939 | 0 | 0 | 0 | 0 | 0 | 0.26068 | FALSE | 10 | 7.2171
4797 | 4798 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFPeA | 37.472109147454 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 3.18 | 3.21 | Baseline | Baseline | 640675.6112 | 76450.7978 | 8.3802 | 66088.0496 | 7753.6691 | 8.5235 | 9.6943 | 9.86 | 66349.5638 | 7949.0864 | 0.98349 | 0.93428 | 3.1798 | 3.1795 | 0.0048473 | 0.00015415 | 2.9629 | 3.0112 | 3.515 | 3.4119 | 0.55208 | 0.40067 | 0.15059 | 0.15479 | 879.9385 | 441.3434 | 0.0022889 | 0.016835 | FALSE | 1.0001 | TRUE | [M-H]- | 93.6803 | 3.0316 | 3.3765 | 0.34495 | 3.0489 | 3.3477 | 0.29875 | -274.0024 | 1.0738 | 1.0845 | 1.8479 | 0 | 97 | 29 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 640675.6112 | No Control Sample | 0 | off | off | 640675.6112 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9772 | 4.6355 | 1.219 | 218.986 | 0.7481 | 0.16382 | 0 | 0 | 0 | 0 | 0 | 0.3515 | FALSE | 10 | 8.3802
4799 | 4800 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFHxA | 37.6603683941267 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 600.8173 | 6.13 | 6.12 | Baseline | Baseline | 1224727.5461 | 139735.5416 | 8.7646 | 198038.7231 | 23283.7353 | 8.5055 | 6.1843 | 6.0014 | 198104.1275 | 23299.5459 | 0.99646 | 0.99445 | 6.136 | 6.1335 | 0.098043 | 0.00601 | 5.9313 | 5.9505 | 6.4106 | 6.3218 | 0.47923 | 0.3713 | 0.097259 | 0.094514 | 258.4183 | 265.6048 | 0.00080314 | 0.0013819 | FALSE | 1.0004 | TRUE | [M-H]- | 94.1509 | 6.0356 | 6.2347 | 0.19915 | 6.0476 | 6.2218 | 0.17416 | 331.8896 | 1.0726 | 1.1554 | 1.2758 | 0 | 53 | 13 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1224727.5461 | No Control Sample | 203843.5852 | off | off | 1224727.5461 | 203843.5852 | off | 0.025 | off | 40 | off | 240.3269 | off | 312.9735 | 2.3463 | 0.73432 | 268.9819 | -1.9128 | -0.5145 | 0 | 0 | 0 | 0 | 0 | 3.5676 | FALSE | 10 | 8.7646
4801 | 4802 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFHpA | 35.9140093740118 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 569.6878 | 7.26 | 7.27 | Baseline | Baseline | 1300761.1728 | 144745.5202 | 8.9865 | 257786.6071 | 32737.6242 | 7.8743 | 5.0459 | 4.4214 | 257890.8052 | 32760.8572 | 0.99623 | 0.99485 | 7.2829 | 7.2841 | 0.31586 | 0.022931 | 7.1428 | 7.1454 | 7.492 | 7.4201 | 0.34923 | 0.2747 | 0.081483 | 0.06871 | 196.3259 | 266.9451 | 0.00094246 | 0.0010029 | FALSE | 0.99984 | TRUE | [M-H]- | 89.785 | 7.2087 | 7.361 | 0.15239 | 7.2173 | 7.3519 | 0.13456 | 695.6766 | 0.98538 | 0.96038 | 1.2673 | 0 | 42 | 12 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Pass | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1300761.1728 | No Control Sample | 228328.7703 | off | off | 1300761.1728 | 228328.7703 | off | 0.025 | off | 40 | off | 227.8751 | off | 362.9703 | 1.9169 | 0.69576 | 318.9783 | -2.5722 | -0.82047 | 0 | 0 | 0 | 0 | 0 | 0.97255 | FALSE | 10 | 8.9865
4803 | 4804 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFOA | 38.6916122491214 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 571.7557 | 7.96 | 7.94 | Baseline | Baseline | 1094771.6118 | 116669.6735 | 9.3835 | 225142.5189 | 26956.0717 | 8.3522 | 4.8626 | 4.3281 | 225280.7784 | 26957.0834 | 0.99542 | 0.99869 | 7.9706 | 7.9687 | 0.13318 | 0.010601 | 7.8339 | 7.8141 | 8.2419 | 8.2016 | 0.40798 | 0.38743 | 0.076475 | 0.065513 | 127.0742 | 105.6149 | 0.00022486 | 9.1714e-05 | FALSE | 1.0002 | TRUE | [M-H]- | 96.729 | 7.9018 | 8.0541 | 0.15225 | 7.9097 | 8.0444 | 0.13467 | -124.0857 | 1.1338 | 1.2714 | 0.95736 | 0 | 30 | 9 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1094771.6118 | No Control Sample | 191475.4056 | off | off | 1094771.6118 | 191475.4056 | off | 0.025 | off | 40 | off | 228.7023 | off | 412.9673 | 2.0774 | 0.85788 | 368.9753 | -2.0537 | -0.75776 | 0 | 0 | 0 | 0 | 0 | 0.3268 | FALSE | 10 | 9.3835
4805 | 4806 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFNA | 42.5186468858916 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 597.4531 | 8.46 | 8.47 | Baseline | Baseline | 950807.0567 | 96260.8396 | 9.8774 | 183480.4984 | 18535.897 | 9.8987 | 5.1821 | 5.1932 | 183511.2404 | 18559.8453 | 0.99807 | 0.99267 | 8.4847 | 8.4816 | 0.2921 | 0.024712 | 8.3065 | 8.2828 | 8.7202 | 8.6395 | 0.41368 | 0.3567 | 0.082064 | 0.081339 | 144.39 | 141.2961 | 0.00030463 | 0.0021007 | FALSE | 1.0004 | TRUE | [M-H]- | 106.2966 | 8.398 | 8.564 | 0.16598 | 8.4064 | 8.5535 | 0.1471 | 135.1134 | 1.0045 | 0.98054 | 0.76724 | 0 | 24 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 950807.0567 | No Control Sample | 159143.3885 | off | off | 950807.0567 | 159143.3885 | off | 0.025 | off | 40 | off | 238.9812 | off | 462.9647 | 3.1064 | 1.4381 | 418.9724 | -1.8976 | -0.79503 | 0 | 0 | 0 | 0 | 0 | 0.70896 | FALSE | 10 | 9.8774
4807 | 4808 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFDA | 38.5570275173021 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.87 | 8.86 | Baseline | Baseline | 810729.4144 | 84701.801 | 9.5716 | 151796.3654 | 17020.3219 | 8.9185 | 5.3409 | 4.9765 | 151908.2233 | 17032.0823 | 0.9951 | 0.99517 | 8.8944 | 8.8927 | 0.27519 | 0.02441 | 8.7396 | 8.7111 | 9.0629 | 9.0711 | 0.32332 | 0.36 | 0.084423 | 0.077299 | 128.1352 | 145.1734 | 0.0013023 | 0.0012539 | FALSE | 1.0002 | TRUE | [M-H]- | 96.3926 | 8.8082 | 8.9819 | 0.17378 | 8.8207 | 8.9709 | 0.15011 | 611.4466 | 0.92586 | 0.84702 | 0.51202 | 0 | 17 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 810729.4144 | No Control Sample | 0 | off | off | 810729.4144 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9624 | 4.5947 | 2.3569 | 468.9693 | -1.5386 | -0.72157 | 0 | 0 | 0 | 0 | 0 | 2.068 | FALSE | 10 | 9.5716
4809 | 4810 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFUdA | 44.3184510629736 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.21 | 9.2 | Baseline | Baseline | 678173.1137 | 69001.6993 | 9.8284 | 135454.6823 | 14326.6997 | 9.4547 | 5.0066 | 4.8163 | 135478.184 | 14331.265 | 0.99771 | 0.99701 | 9.2295 | 9.2315 | 0.21178 | 0.019505 | 9.0452 | 9.072 | 9.5125 | 9.407 | 0.46732 | 0.33503 | 0.078876 | 0.075657 | 259.0485 | 120.8069 | 0.00039845 | 0.00068639 | FALSE | 0.99979 | TRUE | [M-H]- | 110.7961 | 9.1482 | 9.3115 | 0.16332 | 9.1605 | 9.3002 | 0.13969 | 115.4941 | 0.91151 | 0.80796 | 0.53157 | 0 | 27 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 678173.1137 | No Control Sample | 0 | off | off | 678173.1137 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9585 | 2.9193 | 1.6434 | 518.9657 | -3.1405 | -1.6298 | 0 | 0 | 0 | 0 | 0 | 0.51616 | FALSE | 10 | 9.8284
4811 | 4812 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFDoA | 44.8216077581489 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.47 | 9.47 | Baseline | Baseline | 691017.0284 | 59576.3479 | 11.5988 | 156723.4215 | 13698.2569 | 11.4411 | 4.4091 | 4.3492 | 156750.2234 | 13700.0328 | 0.99784 | 0.99358 | 9.4921 | 9.4932 | 0.23333 | 0.022096 | 9.3521 | 9.3418 | 9.7402 | 9.6467 | 0.38805 | 0.30495 | 0.06931 | 0.068162 | 254.1681 | 254.5106 | 0.00041165 | 0.00024188 | FALSE | 0.99988 | TRUE | [M-H]- | 112.054 | 9.4243 | 9.565 | 0.14071 | 9.4342 | 9.5547 | 0.12049 | 166.2565 | 0.93908 | 0.85414 | 1.1227 | 0 | 28 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 691017.0284 | No Control Sample | 0 | off | off | 691017.0284 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9541 | 0.65939 | 0.40418 | 568.9627 | -1.2282 | -0.69882 | 0 | 0 | 0 | 0 | 0 | 8.0996 | FALSE | 10 | 11.5988
4813 | 4814 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFTrDA | 46.5011867034099 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy | 0 | 0 | 40 | 4 | 10 | 0 | 9.68 | 9.47 | Baseline | Baseline | 522577.3906 | 59576.3479 | 8.7716 | 136522.6578 | 13698.2569 | 9.9664 | 3.8278 | 4.3492 | 136594.4369 | 13700.0328 | 0.99554 | 0.99358 | 9.7088 | 9.4932 | 0.29708 | 0.028757 | 9.5566 | 9.3418 | 9.8426 | 9.6467 | 0.286 | 0.30495 | 0.060584 | 0.068162 | 579.2475 | 254.5106 | 0.00096675 | 0.00024188 | FALSE | 1.0227 | TRUE | [M-H]- | 116.253 | 9.646 | 9.7711 | 0.12509 | 9.6549 | 9.7633 | 0.10844 | 461.4814 | 1.0281 | 1.0869 | 0.822 | 0 | 25 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 522577.3906 | No Control Sample | 0 | off | off | 522577.3906 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9518 | 1.9836 | 1.3151 | 618.9602 | -2.3588 | -1.46 | 0 | 0 | 0 | 0 | 0 | 1.8485 | FALSE | 10 | 8.7716
4815 | 4816 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFTeDA | 40.9956124202388 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.87 | 9.86 | Baseline | Baseline | 300503.9826 | 20674.7012 | 14.5349 | 90729.3158 | 5839.091 | 15.5383 | 3.3121 | 3.5407 | 90766.9325 | 5847.4433 | 0.99514 | 0.99196 | 9.8819 | 9.8855 | 0.12049 | 0.011892 | 9.7533 | 9.7953 | 10.0285 | 9.986 | 0.27522 | 0.19062 | 0.051406 | 0.056139 | 1112.3494 | 879.2751 | 0.00082767 | 0.0028895 | FALSE | 0.99964 | TRUE | [M-H]- | 102.489 | 9.834 | 9.9445 | 0.1105 | 9.8395 | 9.9361 | 0.096593 | 272.8543 | 1.1228 | 1.206 | 1.2584 | 0 | 31 | 7 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 300503.9826 | No Control Sample | 0 | off | off | 300503.9826 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9482 | 1.3102 | 0.93409 | 668.9589 | 1.5138 | 1.0126 | 0 | 0 | 0 | 0 | 0 | 7.4896 | FALSE | 10 | 14.5349
4822 | 4823 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFBS | 37.2871484655414 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 3.5 | 3.5 | Baseline | Manual | 432011.5155 | 43772.5864 | 9.8695 | 38424.8449 | 3942.3357 | 9.7467 | 11.243 | 11.1032 | 38520.195 | 3952.8446 | 0.98861 | 0.9882 | 3.8201 | 3.8139 | 9.1443 | 0.32005 | 3.529 | 3.5864 | 4.1859 | 4.125 | 0.65687 | 0.53854 | 0.17482 | 0.17446 | 667.8634 | 837.9877 | 0.0046145 | 0.0016711 | FALSE | 1.0016 | TRUE | [M-H]- | 93.2179 | 3.6396 | 4.0305 | 0.39083 | 3.6578 | 3.9984 | 0.34062 | 269.9364 | 1.153 | 1.2506 | 7.443 | 0 | 102 | 29 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 432011.5155 | No Control Sample | 0 | off | off | 432011.5155 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9435 | 1.836 | 0.54885 | 79.9568 | -4.7729 | -0.38163 | 0 | 0 | 0 | 0 | 0 | 5.1374 | FALSE | 10 | 9.8695
4824 | 4825 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFPeS | 38.8910258192815 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 6.38 | 3.5 | Baseline | Manual | 453696.8318 | 43772.5864 | 10.3649 | 85411.3256 | 3942.3357 | 21.6652 | 5.3119 | 11.1032 | 85497.8682 | 3952.8446 | 0.99149 | 0.9882 | 6.3884 | 3.8139 | 0.13113 | 0.008366 | 6.2318 | 3.5864 | 6.5788 | 4.125 | 0.34702 | 0.53854 | 0.083095 | 0.17446 | 299.331 | 837.9877 | 0.0022071 | 0.0016711 | FALSE | 1.675 | TRUE | [M-H]- | 97.2276 | 6.3043 | 6.4772 | 0.17287 | 6.3151 | 6.4653 | 0.15018 | 543.2395 | 1.0831 | 1.1767 | 5.6809 | 0 | 38 | 12 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 453696.8318 | No Control Sample | 0 | off | off | 453696.8318 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9401 | 0.73564 | 0.25669 | 79.9564 | -9.7819 | -0.78213 | 0 | 0 | 0 | 0 | 0 | 1.1014 | FALSE | 10 | 10.3649
4826 | 4827 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | L-PFHxS | 31.5758798662358 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 32.44 | 4 | 8.11 | 0 | 7.34 | 7.35 | Baseline | Baseline | 340583.0105 | 57743.6093 | 5.8982 | 83762.5296 | 13758.3283 | 6.0881 | 4.0661 | 4.197 | 84652.7841 | 13782.5435 | 0.76304 | 0.9843 | 7.3592 | 7.3593 | 0.2615 | 0.019194 | 7.2712 | 7.2387 | 7.4853 | 7.487 | 0.21403 | 0.24832 | 0.063178 | 0.065871 | 122.7862 | 180.2013 | 0.013003 | 0.0024395 | FALSE | 0.99999 | TRUE | [M-H]- | 97.3363 | 7.2955 | 7.4274 | 0.13185 | 7.3033 | 7.4191 | 0.11574 | -5088.7508 | 0.99396 | 0.97826 | 6.7885 | 0 | 27 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 340583.0105 | No Control Sample | 0 | off | off | 340583.0105 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9373 | 1.6817 | 0.67088 | 79.9564 | -9.5462 | -0.76329 | 0 | 0 | 0 | 0 | 0 | 1.0999 | FALSE | 8.11 | 5.8982
4828 | 4829 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | Br-PFHxS | 6.05083676277693 | Qualifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 7.56 | 4 | 1.89 | 0 | 7.16 | 7.35 | Manual | Baseline | 88672.0512 | 57743.6093 | 1.5356 | 16381.4404 | 13758.3283 | 1.1907 | 5.413 | 4.197 | 84652.7841 | 13782.5435 | 0.19873 | 0.9843 | 7.1749 | 7.3593 | 0.20776 | 0.014875 | 7.0735 | 7.2387 | 7.2693 | 7.487 | 0.19576 | 0.24832 | 0.086311 | 0.065871 | 24.0133 | 180.2013 | 0.0057416 | 0.0024395 | TRUE | 0.97494 | TRUE | [M-H]- | 80.0375 | 7.0925 | 7.2712 | 0.17868 | 7.1002 | 7.2539 | 0.1537 | 480.4697 | 1.0084 | 0.89913 | 0.26035 | 0 | 23 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 88672.0512 | No Control Sample | 0 | off | off | 340583.0105 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9373 | 1.6817 | 0.67088 | 79.6856 | -3397.1689 | -271.6281 | 0 | 0 | 0 | 0 | 0 | 1.0999 | FALSE | 1.89 | 1.5356
4830 | 4831 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFHpS | 38.365529628834 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7.98 | 7.35 | Baseline | Baseline | 337164.862 | 57743.6093 | 5.839 | 77701.0364 | 13758.3283 | 5.6476 | 4.3393 | 4.197 | 77780.1125 | 13782.5435 | 0.992 | 0.9843 | 7.9947 | 7.3593 | 0.18436 | 0.014712 | 7.8637 | 7.2387 | 8.1819 | 7.487 | 0.31812 | 0.24832 | 0.067855 | 0.065871 | 124.7496 | 180.2013 | 0.00039114 | 0.0024395 | FALSE | 1.0863 | TRUE | [M-H]- | 95.9138 | 7.9295 | 8.0747 | 0.14522 | 7.9387 | 8.0654 | 0.12676 | 95.5375 | 1.1204 | 1.2784 | 6.0801 | 0 | 24 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 337164.862 | No Control Sample | 0 | off | off | 337164.862 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9339 | 1.1876 | 0.53317 | 79.9564 | -10.5349 | -0.84234 | 0 | 0 | 0 | 0 | 0 | 1.8281 | FALSE | 10 | 5.839
4832 | 4833 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | L-PFOS | 30.856148912324 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 31.52 | 4 | 7.88 | 0 | 8.46 | 8.45 | Manual | Baseline | 261113.0352 | 32556.0349 | 8.0204 | 51834.675 | 6553.7326 | 7.9092 | 5.0374 | 4.9676 | 51974.0612 | 6559.5091 | 0.65037 | 0.9881 | 8.4799 | 8.4809 | 0.2352 | 0.019898 | 8.3995 | 8.3534 | 8.5829 | 8.7319 | 0.18338 | 0.37848 | 0.078923 | 0.078193 | 55.0496 | 160 | 0.0050386 | 3.1314e-05 | TRUE | 0.99988 | TRUE | [M-H]- | 97.8939 | 8.3823 | 8.56 | 0.17774 | 8.3823 | 8.5501 | 0.16785 | -1424.2208 | 0.86492 | 0.63356 | 7.7927 | 0 | 12 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 261113.0352 | No Control Sample | 0 | off | off | 261113.0352 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9316 | 2.8064 | 1.4002 | 79.9565 | -9.5323 | -0.76217 | 0 | 0 | 0 | 0 | 0 | 1.9725 | FALSE | 7.88 | 8.0204
4834 | 4835 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | Br-PFOS | 7.1548434705079 | Qualifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 8.44 | 4 | 2.11 | 0 | 8.3 | 8.45 | Manual | Baseline | 129619.1009 | 32556.0349 | 3.9814 | 21481.2639 | 6553.7326 | 3.2777 | 6.0341 | 4.9676 | 51974.0612 | 6559.5091 | 0.32292 | 0.9881 | 8.3315 | 8.4809 | 0.37911 | 0.031466 | 8.2119 | 8.3534 | 8.4021 | 8.7319 | 0.19021 | 0.37848 | 0.10051 | 0.078193 | 22.8136 | 160 | 0.0065985 | 3.1314e-05 | TRUE | 0.98238 | TRUE | [M-H]- | 84.773 | 8.2378 | 8.422 | 0.18427 | 8.2494 | 8.422 | 0.17265 | -745.215 | 1.0282 | 1.214 | 0.49641 | 0 | 14 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 129619.1009 | No Control Sample | 0 | off | off | 261113.0352 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9316 | 2.8064 | 1.4002 | 79.9558 | -18.9545 | -1.5156 | 0 | 0 | 0 | 0 | 0 | 1.9725 | FALSE | 2.11 | 3.9814
4836 | 4837 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFNS | 42.1360562940384 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.85 | 8.45 | Baseline | Baseline | 304050.3588 | 32556.0349 | 9.3393 | 60583.5869 | 6553.7326 | 9.2441 | 5.0187 | 4.9676 | 60619.7859 | 6559.5091 | 0.99249 | 0.9881 | 8.8772 | 8.4809 | 0.30721 | 0.027188 | 8.67 | 8.3534 | 9.0859 | 8.7319 | 0.41582 | 0.37848 | 0.078346 | 0.078193 | 118.2064 | 160 | 0.00091953 | 3.1314e-05 | FALSE | 1.0467 | TRUE | [M-H]- | 105.3401 | 8.7992 | 8.9631 | 0.16386 | 8.8079 | 8.9496 | 0.14165 | 133.9742 | 0.9443 | 0.81407 | 7.7632 | 0 | 22 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 304050.3588 | No Control Sample | 0 | off | off | 304050.3588 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9275 | 0.90155 | 0.49488 | 79.9566 | -6.6065 | -0.52824 | 0 | 0 | 0 | 0 | 0 | 2.3507 | FALSE | 10 | 9.3393
4838 | 4839 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFDS | 41.2571041489893 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.19 | 8.45 | Baseline | Baseline | 305882.3911 | 32556.0349 | 9.3956 | 67104.0896 | 6553.7326 | 10.2391 | 4.5583 | 4.9676 | 67121.3639 | 6559.5091 | 0.99499 | 0.9881 | 9.204 | 8.4809 | 0.15274 | 0.014037 | 9.0495 | 8.3534 | 9.4218 | 8.7319 | 0.37225 | 0.37848 | 0.070782 | 0.078193 | 250.5288 | 160 | 0.00058145 | 3.1314e-05 | FALSE | 1.0853 | TRUE | [M-H]- | 103.1428 | 9.1306 | 9.2856 | 0.15499 | 9.1409 | 9.2754 | 0.13454 | 104.8154 | 1.0678 | 1.1604 | 7.385 | 0 | 21 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 305882.3911 | No Control Sample | 0 | off | off | 305882.3911 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9237 | -0.16555 | -0.099151 | 79.9565 | -4.4878 | -0.35883 | 0 | 0 | 0 | 0 | 0 | 1.5552 | FALSE | 10 | 9.3956
4842 | 4843 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FBSA | 34.6390781653805 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 4.6 | 8.86 | Manual | Baseline | 603490.6158 | 66074.0897 | 9.1335 | 51103.6562 | 14444.376 | 3.538 | 11.8091 | 4.5744 | 51361.7554 | 14453.9087 | 0.97918 | 0.99253 | 4.9461 | 9.019 | 7.5237 | 0.34609 | 4.6924 | 8.8942 | 5.2074 | 9.1892 | 0.51506 | 0.29502 | 0.18008 | 0.070823 | 426.2958 | 142.3241 | 0 | 1.3143e-08 | TRUE | 0.54841 | TRUE | [M-H]- | 86.5977 | 4.748 | 5.1619 | 0.41393 | 4.7742 | 5.1304 | 0.35618 | 0 | 1.089 | 1.1734 | 4.5785 | 0 | 94 | 35 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 603490.6158 | No Control Sample | 0 | off | off | 603490.6158 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9598 | 2.6834 | 0.79953 | 77.9648 | -6.9184 | -0.53939 | 0 | 0 | 0 | 0 | 0 | 1.0605 | FALSE | 10 | 9.1335
4844 | 4845 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FPeSA | 42.7334843158375 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7 | 8.86 | Baseline | Baseline | 1156162.3813 | 66074.0897 | 17.498 | 232985.7639 | 14444.376 | 16.1299 | 4.9624 | 4.5744 | 233250.0971 | 14453.9087 | 0.99124 | 0.99253 | 7.1104 | 9.019 | 1.5767 | 0.11037 | 6.9427 | 8.8942 | 7.2869 | 9.1892 | 0.34412 | 0.29502 | 0.077419 | 0.070823 | 247.9273 | 142.3241 | 0.0022506 | 1.3143e-08 | FALSE | 0.78837 | TRUE | [M-H]- | 106.8337 | 7.0294 | 7.1922 | 0.16281 | 7.0405 | 7.181 | 0.14046 | 1523.7487 | 1.0455 | 1.1047 | 1.6526 | 0 | 38 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1156162.3813 | No Control Sample | 0 | off | off | 1156162.3813 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9565 | 1.936 | 0.67363 | 77.9644 | -11.4918 | -0.89596 | 0 | 0 | 0 | 0 | 0 | 1.0429 | FALSE | 10 | 17.498
4846 | 4847 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FHxSA | 39.8446299784787 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7.86 | 8.86 | Baseline | Baseline | 537947.2783 | 66074.0897 | 8.1416 | 118104.5156 | 14444.376 | 8.1765 | 4.5548 | 4.5744 | 118155.8035 | 14453.9087 | 0.995 | 0.99253 | 7.9992 | 9.019 | 1.7714 | 0.13923 | 7.7832 | 8.8942 | 8.2437 | 9.1892 | 0.4605 | 0.29502 | 0.070029 | 0.070823 | 135.4743 | 142.3241 | 0.00088095 | 1.3143e-08 | FALSE | 0.88693 | TRUE | [M-H]- | 99.6116 | 7.9298 | 8.0819 | 0.15203 | 7.9394 | 8.0701 | 0.13069 | 225.9374 | 1.0117 | 0.99466 | 3.0263 | 0 | 34 | 8 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 537947.2783 | No Control Sample | 0 | off | off | 537947.2783 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9536 | 2.6004 | 1.0348 | 77.9645 | -12.1501 | -0.94729 | 0 | 0 | 0 | 0 | 0 | 1.1417 | FALSE | 10 | 8.1416
4848 | 4849 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FHpSA | 40.0481686554052 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 8.58 | 8.86 | Baseline | Baseline | 1119928.1106 | 66074.0897 | 16.9496 | 238600.6604 | 14444.376 | 16.5186 | 4.6937 | 4.5744 | 238649.6453 | 14453.9087 | 0.99713 | 0.99253 | 8.5786 | 9.019 | 0.016244 | 0.0013937 | 8.3256 | 8.8942 | 8.8834 | 9.1892 | 0.55777 | 0.29502 | 0.072637 | 0.070823 | 139.349 | 142.3241 | 0.00042577 | 1.3143e-08 | FALSE | 0.95117 | TRUE | [M-H]- | 100.1204 | 8.5061 | 8.6585 | 0.15241 | 8.5158 | 8.6472 | 0.13142 | 182.1363 | 0.9531 | 0.87088 | 1.5634 | 0 | 31 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1119928.1106 | No Control Sample | 0 | off | off | 1119928.1106 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9499 | 1.0603 | 0.47497 | 77.9642 | -14.7891 | -1.153 | 0 | 0 | 0 | 0 | 0 | 1.0276 | FALSE | 10 | 16.9496
4850 | 4851 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FOSA | 42.134844090691 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.88 | 8.86 | Baseline | Baseline | 633458.3299 | 66074.0897 | 9.5871 | 131959.9353 | 14444.376 | 9.1357 | 4.8004 | 4.5744 | 131994.5107 | 14453.9087 | 0.99331 | 0.99253 | 9.0214 | 9.019 | 1.5921 | 0.14138 | 8.8638 | 8.8942 | 9.2404 | 9.1892 | 0.37653 | 0.29502 | 0.074804 | 0.070823 | 113.3322 | 142.3241 | 0.00058584 | 1.3143e-08 | FALSE | 1.0003 | TRUE | [M-H]- | 105.3371 | 8.942 | 9.0998 | 0.15786 | 8.9504 | 9.0896 | 0.13923 | 205.3149 | 0.91397 | 0.78603 | 3.7757 | 0 | 20 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 633458.3299 | No Control Sample | 0 | off | off | 633458.3299 | 0 | off | 0 | off | 0 | off | 0 | off | 497.947 | 1.6076 | 0.80048 | 77.9645 | -12.037 | -0.93847 | 0 | 0 | 0 | 0 | 0 | 1.5948 | FALSE | 10 | 9.5871
4869 | 4870 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 4:2 FTS | 36.8819667811702 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 5.94 | 5.92 | Baseline | Baseline | 309854.4026 | 16546.0501 | 18.7268 | 52366.7349 | 2778.9173 | 18.8443 | 5.917 | 5.9541 | 52410.7678 | 2787.9661 | 0.9941 | 0.98363 | 5.9411 | 5.9423 | 0.018593 | 0.0011044 | 5.7415 | 5.7885 | 6.1126 | 6.106 | 0.37113 | 0.31753 | 0.091175 | 0.090993 | 216.4662 | 153.9344 | 0.0015134 | 0.0052264 | FALSE | 0.99979 | TRUE | [M-H]- | 92.2049 | 5.8387 | 6.0447 | 0.20595 | 5.8538 | 6.0319 | 0.17813 | 213.5336 | 1.0037 | 1.0354 | 3.3444 | 0 | 46 | 13 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 309854.4026 | No Control Sample | 0 | off | off | 309854.4026 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9753 | 3.1943 | 1.0445 | 306.9682 | 1.937 | 0.59459 | 0 | 0 | 0 | 0 | 0 | 5.3291 | FALSE | 10 | 18.7268
4871 | 4872 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 6:2 FTS | 38.2119248288432 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7.92 | 7.92 | Baseline | Baseline | 249156.3881 | 10746.8354 | 23.1842 | 60065.0143 | 2603.5346 | 23.0706 | 4.1481 | 4.1278 | 60097.4029 | 2610.6783 | 0.99482 | 0.98265 | 7.9323 | 7.9317 | 0.15502 | 0.012278 | 7.7858 | 7.8041 | 8.0895 | 8.0479 | 0.30373 | 0.24385 | 0.06293 | 0.06364 | 112.3889 | 92.2118 | 0.0011145 | 0.0052159 | FALSE | 1.0001 | TRUE | [M-H]- | 95.5298 | 7.8623 | 8.0039 | 0.14165 | 7.8714 | 7.9928 | 0.12144 | 220.3931 | 1.0828 | 1.1554 | 2.207 | 0 | 24 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 249156.3881 | No Control Sample | 0 | off | off | 249156.3881 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9696 | 3.9178 | 1.6728 | 406.9627 | 3.0646 | 1.2472 | 0 | 0 | 0 | 0 | 0 | 8.2742 | FALSE | 10 | 23.1842
4873 | 4874 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 8:2 FTS | 39.1058989595037 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.88 | 8.89 | Baseline | Baseline | 211614.7275 | 11028.8446 | 19.1874 | 40738.0942 | 2034.3846 | 20.0248 | 5.1945 | 5.4212 | 40743.5423 | 2034.3846 | 0.99743 | 1 | 8.8894 | 8.8903 | 0.10538 | 0.0093581 | 8.7263 | 8.6443 | 9.1845 | 9.1153 | 0.4582 | 0.47103 | 0.080742 | 0.084728 | 132.7141 | 112.4286 | 0.00036846 | 0 | FALSE | 0.99989 | TRUE | [M-H]- | 97.7647 | 8.8061 | 8.9813 | 0.1752 | 8.8143 | 8.9673 | 0.15302 | 32.7592 | 1.0733 | 1.0838 | 1.7645 | 0 | 24 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 211614.7275 | No Control Sample | 0 | off | off | 211614.7275 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9639 | 4.5045 | 2.3737 | 506.9564 | 1.4353 | 0.72765 | 0 | 0 | 0 | 0 | 0 | 5.7212 | FALSE | 10 | 19.1874
4875 | 4876 | PFAS CS8 40ng/ml | 35 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 36) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 10:2 FTS | 39.4081311928923 | Quantifiers | 10/27/2023 04:32:10 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.5 | 8.89 | Baseline | Baseline | 221997.915 | 11028.8446 | 20.1288 | 50983.9265 | 2034.3846 | 25.0611 | 4.3543 | 5.4212 | 51002.5453 | 2034.3846 | 0.99578 | 1 | 9.5086 | 8.8903 | 0.090611 | 0.008608 | 9.3568 | 8.6443 | 9.6975 | 9.1153 | 0.34067 | 0.47103 | 0.067662 | 0.084728 | 361.4884 | 112.4286 | 0.00084496 | 0 | FALSE | 1.0695 | TRUE | [M-H]- | 98.5203 | 9.4394 | 9.5813 | 0.14189 | 9.4464 | 9.5707 | 0.12436 | 126.456 | 1.1298 | 1.2296 | 2.6197 | 0 | 24 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 221997.915 | No Control Sample | 0 | off | off | 221997.915 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9565 | 2.2427 | 1.4061 | 606.9497 | 2.9328 | 1.78 | 0 | 0 | 0 | 0 | 0 | 2.1338 | FALSE | 10 | 20.1288
4936 | 4937 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFBA | 39.1058423373554 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 1.49 | 1.48 | Baseline | Baseline | 748688.065 | 106732.7587 | 7.0146 | 94775.4344 | 13795.1571 | 6.8702 | 7.8996 | 7.737 | 94937.6215 | 13835.2154 | 0.98944 | 0.98587 | 1.4654 | 1.4687 | 1.6537 | 0.02464 | 1.3133 | 1.3146 | 1.7238 | 1.7006 | 0.41047 | 0.38603 | 0.11983 | 0.11729 | 250.2385 | 201.3504 | 0.0040163 | 0.0072158 | FALSE | 0.99771 | TRUE | [M-H]- | 97.7646 | 1.3714 | 1.6456 | 0.27412 | 1.3813 | 1.6142 | 0.23283 | 927.3535 | 1.8687 | 2.6692 | 1.4137 | 0 | 83 | 26 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 748688.065 | No Control Sample | 0 | off | off | 748688.065 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9804 | 5.7246 | 1.2192 | 168.9886 | -2.5465 | -0.43033 | 0 | 0 | 0 | 0 | 0 | 0.56334 | FALSE | 10 | 7.0146
4938 | 4939 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFPeA | 36.2090807854418 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 3.18 | 3.21 | Baseline | Manual | 646390.5514 | 79823.2614 | 8.0978 | 62930.9017 | 7659.9723 | 8.2156 | 10.2714 | 10.4208 | 63121.352 | 7673.512 | 0.98551 | 0.99047 | 3.2183 | 3.2177 | 1.2044 | 0.0383 | 2.9405 | 2.9788 | 3.5243 | 3.5404 | 0.58383 | 0.56152 | 0.16308 | 0.15744 | 1035.3686 | 600.1972 | 0.0057708 | 0.0024285 | FALSE | 1.0002 | TRUE | [M-H]- | 90.5227 | 3.0619 | 3.4153 | 0.3534 | 3.0865 | 3.3911 | 0.30464 | 622.0345 | 1.1464 | 1.351 | 1.8768 | 0 | 103 | 31 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 646390.5514 | No Control Sample | 0 | off | off | 646390.5514 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9771 | 4.0605 | 1.0678 | 218.9858 | -0.079076 | -0.017317 | 0 | 0 | 0 | 0 | 0 | 0.51472 | FALSE | 10 | 8.0978
4940 | 4941 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFHxA | 38.1764047836239 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 623.4996 | 6.13 | 6.12 | Baseline | Baseline | 1270963.8428 | 143050.7468 | 8.8847 | 211265.5228 | 23332.1484 | 9.0547 | 6.016 | 6.1311 | 211363.75 | 23358.6023 | 0.99567 | 0.99332 | 6.1445 | 6.1412 | 0.23654 | 0.0145 | 5.9453 | 5.9562 | 6.381 | 6.3287 | 0.43572 | 0.37247 | 0.096408 | 0.097013 | 275.1855 | 198.6949 | 0.00095533 | 0.0022547 | FALSE | 1.0005 | TRUE | [M-H]- | 95.441 | 6.0451 | 6.2388 | 0.19368 | 6.0591 | 6.226 | 0.16688 | 463.2106 | 0.97736 | 0.96214 | 1.2471 | 0 | 48 | 13 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1270963.8428 | No Control Sample | 203843.5852 | off | off | 1270963.8428 | 203843.5852 | off | 0.025 | off | 40 | off | 249.3998 | off | 312.9736 | 2.4942 | 0.78062 | 268.9816 | -2.8715 | -0.77237 | 0 | 0 | 0 | 0 | 0 | 3.9144 | FALSE | 10 | 8.8847
4942 | 4943 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFHpA | 37.5926461182869 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 592.9053 | 7.26 | 7.27 | Baseline | Baseline | 1353773.2674 | 143917.7947 | 9.4066 | 261951.5787 | 33359.8563 | 7.8523 | 5.168 | 4.3141 | 262072.1713 | 33377.8686 | 0.99636 | 0.99461 | 7.2917 | 7.2909 | 0.43622 | 0.03167 | 7.1409 | 7.1527 | 7.4825 | 7.4357 | 0.34158 | 0.283 | 0.082523 | 0.067218 | 172.97 | 232.7237 | 0.00093113 | 0.0010547 | FALSE | 1.0001 | TRUE | [M-H]- | 93.9816 | 7.2153 | 7.3674 | 0.1521 | 7.2243 | 7.3587 | 0.1344 | 714.0567 | 0.97414 | 0.9455 | 1.1563 | 0 | 41 | 12 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Pass | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1353773.2674 | No Control Sample | 228328.7703 | off | off | 1353773.2674 | 228328.7703 | off | 0.025 | off | 40 | off | 237.1621 | off | 362.9706 | 2.5733 | 0.93403 | 318.9785 | -1.9226 | -0.61328 | 0 | 0 | 0 | 0 | 0 | 1.1865 | FALSE | 10 | 9.4066
4944 | 4945 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFOA | 37.2334142482271 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 576.9343 | 7.96 | 7.94 | Baseline | Baseline | 1104687.3124 | 122336.9886 | 9.0299 | 239369.0203 | 27776.4455 | 8.6177 | 4.615 | 4.4043 | 239501.5038 | 27787.8015 | 0.99582 | 0.99587 | 7.9781 | 7.9768 | 0.22743 | 0.018103 | 7.8468 | 7.8396 | 8.3049 | 8.169 | 0.45817 | 0.32942 | 0.0721 | 0.067714 | 123.5856 | 132.0841 | 0.00030924 | 0.00096215 | FALSE | 1.0002 | TRUE | [M-H]- | 93.0835 | 7.9113 | 8.0631 | 0.15171 | 7.9186 | 8.0504 | 0.1318 | -161.5604 | 1.1092 | 1.1568 | 0.76485 | 0 | 33 | 8 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Pass | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1104687.3124 | No Control Sample | 191475.4056 | off | off | 1104687.3124 | 191475.4056 | off | 0.025 | off | 40 | off | 230.7737 | off | 412.9676 | 2.8265 | 1.1672 | 368.9752 | -2.3104 | -0.85247 | 0 | 0 | 0 | 0 | 0 | 1.6347 | FALSE | 10 | 9.0299
4946 | 4947 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFNA | 39.3295621750624 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 614.4296 | 8.46 | 8.47 | Baseline | Baseline | 977824.04 | 107023.2864 | 9.1366 | 184858.1308 | 21340.7629 | 8.6622 | 5.2896 | 5.015 | 184890.3637 | 21348.8894 | 0.99806 | 0.99665 | 8.4922 | 8.4925 | 0.38067 | 0.032205 | 8.3232 | 8.3498 | 8.7568 | 8.746 | 0.43357 | 0.39618 | 0.083585 | 0.079062 | 156.5934 | 160.9131 | 0.00041773 | 0.00017251 | FALSE | 0.99997 | TRUE | [M-H]- | 98.3239 | 8.4038 | 8.5751 | 0.17129 | 8.417 | 8.564 | 0.14708 | 178.1065 | 0.89183 | 0.77465 | 1.0245 | 0 | 25 | 6 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Pass | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 977824.04 | No Control Sample | 159143.3885 | off | off | 977824.04 | 159143.3885 | off | 0.025 | off | 40 | off | 245.7718 | off | 462.9647 | 3.2176 | 1.4896 | 418.9725 | -1.6436 | -0.68862 | 0 | 0 | 0 | 0 | 0 | 1.5949 | FALSE | 10 | 9.1366
4948 | 4949 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFDA | 39.8291676430167 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.87 | 8.86 | Baseline | Baseline | 826392.3696 | 83580.5698 | 9.8874 | 161184.128 | 16163.5345 | 9.9721 | 5.127 | 5.1709 | 161203.4483 | 16171.2824 | 0.99867 | 0.99642 | 8.9046 | 8.9027 | 0.39018 | 0.034609 | 8.7205 | 8.7105 | 9.2174 | 9.1081 | 0.49697 | 0.3976 | 0.080042 | 0.080594 | 136.5011 | 113.8539 | 0.00020486 | 0.00093832 | FALSE | 1.0002 | TRUE | [M-H]- | 99.5729 | 8.8206 | 8.9908 | 0.17014 | 8.8292 | 8.9798 | 0.15061 | 66.4429 | 1.0508 | 1.0812 | 0.56426 | 0 | 26 | 6 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 826392.3696 | No Control Sample | 0 | off | off | 826392.3696 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9623 | 4.3214 | 2.2167 | 468.9691 | -1.9337 | -0.90685 | 0 | 0 | 0 | 0 | 0 | 2.5146 | FALSE | 10 | 9.8874
4950 | 4951 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFUdA | 38.3094112513344 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.21 | 9.2 | Baseline | Baseline | 650283.4962 | 76542.2185 | 8.4957 | 136048.8901 | 16125.3614 | 8.437 | 4.7798 | 4.7467 | 136104.764 | 16132.7183 | 0.99674 | 0.99551 | 9.2383 | 9.2377 | 0.30721 | 0.028294 | 9.082 | 9.0902 | 9.4019 | 9.3917 | 0.31995 | 0.30148 | 0.075191 | 0.074278 | 200.8031 | 156.4543 | 0.00073376 | 0.00086702 | FALSE | 1.0001 | TRUE | [M-H]- | 95.7735 | 9.162 | 9.319 | 0.15698 | 9.1702 | 9.3102 | 0.13994 | 312.0076 | 1.036 | 1.0716 | 0.56364 | 0 | 18 | 5 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 650283.4962 | No Control Sample | 0 | off | off | 650283.4962 | 0 | off | 0 | off | 0 | off | 0 | off | 562.959 | 3.8523 | 2.1687 | 518.9669 | -0.77509 | -0.40225 | 0 | 0 | 0 | 0 | 0 | 0.73928 | FALSE | 10 | 8.4957
4952 | 4953 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFDoA | 40.3595734858134 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.47 | 9.47 | Baseline | Baseline | 637986.8668 | 61085.4431 | 10.4442 | 149555.4512 | 14315.0199 | 10.4474 | 4.2659 | 4.2672 | 149572.2205 | 14317.6186 | 0.99825 | 0.99177 | 9.5028 | 9.5048 | 0.34648 | 0.032812 | 9.3525 | 9.3758 | 9.7518 | 9.6327 | 0.39935 | 0.25693 | 0.067022 | 0.067786 | 338.065 | 261.6854 | 0.00030045 | 0.0003072 | FALSE | 0.99979 | TRUE | [M-H]- | 100.8989 | 9.4356 | 9.576 | 0.14036 | 9.4417 | 9.5665 | 0.12479 | 112.5182 | 1.0963 | 1.1556 | 1.3799 | 0 | 29 | 6 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Pass | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 637986.8668 | No Control Sample | 0 | off | off | 637986.8668 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9542 | 0.82466 | 0.50548 | 568.9638 | 0.63709 | 0.36248 | 0 | 0 | 0 | 0 | 0 | 9.0496 | FALSE | 10 | 10.4442
4954 | 4955 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFTrDA | 42.9033111341219 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.68 | 9.47 | Baseline | Baseline | 494357.6311 | 61085.4431 | 8.0929 | 129202.7106 | 14315.0199 | 9.0257 | 3.8262 | 4.2672 | 129239.5153 | 14317.6186 | 0.99673 | 0.99177 | 9.7174 | 9.5048 | 0.38655 | 0.037418 | 9.5855 | 9.3758 | 9.9117 | 9.6327 | 0.32622 | 0.25693 | 0.060651 | 0.067786 | 491.7158 | 261.6854 | 0.00045828 | 0.0003072 | FALSE | 1.0224 | TRUE | [M-H]- | 107.2583 | 9.6584 | 9.7814 | 0.12297 | 9.6641 | 9.7731 | 0.109 | 181.5087 | 1.0232 | 1.0038 | 0.99798 | 0 | 30 | 7 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 494357.6311 | No Control Sample | 0 | off | off | 494357.6311 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9512 | 1.1681 | 0.77437 | 618.9599 | -2.8729 | -1.7782 | 0 | 0 | 0 | 0 | 0 | 2.6405 | FALSE | 10 | 8.0929
4956 | 4957 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFTeDA | 40.7487412046589 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 9.87 | 9.86 | Baseline | Baseline | 294774.9623 | 20403.4113 | 14.4473 | 84091.8298 | 5711.249 | 14.7239 | 3.5054 | 3.5725 | 84199.4504 | 5719.396 | 0.99107 | 0.98962 | 9.8935 | 9.8945 | 0.23779 | 0.02347 | 9.7829 | 9.8051 | 9.9928 | 9.9858 | 0.2099 | 0.18072 | 0.055412 | 0.056206 | 887.9105 | 515.5166 | 0.0023378 | 0.0027344 | FALSE | 0.9999 | TRUE | [M-H]- | 101.8719 | 9.8435 | 9.9553 | 0.11188 | 9.8488 | 9.948 | 0.099199 | 936.5697 | 1.129 | 1.2452 | 1.3002 | 0 | 23 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 294774.9623 | No Control Sample | 0 | off | off | 294774.9623 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9485 | 1.778 | 1.2676 | 668.9581 | 0.29712 | 0.19876 | 0 | 0 | 0 | 0 | 0 | 7.4822 | FALSE | 10 | 14.4473
4963 | 4964 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFBS | 36.7155079728007 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 3.5 | 3.5 | Baseline | Manual | 431722.5884 | 44424.3711 | 9.7181 | 38208.1849 | 3892.0197 | 9.8171 | 11.2992 | 11.4142 | 38349.2023 | 3899.5924 | 0.98583 | 0.99087 | 3.8669 | 3.864 | 10.4816 | 0.36685 | 3.5383 | 3.6203 | 4.1955 | 4.1751 | 0.65727 | 0.55475 | 0.17195 | 0.17453 | 605.9903 | 775.0745 | 0.0068801 | 0.0012397 | FALSE | 1.0007 | TRUE | [M-H]- | 91.7888 | 3.6821 | 4.0849 | 0.4028 | 3.7084 | 4.0519 | 0.34352 | 399.9553 | 1.0706 | 1.1227 | 7.6945 | 0 | 102 | 29 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 431722.5884 | No Control Sample | 0 | off | off | 431722.5884 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9436 | 1.9753 | 0.5905 | 79.9565 | -8.5193 | -0.68118 | 0 | 0 | 0 | 0 | 0 | 1.2784 | FALSE | 10 | 9.7181
4965 | 4966 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFPeS | 39.3384966038189 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 6.38 | 3.5 | Baseline | Manual | 465750.3459 | 44424.3711 | 10.4841 | 89937.2387 | 3892.0197 | 23.1081 | 5.1786 | 11.4142 | 89977.7492 | 3899.5924 | 0.9948 | 0.99087 | 6.3947 | 3.864 | 0.23075 | 0.014722 | 6.2104 | 3.6203 | 6.6387 | 4.1751 | 0.42832 | 0.55475 | 0.080819 | 0.17453 | 320.9318 | 775.0745 | 0.0010373 | 0.0012397 | FALSE | 1.6549 | TRUE | [M-H]- | 98.3462 | 6.3118 | 6.4819 | 0.17003 | 6.3225 | 6.4708 | 0.14836 | 217.8056 | 1.0738 | 1.1648 | 4.6227 | 0 | 48 | 10 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 465750.3459 | No Control Sample | 0 | off | off | 465750.3459 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9401 | 0.98933 | 0.34522 | 79.9564 | -9.8255 | -0.78562 | 0 | 0 | 0 | 0 | 0 | 1.3329 | FALSE | 10 | 10.4841
4967 | 4968 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | L-PFHxS | 32.8186870055146 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 32.44 | 4 | 8.11 | 0 | 7.34 | 7.35 | Baseline | Baseline | 347478.7177 | 56681.7716 | 6.1303 | 84576.3653 | 13724.2243 | 6.1626 | 4.1085 | 4.1301 | 85489.0139 | 13737.0751 | 0.76255 | 0.98938 | 7.3667 | 7.3661 | 0.36357 | 0.026686 | 7.2786 | 7.2273 | 7.5032 | 7.5513 | 0.22462 | 0.32402 | 0.064321 | 0.064018 | 106.9742 | 149.5302 | 0.014468 | 0.0011773 | FALSE | 1.0001 | TRUE | [M-H]- | 101.1673 | 7.3037 | 7.4361 | 0.1324 | 7.3108 | 7.4263 | 0.11551 | -5447.8912 | 1.0097 | 0.97582 | 6.4894 | 0 | 28 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 347478.7177 | No Control Sample | 0 | off | off | 347478.7177 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9371 | 1.2717 | 0.50734 | 79.9565 | -9.2745 | -0.74156 | 0 | 0 | 0 | 0 | 0 | 1.4686 | FALSE | 8.11 | 6.1303
4969 | 4970 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | Br-PFHxS | 6.2863704483739 | Qualifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 7.56 | 4 | 1.89 | 0 | 7.16 | 7.35 | Manual | Baseline | 90429.6348 | 56681.7716 | 1.5954 | 16228.8235 | 13724.2243 | 1.1825 | 5.5722 | 4.1301 | 85489.0139 | 13737.0751 | 0.19854 | 0.98938 | 7.1812 | 7.3661 | 0.29603 | 0.021196 | 7.0808 | 7.2273 | 7.2786 | 7.5513 | 0.19776 | 0.32402 | 0.088605 | 0.064018 | 20.5266 | 149.5302 | 0.018091 | 0.0011773 | TRUE | 0.97489 | TRUE | [M-H]- | 83.153 | 7.0984 | 7.2786 | 0.18018 | 7.1062 | 7.2615 | 0.15523 | 1484.5846 | 0.99991 | 0.88642 | 0.26025 | 0 | 23 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 90429.6348 | No Control Sample | 0 | off | off | 347478.7177 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9371 | 1.2717 | 0.50734 | 79.9488 | -105.059 | -8.4002 | 0 | 0 | 0 | 0 | 0 | 1.4686 | FALSE | 1.89 | 1.5954
4971 | 4972 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFHpS | 38.8541757761544 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7.98 | 7.35 | Baseline | Baseline | 335180.1567 | 56681.7716 | 5.9134 | 73661.9751 | 13724.2243 | 5.3673 | 4.5502 | 4.1301 | 73732.3605 | 13737.0751 | 0.99237 | 0.98938 | 8.0039 | 7.3661 | 0.29953 | 0.023903 | 7.8895 | 7.2273 | 8.2765 | 7.5513 | 0.38697 | 0.32402 | 0.071508 | 0.064018 | 138.0922 | 149.5302 | 0.00040462 | 0.0011773 | FALSE | 1.0866 | TRUE | [M-H]- | 97.1354 | 7.9378 | 8.0849 | 0.14705 | 7.9457 | 8.0741 | 0.1284 | -77.0228 | 1.063 | 1.0938 | 5.7226 | 0 | 28 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 335180.1567 | No Control Sample | 0 | off | off | 335180.1567 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9344 | 2.2646 | 1.0167 | 79.9565 | -8.2734 | -0.66152 | 0 | 0 | 0 | 0 | 0 | 1.0864 | FALSE | 10 | 5.9134
4973 | 4974 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | L-PFOS | 30.3617818626371 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 31.52 | 4 | 7.88 | 0 | 8.46 | 8.45 | Manual | Baseline | 258470.2058 | 32751.252 | 7.8919 | 52149.9127 | 6645.5946 | 7.8473 | 4.9563 | 4.9283 | 52684.8249 | 6649.6755 | 0.64003 | 0.99144 | 8.4903 | 8.4913 | 0.3587 | 0.030346 | 8.4092 | 8.3528 | 8.5935 | 8.6579 | 0.18432 | 0.30512 | 0.077472 | 0.077748 | 44.8772 | 153.7285 | 0.0093977 | 0.00086866 | TRUE | 0.99988 | TRUE | [M-H]- | 96.3255 | 8.4001 | 8.5678 | 0.16774 | 8.4001 | 8.558 | 0.15785 | -2658.9427 | 0.99353 | 0.86991 | 10.0678 | 0 | 12 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 258470.2058 | No Control Sample | 0 | off | off | 258470.2058 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9306 | 0.7526 | 0.3755 | 79.9564 | -10.8337 | -0.86623 | 0 | 0 | 0 | 0 | 0 | 1.5993 | FALSE | 7.88 | 7.8919
4975 | 4976 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | Br-PFOS | 6.92672355081363 | Qualifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 8.44 | 4 | 2.11 | 0 | 8.3 | 8.45 | Manual | Baseline | 126238.8776 | 32751.252 | 3.8545 | 21279.6149 | 6645.5946 | 3.2021 | 5.9324 | 4.9283 | 52684.8249 | 6649.6755 | 0.31278 | 0.99144 | 8.3432 | 8.4913 | 0.52082 | 0.043228 | 8.23 | 8.3528 | 8.4126 | 8.6579 | 0.18261 | 0.30512 | 0.099548 | 0.077748 | 18.312 | 153.7285 | 0.013754 | 0.00086866 | TRUE | 0.98256 | TRUE | [M-H]- | 82.0702 | 8.2521 | 8.4212 | 0.16908 | 8.2624 | 8.4212 | 0.15879 | 1602.7801 | 0.90988 | 0.9218 | 0.48841 | 0 | 12 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 126238.8776 | No Control Sample | 0 | off | off | 258470.2058 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9306 | 0.7526 | 0.3755 | 79.9581 | 10.4783 | 0.83782 | 0 | 0 | 0 | 0 | 0 | 1.5993 | FALSE | 2.11 | 3.8545
4977 | 4978 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFNS | 42.19179079785 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.85 | 8.45 | Baseline | Baseline | 306278.136 | 32751.252 | 9.3516 | 63172.558 | 6645.5946 | 9.5059 | 4.8483 | 4.9283 | 63201.6345 | 6649.6755 | 0.99292 | 0.99144 | 8.8857 | 8.4913 | 0.40361 | 0.03572 | 8.7433 | 8.3528 | 9.1042 | 8.6579 | 0.36088 | 0.30512 | 0.07555 | 0.077748 | 169.4667 | 153.7285 | 0.00055507 | 0.00086866 | FALSE | 1.0464 | TRUE | [M-H]- | 105.4795 | 8.8074 | 8.9694 | 0.16201 | 8.8172 | 8.9605 | 0.14323 | 97.1645 | 1.0367 | 1.0982 | 6.3356 | 0 | 19 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 306278.136 | No Control Sample | 0 | off | off | 306278.136 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9276 | 0.99665 | 0.54709 | 79.9566 | -5.8834 | -0.47042 | 0 | 0 | 0 | 0 | 0 | 1.5779 | FALSE | 10 | 9.3516
4979 | 4980 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | PFDS | 41.5869840331767 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.19 | 8.45 | Baseline | Baseline | 310176.9793 | 32751.252 | 9.4707 | 65018.738 | 6645.5946 | 9.7837 | 4.7706 | 4.9283 | 65034.1497 | 6649.6755 | 0.99438 | 0.99144 | 9.2142 | 8.4913 | 0.26331 | 0.024199 | 9.0678 | 8.3528 | 9.4383 | 8.6579 | 0.3705 | 0.30512 | 0.07435 | 0.077748 | 202.4267 | 153.7285 | 0.00055769 | 0.00086866 | FALSE | 1.0851 | TRUE | [M-H]- | 103.9675 | 9.1412 | 9.2965 | 0.15524 | 9.1491 | 9.2847 | 0.13565 | 97.8685 | 0.95328 | 0.84307 | 7.1814 | 0 | 21 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 310176.9793 | No Control Sample | 0 | off | off | 310176.9793 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9241 | 0.45897 | 0.27489 | 79.9566 | -4.2991 | -0.34375 | 0 | 0 | 0 | 0 | 0 | 2.5788 | FALSE | 10 | 9.4707
4983 | 4984 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FBSA | 34.4699677676117 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 4.6 | 8.86 | Baseline | Baseline | 620508.5165 | 68270.6224 | 9.089 | 53269.0321 | 14470.4775 | 3.6812 | 11.6486 | 4.7179 | 53392.2639 | 14473.6242 | 0.99133 | 0.99692 | 4.9885 | 9.0297 | 8.4451 | 0.38847 | 4.5962 | 8.8318 | 5.3182 | 9.2283 | 0.72197 | 0.39648 | 0.1811 | 0.072364 | 498.4094 | 108.2522 | 0.004287 | 0.00040943 | FALSE | 0.55245 | TRUE | [M-H]- | 86.1749 | 4.79 | 5.201 | 0.41093 | 4.8184 | 5.1711 | 0.35271 | 316.3083 | 1.0527 | 1.1145 | 4.2352 | 0 | 130 | 35 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 620508.5165 | No Control Sample | 0 | off | off | 620508.5165 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9598 | 2.6687 | 0.79517 | 77.9648 | -6.0898 | -0.47479 | 0 | 0 | 0 | 0 | 0 | 0.47089 | FALSE | 10 | 9.089
4985 | 4986 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FPeSA | 41.6992637512528 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7 | 8.86 | Baseline | Baseline | 1165686.0264 | 68270.6224 | 17.0745 | 228184.7023 | 14470.4775 | 15.769 | 5.1085 | 4.7179 | 228353.1662 | 14473.6242 | 0.99394 | 0.99692 | 7.1189 | 9.0297 | 1.6992 | 0.11894 | 6.9507 | 8.8318 | 7.3273 | 9.2283 | 0.37653 | 0.39648 | 0.079693 | 0.072364 | 260.6261 | 108.2522 | 0.0015609 | 0.00040943 | FALSE | 0.78839 | TRUE | [M-H]- | 104.2482 | 7.0372 | 7.2023 | 0.16506 | 7.0489 | 7.1909 | 0.14197 | 945.9426 | 0.95645 | 0.90335 | 1.9655 | 0 | 42 | 10 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1165686.0264 | No Control Sample | 0 | off | off | 1165686.0264 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9564 | 1.7317 | 0.60257 | 77.9641 | -14.8479 | -1.1576 | 0 | 0 | 0 | 0 | 0 | 1.6377 | FALSE | 10 | 17.0745
4987 | 4988 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FHxSA | 38.6561110478082 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7.86 | 8.86 | Baseline | Baseline | 539250.741 | 68270.6224 | 7.8987 | 115502.1028 | 14470.4775 | 7.9819 | 4.6688 | 4.7179 | 115560.9359 | 14473.6242 | 0.99448 | 0.99692 | 8.009 | 9.0297 | 1.896 | 0.14902 | 7.8745 | 8.8318 | 8.2748 | 9.2283 | 0.40023 | 0.39648 | 0.072818 | 0.072364 | 142.5509 | 108.2522 | 0.00094685 | 0.00040943 | FALSE | 0.88696 | TRUE | [M-H]- | 96.6403 | 7.9392 | 8.0927 | 0.15351 | 7.9487 | 8.0816 | 0.13294 | 273.248 | 1.1779 | 1.388 | 2.8424 | 0 | 29 | 7 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 539250.741 | No Control Sample | 0 | off | off | 539250.741 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9539 | 3.3045 | 1.315 | 77.9646 | -10.3218 | -0.80474 | 0 | 0 | 0 | 0 | 0 | 5.3574 | FALSE | 10 | 7.8987
4989 | 4990 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FHpSA | 39.0589025773011 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.58 | 8.86 | Baseline | Baseline | 1128574.4014 | 68270.6224 | 16.5309 | 253963.1133 | 14470.4775 | 17.5504 | 4.4439 | 4.7179 | 254375.2968 | 14473.6242 | 0.98963 | 0.99692 | 8.5901 | 9.0297 | 0.11766 | 0.010095 | 8.4619 | 8.8318 | 8.9033 | 9.2283 | 0.4414 | 0.39648 | 0.069501 | 0.072364 | 117.1708 | 108.2522 | 0.0016707 | 0.00040943 | FALSE | 0.95131 | TRUE | [M-H]- | 97.6473 | 8.5192 | 8.6658 | 0.14661 | 8.5262 | 8.6569 | 0.13061 | -961.2371 | 1.1055 | 1.2055 | 1.705 | 0 | 25 | 6 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1128574.4014 | No Control Sample | 0 | off | off | 1128574.4014 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9498 | 0.91723 | 0.41087 | 77.9641 | -16.4854 | -1.2853 | 0 | 0 | 0 | 0 | 0 | 1.1908 | FALSE | 10 | 16.5309
4991 | 4992 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | FOSA | 40.0501639973065 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 8.88 | 8.86 | Baseline | Baseline | 622133.565 | 68270.6224 | 9.1128 | 134137.9537 | 14470.4775 | 9.2698 | 4.638 | 4.7179 | 134182.7862 | 14473.6242 | 0.99301 | 0.99692 | 9.0293 | 9.0297 | 1.6814 | 0.14931 | 8.8218 | 8.8318 | 9.2794 | 9.2283 | 0.45762 | 0.39648 | 0.071473 | 0.072364 | 169.5062 | 108.2522 | 0.00067931 | 0.00040943 | FALSE | 0.99995 | TRUE | [M-H]- | 100.1254 | 8.9498 | 9.1098 | 0.16001 | 8.9601 | 9.0985 | 0.13843 | 199.1212 | 1.0221 | 1.0378 | 3.9707 | 0 | 24 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 622133.565 | No Control Sample | 0 | off | off | 622133.565 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9458 | -0.83247 | -0.41453 | 77.9642 | -14.9094 | -1.1624 | 0 | 0 | 0 | 0 | 0 | 0.77821 | FALSE | 10 | 9.1128
5010 | 5011 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 4:2 FTS | 40.2632110466294 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 5.94 | 5.92 | Baseline | Baseline | 321819.4231 | 15741.8069 | 20.4436 | 51651.2413 | 2367.0615 | 21.8208 | 6.2306 | 6.6504 | 51691.9954 | 2369.9298 | 0.99414 | 0.98893 | 5.9511 | 5.9538 | 0.18729 | 0.011125 | 5.7378 | 5.7785 | 6.156 | 6.131 | 0.41817 | 0.35257 | 0.095126 | 0.10643 | 210.3327 | 141.6743 | 0.0015313 | 0.0025415 | FALSE | 0.99954 | TRUE | [M-H]- | 100.658 | 5.8477 | 6.0537 | 0.20598 | 5.8648 | 6.0417 | 0.17689 | 189.1424 | 1.037 | 1.1509 | 3.3392 | 0 | 51 | 14 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 321819.4231 | No Control Sample | 0 | off | off | 321819.4231 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9751 | 2.573 | 0.84132 | 306.9678 | 0.72609 | 0.22289 | 0 | 0 | 0 | 0 | 0 | 5.3291 | FALSE | 10 | 20.4436
5012 | 5013 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 6:2 FTS | 37.8284877724273 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 7.92 | 7.92 | Baseline | Baseline | 249669.5524 | 10878.1261 | 22.9515 | 61234.4574 | 2503.5551 | 24.459 | 4.0773 | 4.3451 | 61237.4815 | 2503.5551 | 0.99877 | 0.99575 | 7.9403 | 7.9403 | 0.25642 | 0.020308 | 7.7558 | 7.7469 | 8.2146 | 8.1887 | 0.45878 | 0.44178 | 0.061523 | 0.065739 | 112.4322 | 93.6137 | 0.0001204 | 0 | FALSE | 1 | TRUE | [M-H]- | 94.5712 | 7.8695 | 8.0123 | 0.14277 | 7.8803 | 8.0012 | 0.12094 | 16.0693 | 1.0347 | 1.076 | 2.4974 | 0 | 34 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 249669.5524 | No Control Sample | 0 | off | off | 249669.5524 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9701 | 5.0373 | 2.1507 | 406.9622 | 1.6778 | 0.68282 | 0 | 0 | 0 | 0 | 0 | 4.4703 | FALSE | 10 | 22.9515
5014 | 5015 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 8:2 FTS | 41.0435194380537 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Total Width | 0 | 0 | 40 | 4 | 10 | 0 | 8.88 | 8.89 | Baseline | Baseline | 216194.2794 | 10735.5922 | 20.1381 | 39889.7588 | 2152.2874 | 18.5337 | 5.4198 | 4.988 | 39891.9338 | 2152.2874 | 0.99912 | 1 | 8.8997 | 8.9002 | 0.22171 | 0.019688 | 8.7257 | 8.6783 | 9.243 | 9.1935 | 0.5173 | 0.51522 | 0.084252 | 0.077792 | 117.6645 | 124.9451 | 0.00015273 | 0 | FALSE | 0.99994 | TRUE | [M-H]- | 102.6088 | 8.8117 | 8.9902 | 0.17843 | 8.8252 | 8.9796 | 0.15438 | 11.777 | 0.92946 | 0.87154 | 1.5664 | 0 | 27 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Fail | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 216194.2794 | No Control Sample | 0 | off | off | 216194.2794 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9641 | 4.8271 | 2.5437 | 506.9558 | 0.25272 | 0.12812 | 0 | 0 | 0 | 0 | 0 | 7.0159 | FALSE | 10 | 20.1381
5016 | 5017 | PFAS CS8 40ng/ml | 36 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 37) | CS8 | 0 | 0 | Standard | 0 | 0 | 0 | 9 | 1 | 10:2 FTS | 42.9365070307672 | Quantifiers | 10/27/2023 04:48:50 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 40 | 4 | 10 | 0 | 9.5 | 8.89 | Baseline | Baseline | 235442.9846 | 10735.5922 | 21.9311 | 56789.3792 | 2152.2874 | 26.3856 | 4.1459 | 4.988 | 56812.592 | 2152.2874 | 0.9959 | 1 | 9.5206 | 8.9002 | 0.21729 | 0.020643 | 9.3908 | 8.6783 | 9.6456 | 9.1935 | 0.25477 | 0.51522 | 0.064124 | 0.077792 | 443.8858 | 124.9451 | 0.00080966 | 0 | FALSE | 1.0697 | TRUE | [M-H]- | 107.3413 | 9.449 | 9.5892 | 0.14014 | 9.458 | 9.5804 | 0.12238 | 180.4801 | 1.0339 | 1.0822 | 2.8425 | 0 | 18 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 235442.9846 | No Control Sample | 0 | off | off | 235442.9846 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9568 | 2.5861 | 1.6214 | 606.9497 | 2.9413 | 1.7852 | 0 | 0 | 0 | 0 | 0 | 1.8925 | FALSE | 10 | 21.9311
5077 | 5078 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFBA | 284.609465633191 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 1.49 | 1.48 | Manual | Manual | 2842848.1708 | 55685.5862 | 51.0518 | 280601.3951 | 6703.4323 | 41.8594 | 10.1313 | 8.307 | 282419.1702 | 6717.1526 | 0.98285 | 0.99143 | 1.4741 | 1.4676 | 1.0697 | 0.015939 | 1.3368 | 1.3403 | 1.7213 | 1.7032 | 0.38448 | 0.36288 | 0.17828 | 0.13348 | 97.3243 | 153.8922 | 0.0038123 | 0.00088563 | TRUE | 1.0044 | FALSE | [M-H]- | 71.1524 | 1.3597 | 1.6535 | 0.29371 | 1.3683 | 1.6367 | 0.26843 | -2782.3323 | 1.593 | 2.2092 | 1.0394 | 0 | 81 | 38 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2842848.1708 | No Control Sample | 0 | off | off | 2842848.1708 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9805 | 6.2381 | 1.3286 | 168.988 | -6.0121 | -1.016 | 0 | 0 | 0 | 0 | 0 | 0.53556 | FALSE | 100 | 51.0518
5079 | 5080 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFPeA | 357.99796780702 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 3.18 | 3.21 | Manual | Manual | 3464001.0892 | 43266.2671 | 80.0624 | 270986.0555 | 4142.4032 | 65.4176 | 12.7829 | 10.4447 | 271665.1959 | 4154.519 | 0.98974 | 0.98623 | 3.211 | 3.2042 | 0.97624 | 0.031044 | 2.9671 | 2.9828 | 3.56 | 3.533 | 0.5929 | 0.55022 | 0.20833 | 0.17207 | 809.5459 | 399.1521 | 0 | 0.003343 | TRUE | 1.0021 | FALSE | [M-H]- | 89.4995 | 3.0331 | 3.4193 | 0.38617 | 3.0555 | 3.3968 | 0.34134 | 0 | 1.0584 | 1.1318 | 1.1352 | 0 | 106 | 40 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3464001.0892 | No Control Sample | 0 | off | off | 3464001.0892 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9771 | 4.3252 | 1.1374 | 218.9846 | -5.3375 | -1.1688 | 0 | 0 | 0 | 0 | 0 | 0.53809 | FALSE | 100 | 80.0624
5081 | 5082 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFHxA | 183.724514897932 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 1474.3282 | 6.13 | 6.12 | Baseline | Baseline | 3005323.3642 | 70287.1702 | 42.7578 | 292797.6139 | 10482.9065 | 27.931 | 10.2642 | 6.7049 | 293182.8321 | 10518.3687 | 0.99138 | 0.98488 | 6.1561 | 6.1521 | 0.42651 | 0.026145 | 5.9605 | 5.9713 | 6.4271 | 6.3205 | 0.46667 | 0.34918 | 0.16455 | 0.10795 | 109.3819 | 196.055 | 0.0030757 | 0.0063387 | FALSE | 1.0007 | FALSE | [M-H]- | 45.9311 | 6.0279 | 6.2868 | 0.25882 | 6.0393 | 6.2713 | 0.23201 | 1929.7597 | 1.1718 | 1.3415 | 1.3077 | 0 | 52 | 20 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3005323.3642 | No Control Sample | 203843.5852 | off | off | 3005323.3642 | 203843.5852 | off | 0.0025 | off | 400 | off | 5897.3126 | off | 312.9732 | 1.2926 | 0.40456 | 268.9808 | -5.8607 | -1.5764 | 0 | 0 | 0 | 0 | 0 | 1.1962 | FALSE | 100 | 42.7578
5083 | 5084 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFHpA | 146.570028103369 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 972.088 | 7.26 | 7.27 | Baseline | Baseline | 2219556.5389 | 60519.1053 | 36.6753 | 289793.7889 | 13460.7643 | 21.5288 | 7.6591 | 4.496 | 290895.5825 | 13489.6683 | 0.98257 | 0.9857 | 7.2956 | 7.2959 | 0.49077 | 0.03563 | 7.146 | 7.1578 | 7.4595 | 7.4131 | 0.3135 | 0.25528 | 0.12707 | 0.07324 | 79.7195 | 210.3935 | 0.0062502 | 0.0037399 | FALSE | 0.99997 | FALSE | [M-H]- | 36.6425 | 7.2029 | 7.391 | 0.18811 | 7.2107 | 7.382 | 0.17129 | 5777.5681 | 0.80157 | 0.5646 | 1.6904 | 0 | 38 | 17 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2219556.5389 | No Control Sample | 228328.7703 | off | off | 2219556.5389 | 228328.7703 | off | 0.0025 | off | 400 | off | 3888.3519 | off | 362.9703 | 1.9031 | 0.69075 | 318.978 | -3.544 | -1.1305 | 0 | 0 | 0 | 0 | 0 | 1.1052 | FALSE | 100 | 36.6753
5085 | 5086 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFOA | 186.528490368322 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 955.1414 | 7.96 | 7.94 | Baseline | Baseline | 1828860.9397 | 40428.4221 | 45.237 | 264864.5934 | 9348.5959 | 28.332 | 6.9049 | 4.3245 | 265566.7073 | 9356.5466 | 0.98714 | 0.99016 | 7.9833 | 7.9787 | 0.29334 | 0.02335 | 7.8587 | 7.852 | 8.3683 | 8.1197 | 0.50953 | 0.26765 | 0.11118 | 0.066098 | 56.6825 | 118.5063 | 0.0018686 | 0.0018148 | FALSE | 1.0006 | FALSE | [M-H]- | 46.6321 | 7.9012 | 8.0928 | 0.19167 | 7.91 | 8.0788 | 0.16878 | -971.3061 | 1.2514 | 1.4902 | 2.0082 | 0 | 36 | 10 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1828860.9397 | No Control Sample | 191475.4056 | off | off | 1828860.9397 | 191475.4056 | off | 0.0025 | off | 400 | off | 3820.5657 | off | 412.9679 | 3.455 | 1.4268 | 368.9743 | -4.983 | -1.8386 | 0 | 0 | 0 | 0 | 0 | 0.20876 | FALSE | 100 | 45.237
5087 | 5088 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFNA | 204.830452049779 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 1153.19 | 8.46 | 8.47 | Baseline | Baseline | 1835225.5781 | 38568.3993 | 47.5837 | 251004.036 | 7216.0983 | 34.7839 | 7.3115 | 5.3448 | 251158.8638 | 7227.6554 | 0.99564 | 0.98775 | 8.4989 | 8.4948 | 0.46002 | 0.038918 | 8.3008 | 8.3601 | 8.7466 | 8.6632 | 0.44582 | 0.30317 | 0.12 | 0.084786 | 73.7776 | 124.4432 | 0.0012198 | 0.0027233 | FALSE | 1.0005 | FALSE | [M-H]- | 51.2076 | 8.392 | 8.5946 | 0.20251 | 8.4016 | 8.5855 | 0.18395 | 686.7998 | 0.86871 | 0.71894 | 1.2508 | 0 | 26 | 8 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1835225.5781 | No Control Sample | 159143.3885 | off | off | 1835225.5781 | 159143.3885 | off | 0.0025 | off | 400 | off | 4612.7598 | off | 462.9639 | 1.4824 | 0.6863 | 418.971 | -5.3573 | -2.2446 | 0 | 0 | 0 | 0 | 0 | 1.7314 | FALSE | 100 | 47.5837
5089 | 5090 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFDA | 211.233818923771 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 8.87 | 8.86 | Baseline | Baseline | 1691475.5811 | 32256.8936 | 52.4376 | 238275.1231 | 6052.1851 | 39.3701 | 7.0988 | 5.3298 | 238337.6105 | 6052.9465 | 0.99781 | 0.99734 | 8.9086 | 8.9028 | 0.43516 | 0.038599 | 8.7289 | 8.6826 | 9.261 | 9.154 | 0.53217 | 0.47137 | 0.1148 | 0.082182 | 70.5758 | 103.2885 | 0.00066767 | 0.00027178 | FALSE | 1.0007 | FALSE | [M-H]- | 52.8085 | 8.8048 | 9.011 | 0.20616 | 8.8144 | 9.0007 | 0.18623 | 298.9466 | 0.97213 | 0.93134 | 1.1164 | 0 | 28 | 7 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Fail | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1691475.5811 | No Control Sample | 0 | off | off | 1691475.5811 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9615 | 2.9267 | 1.5013 | 468.9679 | -4.4092 | -2.0678 | 0 | 0 | 0 | 0 | 0 | 0.57345 | FALSE | 100 | 52.4376
5091 | 5092 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFUdA | 268.361908569493 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.21 | 9.2 | Baseline | Baseline | 1412908.4696 | 23740.8874 | 59.5137 | 230326.2137 | 4559.0928 | 50.5202 | 6.1344 | 5.2074 | 230669.0435 | 4559.1911 | 0.99113 | 0.99658 | 9.2449 | 9.2414 | 0.37858 | 0.034867 | 9.0704 | 9.0786 | 9.4044 | 9.41 | 0.33402 | 0.3314 | 0.099879 | 0.081789 | 116.597 | 128.1293 | 0.0028362 | 4.0321e-05 | FALSE | 1.0004 | FALSE | [M-H]- | 67.0905 | 9.1523 | 9.3363 | 0.18403 | 9.1647 | 9.3287 | 0.16404 | 1955.7708 | 1.0211 | 1.1099 | 1.0457 | 0 | 19 | 7 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1412908.4696 | No Control Sample | 0 | off | off | 1412908.4696 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9592 | 4.171 | 2.3481 | 518.9652 | -3.9674 | -2.059 | 0 | 0 | 0 | 0 | 0 | 2.0413 | FALSE | 100 | 59.5137
5093 | 5094 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFDoA | 184.656782493292 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.47 | 9.47 | Baseline | Baseline | 1347582.2689 | 28200.8749 | 47.7851 | 225722.7429 | 6178.5361 | 36.5334 | 5.9701 | 4.5643 | 225818.8334 | 6197.4957 | 0.99556 | 0.95455 | 9.5071 | 9.5013 | 0.39152 | 0.037077 | 9.355 | 9.3951 | 9.728 | 9.6185 | 0.37303 | 0.22345 | 0.097125 | 0.072516 | 139.8418 | 108.0511 | 0.0010727 | 0.0055213 | FALSE | 1.0006 | FALSE | [M-H]- | 46.1642 | 9.4235 | 9.5928 | 0.16931 | 9.4303 | 9.5827 | 0.15234 | 649.0829 | 1.0918 | 1.1549 | 1.6821 | 0 | 27 | 8 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1347582.2689 | No Control Sample | 0 | off | off | 1347582.2689 | 0 | off | 0 | off | 0 | off | 0 | off | 612.954 | 0.64847 | 0.39748 | 568.9614 | -3.5731 | -2.033 | 0 | 0 | 0 | 0 | 0 | 2.3581 | FALSE | 100 | 47.7851
5095 | 5096 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFTrDA | 231.600218288897 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.68 | 9.47 | Baseline | Baseline | 1232010.0464 | 28200.8749 | 43.6869 | 214408.6104 | 6178.5361 | 34.7022 | 5.7461 | 4.5643 | 214565.1989 | 6197.4957 | 0.99307 | 0.95455 | 9.7247 | 9.5013 | 0.46226 | 0.044747 | 9.5583 | 9.3951 | 9.8987 | 9.6185 | 0.34042 | 0.22345 | 0.094063 | 0.072516 | 217.6614 | 108.0511 | 0.0014104 | 0.0055213 | FALSE | 1.0235 | FALSE | [M-H]- | 57.9001 | 9.646 | 9.8013 | 0.15531 | 9.6542 | 9.7938 | 0.13956 | 888.3142 | 0.93732 | 0.86888 | 1.0978 | 0 | 31 | 11 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1232010.0464 | No Control Sample | 0 | off | off | 1232010.0464 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9511 | 0.92038 | 0.61017 | 618.9589 | -4.6029 | -2.849 | 0 | 0 | 0 | 0 | 0 | 2.3173 | FALSE | 100 | 43.6869
5097 | 5098 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFTeDA | 294.552501190887 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.87 | 9.86 | Baseline | Baseline | 1005965.173 | 9632.6686 | 104.4327 | 210265.1216 | 2629.3414 | 79.9687 | 4.7843 | 3.6635 | 210455.8917 | 2637.6705 | 0.99258 | 0.98286 | 9.8991 | 9.8993 | 0.29462 | 0.029079 | 9.7619 | 9.7928 | 10.071 | 10.0029 | 0.30912 | 0.21013 | 0.079243 | 0.057835 | 531.7158 | 520.6812 | 0.0019666 | 0.0060444 | FALSE | 0.99998 | FALSE | [M-H]- | 73.6381 | 9.8343 | 9.9698 | 0.13549 | 9.8416 | 9.9615 | 0.11987 | 1337.7286 | 1.0382 | 1.0668 | 0.90049 | 0 | 37 | 11 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1005965.173 | No Control Sample | 0 | off | off | 1005965.173 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9478 | 0.77718 | 0.55409 | 668.9561 | -2.643 | -1.7681 | 0 | 0 | 0 | 0 | 0 | 0.37686 | FALSE | 100 | 104.4327
5104 | 5105 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFBS | 404.574563120114 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 3.5 | 3.5 | Baseline | Manual | 2670000.4046 | 24933.2453 | 107.086 | 224263.7589 | 2069.2067 | 108.3815 | 11.9056 | 12.0497 | 224612.7837 | 2074.7077 | 0.99355 | 0.98915 | 3.8497 | 3.8461 | 9.9907 | 0.34968 | 3.5065 | 3.5969 | 4.2261 | 4.1857 | 0.71968 | 0.58882 | 0.18809 | 0.1749 | 944.5861 | 435.6383 | 0.0032316 | 0 | FALSE | 1.0009 | TRUE | [M-H]- | 101.1436 | 3.6612 | 4.074 | 0.41276 | 3.6882 | 4.0412 | 0.35295 | 1007.0228 | 1.1383 | 1.2874 | 4.0543 | 0 | 113 | 32 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2670000.4046 | No Control Sample | 0 | off | off | 2670000.4046 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9435 | 1.7695 | 0.52899 | 79.9561 | -13.573 | -1.0853 | 0 | 0 | 0 | 0 | 0 | 1.3149 | FALSE | 100 | 107.086
5106 | 5107 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFPeS | 297.792803218143 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 6.38 | 3.5 | Baseline | Manual | 1978823.9039 | 24933.2453 | 79.3649 | 297871.2547 | 2069.2067 | 143.9543 | 6.6432 | 12.0497 | 298146.0509 | 2074.7077 | 0.99183 | 0.98915 | 6.4067 | 3.8461 | 0.4179 | 0.026662 | 6.2221 | 3.5969 | 6.6234 | 4.1857 | 0.40137 | 0.58882 | 0.10734 | 0.1749 | 224.1535 | 435.6383 | 0.001823 | 0 | FALSE | 1.6658 | FALSE | [M-H]- | 74.4482 | 6.3113 | 6.5083 | 0.19699 | 6.3215 | 6.4961 | 0.17459 | 1352.9395 | 0.91528 | 0.79179 | 4.5529 | 0 | 45 | 14 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1978823.9039 | No Control Sample | 0 | off | off | 1978823.9039 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9401 | 0.82598 | 0.28822 | 79.956 | -14.9959 | -1.199 | 0 | 0 | 0 | 0 | 0 | 1.4161 | FALSE | 100 | 79.3649
5108 | 5109 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | L-PFHxS | 337.73743140853 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 324.4 | 4 | 81.1 | 0 | 7.34 | 7.35 | Baseline | Baseline | 1323651.5377 | 20981.2269 | 63.0874 | 288464.9126 | 4965.741 | 58.091 | 4.5886 | 4.2252 | 293628.6228 | 4984.6529 | 0.63894 | 0.96879 | 7.372 | 7.369 | 0.43655 | 0.032043 | 7.2892 | 7.2662 | 7.5042 | 7.4794 | 0.21505 | 0.21315 | 0.073044 | 0.066207 | 68.3638 | 204.5471 | 0.019574 | 0.0071359 | FALSE | 1.0004 | TRUE | [M-H]- | 104.1114 | 7.3073 | 7.4442 | 0.13692 | 7.3144 | 7.4356 | 0.12123 | -26256.669 | 1.0722 | 1.1374 | 6.8882 | 0 | 27 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1323651.5377 | No Control Sample | 0 | off | off | 1323651.5377 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9369 | 0.69107 | 0.27569 | 79.9559 | -16.6953 | -1.3349 | 0 | 0 | 0 | 0 | 0 | 1.4416 | FALSE | 81.1 | 63.0874
5110 | 5111 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | Br-PFHxS | 123.750210699257 | Qualifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 75.6 | 4 | 18.9 | 0 | 7.16 | 7.35 | Manual | Baseline | 658937.4915 | 20981.2269 | 31.4061 | 118614.681 | 4965.741 | 23.8866 | 5.5553 | 4.2252 | 293628.6228 | 4984.6529 | 0.3182 | 0.96879 | 7.1889 | 7.369 | 0.40361 | 0.028899 | 7.0676 | 7.2662 | 7.2873 | 7.4794 | 0.21968 | 0.21315 | 0.089296 | 0.066207 | 28.1107 | 204.5471 | 0.010314 | 0.0071359 | TRUE | 0.97556 | FALSE | [M-H]- | 163.6908 | 7.1026 | 7.2892 | 0.18657 | 7.1115 | 7.2665 | 0.15498 | 5569.0942 | 1.0279 | 0.89456 | 0.49782 | 0 | 26 | 12 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 658937.4915 | No Control Sample | 0 | off | off | 1323651.5377 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9369 | 0.69107 | 0.27569 | 79.9582 | 11.906 | 0.95197 | 0 | 0 | 0 | 0 | 0 | 1.4416 | FALSE | 18.9 | 31.4061
5112 | 5113 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFHpS | 396.205379648067 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 400 | 4 | 100 | 0 | 7.98 | 7.35 | Baseline | Baseline | 1265168.6776 | 20981.2269 | 60.3 | 244584.2832 | 4965.741 | 49.2543 | 5.1727 | 4.2252 | 244892.8267 | 4984.6529 | 0.98993 | 0.96879 | 8.0093 | 7.369 | 0.36657 | 0.029252 | 7.8877 | 7.2662 | 8.2721 | 7.4794 | 0.38433 | 0.21315 | 0.08273 | 0.066207 | 127.8605 | 204.5471 | 0.00029476 | 0.0071359 | FALSE | 1.0869 | TRUE | [M-H]- | 99.0513 | 7.9382 | 8.0985 | 0.1603 | 7.947 | 8.0874 | 0.14042 | 187.5827 | 1.1988 | 1.4199 | 8.2756 | 0 | 28 | 8 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1265168.6776 | No Control Sample | 0 | off | off | 1265168.6776 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9342 | 1.7052 | 0.76553 | 79.9562 | -13.0775 | -1.0456 | 0 | 0 | 0 | 0 | 0 | 1.1539 | FALSE | 100 | 60.3
5114 | 5115 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | L-PFOS | 313.205734796443 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 315.2 | 4 | 78.8 | 0 | 8.46 | 8.45 | Manual | Baseline | 1106908.8386 | 13596.4905 | 81.4114 | 207934.8184 | 2447.7048 | 84.9509 | 5.3233 | 5.5548 | 209606.4734 | 2449.9224 | 0.51737 | 0.9626 | 8.4919 | 8.4916 | 0.37666 | 0.031865 | 8.4154 | 8.3449 | 8.6039 | 8.6503 | 0.18854 | 0.30543 | 0.084725 | 0.088071 | 35.2157 | 70.797 | 0.0050655 | 4.0906e-05 | TRUE | 1 | TRUE | [M-H]- | 99.3673 | 8.41 | 8.5763 | 0.16636 | 8.41 | 8.5657 | 0.15569 | -5586.7275 | 0.99856 | 0.86909 | 5.7506 | 0 | 12 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1106908.8386 | No Control Sample | 0 | off | off | 1106908.8386 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9302 | 0.060892 | 0.030381 | 79.956 | -16.3247 | -1.3053 | 0 | 0 | 0 | 0 | 0 | 1.0664 | FALSE | 78.8 | 81.4114
5116 | 5117 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | Br-PFOS | 122.488101011146 | Qualifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 84.4 | 4 | 21.1 | 0 | 8.3 | 8.45 | Manual | Baseline | 926740.3832 | 13596.4905 | 68.1603 | 150237.7548 | 2447.7048 | 61.379 | 6.1685 | 5.5548 | 209606.4734 | 2449.9224 | 0.43329 | 0.9626 | 8.3421 | 8.4916 | 0.5068 | 0.042064 | 8.2294 | 8.3449 | 8.4213 | 8.6503 | 0.19192 | 0.30543 | 0.10102 | 0.088071 | 25.4441 | 70.797 | 0.010781 | 4.0906e-05 | TRUE | 0.98239 | FALSE | [M-H]- | 145.1281 | 8.256 | 8.4308 | 0.17485 | 8.2628 | 8.4308 | 0.16805 | 8439.2693 | 1.0764 | 1.2582 | 0.83723 | 0 | 13 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 926740.3832 | No Control Sample | 0 | off | off | 1106908.8386 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9302 | 0.060892 | 0.030381 | 79.9566 | -8.1523 | -0.65183 | 0 | 0 | 0 | 0 | 0 | 1.0664 | FALSE | 21.1 | 68.1603
5118 | 5119 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFNS | 378.23934705013 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 400 | 4 | 100 | 0 | 8.85 | 8.45 | Baseline | Baseline | 1139865.6292 | 13596.4905 | 83.8353 | 201162.1878 | 2447.7048 | 82.184 | 5.6664 | 5.5548 | 201388.8416 | 2449.9224 | 0.99082 | 0.9626 | 8.8854 | 8.4916 | 0.3998 | 0.035382 | 8.7332 | 8.3449 | 9.0739 | 8.6503 | 0.34067 | 0.30543 | 0.088454 | 0.088071 | 90.6279 | 70.797 | 0.0016264 | 4.0906e-05 | FALSE | 1.0464 | TRUE | [M-H]- | 94.5598 | 8.8003 | 8.9821 | 0.18185 | 8.8128 | 8.9711 | 0.15825 | 960.4062 | 1.2258 | 1.5688 | 7.2787 | 0 | 18 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1139865.6292 | No Control Sample | 0 | off | off | 1139865.6292 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9271 | 0.10419 | 0.057191 | 79.956 | -13.2865 | -1.0623 | 0 | 0 | 0 | 0 | 0 | 1.4312 | FALSE | 100 | 83.8353
5120 | 5121 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFDS | 362.492653028941 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 400 | 4 | 100 | 0 | 9.19 | 8.45 | Baseline | Baseline | 1122406.7306 | 13596.4905 | 82.5512 | 220673.8001 | 2447.7048 | 90.1554 | 5.0863 | 5.5548 | 220795.7153 | 2449.9224 | 0.99225 | 0.9626 | 9.2136 | 8.4916 | 0.25648 | 0.02357 | 9.0563 | 8.3449 | 9.3921 | 8.6503 | 0.33582 | 0.30543 | 0.080222 | 0.088071 | 192.007 | 70.797 | 0.0012048 | 4.0906e-05 | FALSE | 1.085 | TRUE | [M-H]- | 90.6232 | 9.1358 | 9.303 | 0.16726 | 9.1453 | 9.2917 | 0.14645 | 791.6935 | 1.1602 | 1.3413 | 6.0939 | 0 | 19 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1122406.7306 | No Control Sample | 0 | off | off | 1122406.7306 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9241 | 0.42381 | 0.25383 | 79.9561 | -10.3734 | -0.82943 | 0 | 0 | 0 | 0 | 0 | 2.2852 | FALSE | 100 | 82.5512
5124 | 5125 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FBSA | 525.031163830049 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 4.6 | 8.86 | Manual | Baseline | 4496336.9285 | 32478.8586 | 138.4389 | 347256.6931 | 6779.0137 | 51.2253 | 12.9482 | 4.7911 | 347906.8056 | 6779.0173 | 0.99328 | 0.98399 | 4.9724 | 9.0287 | 8.096 | 0.37241 | 4.647 | 8.8416 | 5.3215 | 9.2524 | 0.67451 | 0.41073 | 0.20496 | 0.073531 | 524.5474 | 83.6982 | 0 | 1.1887e-06 | TRUE | 0.55074 | FALSE | [M-H]- | 131.2578 | 4.7649 | 5.185 | 0.4201 | 4.789 | 5.1591 | 0.37006 | 0 | 0.88379 | 0.73238 | 2.2606 | 0 | 125 | 40 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4496336.9285 | No Control Sample | 0 | off | off | 4496336.9285 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9594 | 1.4033 | 0.41811 | 77.9642 | -14.3909 | -1.122 | 0 | 0 | 0 | 0 | 0 | 1.138 | FALSE | 100 | 138.4389
5126 | 5127 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FPeSA | 229.181454770384 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 7 | 8.86 | Baseline | Baseline | 3047891.9783 | 32478.8586 | 93.8423 | 374267.0169 | 6779.0137 | 55.2097 | 8.1436 | 4.7911 | 374508.9175 | 6779.0173 | 0.99255 | 0.98399 | 7.1243 | 9.0287 | 1.7759 | 0.12431 | 6.9374 | 8.8416 | 7.444 | 9.2524 | 0.50663 | 0.41073 | 0.13483 | 0.073531 | 118.0924 | 83.6982 | 0.0019309 | 1.1887e-06 | FALSE | 0.78908 | FALSE | [M-H]- | 57.2954 | 7.0215 | 7.2334 | 0.21187 | 7.0311 | 7.221 | 0.1899 | 1426.4343 | 1.3257 | 1.6989 | 2.7074 | 0 | 58 | 17 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3047891.9783 | No Control Sample | 0 | off | off | 3047891.9783 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9561 | 0.95666 | 0.33288 | 77.9638 | -19.7485 | -1.5397 | 0 | 0 | 0 | 0 | 0 | 1.7595 | FALSE | 100 | 93.8423
5128 | 5129 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FHxSA | 279.824213626181 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 7.86 | 8.86 | Baseline | Baseline | 1857055.0499 | 32478.8586 | 57.1773 | 294515.5301 | 6779.0137 | 43.4452 | 6.3055 | 4.7911 | 294767.791 | 6779.0173 | 0.99288 | 0.98399 | 8.0132 | 9.0287 | 1.9489 | 0.15319 | 7.8073 | 8.8416 | 8.3172 | 9.2524 | 0.5099 | 0.41073 | 0.10144 | 0.073531 | 77.7252 | 83.6982 | 0.0018534 | 1.1887e-06 | FALSE | 0.88753 | FALSE | [M-H]- | 69.9561 | 7.9339 | 8.1108 | 0.17691 | 7.9427 | 8.0985 | 0.15583 | 1070.5076 | 1.0419 | 1.0475 | 5.7122 | 0 | 37 | 9 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1857055.0499 | No Control Sample | 0 | off | off | 1857055.0499 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9534 | 1.9478 | 0.77512 | 77.964 | -17.5521 | -1.3685 | 0 | 0 | 0 | 0 | 0 | 1.6505 | FALSE | 100 | 57.1773
5130 | 5131 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FHpSA | 167.480073337798 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 8.58 | 8.86 | Baseline | Baseline | 2302184.8755 | 32478.8586 | 70.8826 | 347715.6401 | 6779.0137 | 51.293 | 6.6209 | 4.7911 | 350263.0092 | 6779.0173 | 0.9695 | 0.98399 | 8.5972 | 9.0287 | 0.20095 | 0.017242 | 8.4707 | 8.8416 | 8.7492 | 9.2524 | 0.27853 | 0.41073 | 0.10954 | 0.073531 | 61.7708 | 83.6982 | 0.0038866 | 1.1887e-06 | FALSE | 0.95221 | FALSE | [M-H]- | 41.87 | 8.504 | 8.6904 | 0.18642 | 8.513 | 8.6806 | 0.16756 | -4852.0035 | 1.036 | 1.0707 | 3.8746 | 0 | 17 | 9 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2302184.8755 | No Control Sample | 0 | off | off | 2302184.8755 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9496 | 0.48629 | 0.21783 | 77.9638 | -20.3049 | -1.5831 | 0 | 0 | 0 | 0 | 0 | 2.1336 | FALSE | 100 | 70.8826
5132 | 5133 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FOSA | 230.355590551553 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 8.88 | 8.86 | Baseline | Baseline | 1702334.8358 | 32478.8586 | 52.4136 | 279435.5676 | 6779.0137 | 41.2207 | 6.092 | 4.7911 | 279677.6608 | 6779.0173 | 0.98889 | 0.98399 | 9.0337 | 9.0287 | 1.7312 | 0.15373 | 8.8317 | 8.8416 | 9.245 | 9.2524 | 0.41335 | 0.41073 | 0.097917 | 0.073531 | 90.3319 | 83.6982 | 0.0016946 | 1.1887e-06 | FALSE | 1.0006 | FALSE | [M-H]- | 57.5889 | 8.94 | 9.126 | 0.18605 | 8.9511 | 9.1154 | 0.16433 | 1145.5655 | 0.97729 | 0.9544 | 5.4836 | 0 | 22 | 7 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1702334.8358 | No Control Sample | 0 | off | off | 1702334.8358 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9466 | 0.75077 | 0.37384 | 77.964 | -18.5079 | -1.443 | 0 | 0 | 0 | 0 | 0 | 0.85588 | FALSE | 100 | 52.4136
5151 | 5152 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 4:2 FTS | 300.05548642829 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 5.94 | 5.92 | Baseline | Baseline | 1471946.6766 | 9661.4262 | 152.3529 | 218106.746 | 1630.0473 | 133.8039 | 6.7487 | 5.9271 | 218246.5711 | 1648.2247 | 0.9949 | 0.9659 | 5.9665 | 5.9651 | 0.44535 | 0.026454 | 5.718 | 5.8584 | 6.1684 | 6.1072 | 0.45035 | 0.24877 | 0.10601 | 0.093479 | 200.5042 | 209.5593 | 0.0011397 | 0.018246 | FALSE | 1.0002 | FALSE | [M-H]- | 75.0139 | 5.8632 | 6.0706 | 0.20744 | 5.8768 | 6.0584 | 0.1816 | 551.9811 | 1.047 | 1.1264 | 2.0413 | 0 | 56 | 14 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1471946.6766 | No Control Sample | 0 | off | off | 1471946.6766 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9748 | 1.5032 | 0.49152 | 306.9664 | -3.9158 | -1.202 | 0 | 0 | 0 | 0 | 0 | 1.0805 | FALSE | 100 | 152.3529
5153 | 5154 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 6:2 FTS | 252.880980562522 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 7.92 | 7.92 | Baseline | Baseline | 939012.2749 | 6120.1574 | 153.4294 | 205372.6383 | 1482.668 | 138.5156 | 4.5722 | 4.1278 | 205467.5607 | 1482.668 | 0.99471 | 0.99378 | 7.9448 | 7.9457 | 0.31278 | 0.024772 | 7.7673 | 7.7439 | 8.2103 | 8.1389 | 0.44305 | 0.39505 | 0.070566 | 0.062872 | 110.1791 | 104.3683 | 0.0011041 | 1.2322e-13 | FALSE | 0.99989 | FALSE | [M-H]- | 63.2202 | 7.8704 | 8.0223 | 0.15186 | 7.8811 | 8.0109 | 0.1298 | 511.7996 | 0.97867 | 0.93983 | 2.0054 | 0 | 33 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 939012.2749 | No Control Sample | 0 | off | off | 939012.2749 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9701 | 5.1556 | 2.2013 | 406.9608 | -1.7532 | -0.71348 | 0 | 0 | 0 | 0 | 0 | 2.7646 | FALSE | 100 | 153.4294
5155 | 5156 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 8:2 FTS | 271.878091099487 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 8.88 | 8.89 | Baseline | Baseline | 849455.9355 | 6367.8532 | 133.3975 | 155300.1801 | 1085.5237 | 143.0648 | 5.4698 | 5.8662 | 155310.6511 | 1089.8657 | 0.9988 | 0.97687 | 8.9009 | 8.9 | 0.23542 | 0.020906 | 8.7156 | 8.7423 | 9.2282 | 9.0282 | 0.51265 | 0.28592 | 0.085242 | 0.093465 | 123.0892 | 95.2247 | 0.00019039 | 0.0068493 | FALSE | 1.0001 | FALSE | [M-H]- | 67.9695 | 8.8124 | 8.9947 | 0.18231 | 8.8219 | 8.9808 | 0.1589 | 57.6756 | 1.0826 | 1.1279 | 1.9525 | 0 | 27 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Fail | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 849455.9355 | No Control Sample | 0 | off | off | 849455.9355 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9634 | 3.5218 | 1.8559 | 506.954 | -3.3392 | -1.6928 | 0 | 0 | 0 | 0 | 0 | 0.38739 | FALSE | 100 | 133.3975
5157 | 5158 | PFAS CS9 100ng/ml | 37 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 38) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 10:2 FTS | 222.269075732932 | Quantifiers | 10/27/2023 05:05:26 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.5 | 8.89 | Baseline | Baseline | 722944.7802 | 6367.8532 | 113.5304 | 143173.3971 | 1085.5237 | 131.8934 | 5.0494 | 5.8662 | 143214.1923 | 1089.8657 | 0.99597 | 0.97687 | 9.5246 | 8.9 | 0.25886 | 0.024592 | 9.3765 | 8.7423 | 9.7083 | 9.0282 | 0.33172 | 0.28592 | 0.080367 | 0.093465 | 254.1528 | 95.2247 | 0.00066082 | 0.0068493 | FALSE | 1.0702 | FALSE | [M-H]- | 55.5673 | 9.4452 | 9.6026 | 0.15743 | 9.4553 | 9.5942 | 0.13887 | 285.2181 | 1.0796 | 1.2119 | 2.2822 | 0 | 24 | 8 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 722944.7802 | No Control Sample | 0 | off | off | 722944.7802 | 0 | off | 0 | off | 0 | off | 0 | off | 626.956 | 1.3968 | 0.87572 | 606.9472 | -1.1954 | -0.72556 | 0 | 0 | 0 | 0 | 0 | 2.2695 | FALSE | 100 | 113.5304
5218 | 5219 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFBA | 304.231664082422 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 1.49 | 1.48 | Manual | Manual | 3129711.0023 | 57350.6433 | 54.5715 | 285843.3236 | 6659.3835 | 42.9234 | 10.949 | 8.612 | 286760.627 | 6672.5261 | 0.98902 | 0.9922 | 1.4625 | 1.4539 | 1.8471 | 0.027522 | 1.3215 | 1.3199 | 1.7195 | 1.7065 | 0.39803 | 0.38663 | 0.18747 | 0.13879 | 140.8231 | 142.0578 | 0 | 0 | TRUE | 1.0059 | FALSE | [M-H]- | 76.0579 | 1.3466 | 1.6492 | 0.30265 | 1.3552 | 1.6306 | 0.27539 | 0 | 1.3488 | 1.66 | 0.98428 | 0 | 83 | 41 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3129711.0023 | No Control Sample | 0 | off | off | 3129711.0023 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9802 | 4.4887 | 0.95601 | 168.9877 | -7.5984 | -1.2841 | 0 | 0 | 0 | 0 | 0 | 0.51639 | FALSE | 100 | 54.5715
5220 | 5221 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFPeA | 338.832652462264 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 3.18 | 3.21 | Manual | Manual | 3550338.9473 | 46852.9056 | 75.7763 | 270053.973 | 4047.7996 | 66.7162 | 13.1468 | 11.5749 | 270677.9832 | 4051.1755 | 0.99078 | 0.99369 | 3.1835 | 3.1779 | 0.10886 | 0.0034619 | 2.9437 | 2.9502 | 3.5369 | 3.4995 | 0.59319 | 0.54939 | 0.22046 | 0.18067 | 605.8863 | 326.9241 | 0.0021767 | 0 | TRUE | 1.0018 | FALSE | [M-H]- | 84.7082 | 3.0035 | 3.4041 | 0.4006 | 3.0228 | 3.3774 | 0.35456 | -990.9435 | 0.99995 | 0.95905 | 1.2236 | 0 | 107 | 43 | C5HF9O2 | 262.976 | 218.9858 | Gray | Fail | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3550338.9473 | No Control Sample | 0 | off | off | 3550338.9473 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9771 | 4.039 | 1.0622 | 218.9845 | -5.9506 | -1.3031 | 0 | 0 | 0 | 0 | 0 | 0.54626 | FALSE | 100 | 75.7763
5222 | 5223 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFHxA | 186.251029041933 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 1511.5096 | 6.13 | 6.12 | Baseline | Baseline | 3081115.4581 | 71082.2641 | 43.3458 | 300829.1183 | 10016.036 | 30.0347 | 10.2421 | 7.0968 | 301214.3755 | 10028.2346 | 0.99169 | 0.9929 | 6.1208 | 6.1184 | 0.1506 | 0.0092316 | 5.9216 | 5.9407 | 6.3944 | 6.315 | 0.47275 | 0.37433 | 0.16925 | 0.11152 | 104.875 | 134.598 | 0.0025449 | 0.0024343 | FALSE | 1.0004 | FALSE | [M-H]- | 46.5628 | 5.9944 | 6.2515 | 0.25715 | 6.005 | 6.2388 | 0.23381 | 1619.4342 | 0.99271 | 0.96676 | 1.5288 | 0 | 53 | 21 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3081115.4581 | No Control Sample | 203843.5852 | off | off | 3081115.4581 | 203843.5852 | off | 0.0025 | off | 400 | off | 6046.0386 | off | 312.9736 | 2.5275 | 0.79104 | 268.9808 | -6.0467 | -1.6265 | 0 | 0 | 0 | 0 | 0 | 1.1968 | FALSE | 100 | 43.3458
5224 | 5225 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFHpA | 135.128120654441 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 989.6491 | 7.26 | 7.27 | Baseline | Baseline | 2259653.5215 | 66829.4019 | 33.8123 | 289442.5493 | 14445.88 | 20.0363 | 7.8069 | 4.6262 | 290649.2378 | 14457.7393 | 0.98059 | 0.99224 | 7.2728 | 7.2719 | 0.17694 | 0.012846 | 7.1234 | 7.126 | 7.4234 | 7.4096 | 0.30005 | 0.28365 | 0.12902 | 0.073452 | 85.1195 | 195.5782 | 0.0068822 | 0.0016031 | FALSE | 1.0001 | FALSE | [M-H]- | 33.782 | 7.1778 | 7.3659 | 0.18809 | 7.1846 | 7.3575 | 0.1729 | 6638.8611 | 0.88607 | 0.74036 | 1.6415 | 0 | 36 | 17 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2259653.5215 | No Control Sample | 228328.7703 | off | off | 2259653.5215 | 228328.7703 | off | 0.0025 | off | 400 | off | 3958.5962 | off | 362.9698 | 0.54287 | 0.19704 | 318.9777 | -4.5002 | -1.4355 | 0 | 0 | 0 | 0 | 0 | 0.99683 | FALSE | 100 | 33.8123
5226 | 5227 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFOA | 185.143884669762 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 1009.4334 | 7.96 | 7.94 | Baseline | Baseline | 1932816.6679 | 43045.9773 | 44.9012 | 280380.4355 | 10138.3739 | 27.6554 | 6.8936 | 4.2458 | 281316.0366 | 10153.0097 | 0.98411 | 0.9873 | 7.9629 | 7.9606 | 0.036437 | 0.0029004 | 7.8296 | 7.8361 | 8.3134 | 8.1175 | 0.48378 | 0.28142 | 0.11136 | 0.063062 | 66.3555 | 91.577 | 0.0023246 | 0.0032384 | FALSE | 1.0003 | FALSE | [M-H]- | 46.286 | 7.8802 | 8.0733 | 0.19313 | 7.888 | 8.0597 | 0.17161 | -1347.2462 | 1.1534 | 1.2614 | 2.5166 | 0 | 35 | 11 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1932816.6679 | No Control Sample | 191475.4056 | off | off | 1932816.6679 | 191475.4056 | off | 0.0025 | off | 400 | off | 4037.7335 | off | 412.9668 | 0.90464 | 0.37359 | 368.9743 | -4.7877 | -1.7666 | 0 | 0 | 0 | 0 | 0 | 1.0067 | FALSE | 100 | 44.9012
5228 | 5229 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFNA | 189.280607796345 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 1176.9684 | 8.46 | 8.47 | Baseline | Baseline | 1873067.3745 | 42597.4853 | 43.9713 | 246308.3205 | 7929.1119 | 31.0638 | 7.6046 | 5.3723 | 246446.8742 | 7944.227 | 0.99559 | 0.9889 | 8.4802 | 8.471 | 0.23881 | 0.020204 | 8.2833 | 8.3069 | 8.7269 | 8.6302 | 0.4436 | 0.32333 | 0.12476 | 0.085314 | 78.4641 | 119.5709 | 0.0011866 | 0.0022027 | FALSE | 1.0011 | FALSE | [M-H]- | 47.3202 | 8.3735 | 8.5815 | 0.20801 | 8.3848 | 8.5703 | 0.18553 | 658.8454 | 0.91547 | 0.81368 | 1.4113 | 0 | 26 | 8 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1873067.3745 | No Control Sample | 159143.3885 | off | off | 1873067.3745 | 159143.3885 | off | 0.0025 | off | 400 | off | 4707.8736 | off | 462.963 | -0.52324 | -0.24224 | 418.9723 | -2.1159 | -0.88651 | 0 | 0 | 0 | 0 | 0 | inf | FALSE | 100 | 43.9713
5230 | 5231 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFDA | 197.079342087151 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 8.87 | 8.86 | Baseline | Baseline | 1670395.0135 | 34142.7376 | 48.9239 | 239607.0998 | 6556.566 | 36.5446 | 6.9714 | 5.2074 | 239670.9583 | 6571.2881 | 0.99752 | 0.98656 | 8.8913 | 8.8978 | 0.24003 | 0.021291 | 8.6914 | 8.7365 | 9.2217 | 9.04 | 0.53028 | 0.30355 | 0.11194 | 0.080787 | 70.5976 | 100.6734 | 0.00070751 | 0.0038173 | FALSE | 0.99926 | FALSE | [M-H]- | 49.2698 | 8.7915 | 8.9989 | 0.20744 | 8.8022 | 8.9881 | 0.1859 | 319.6881 | 1.0565 | 1.124 | 1.3196 | 0 | 28 | 7 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Fail | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1670395.0135 | No Control Sample | 0 | off | off | 1670395.0135 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9619 | 3.5811 | 1.837 | 468.968 | -4.1972 | -1.9684 | 0 | 0 | 0 | 0 | 0 | 1.0118 | FALSE | 100 | 48.9239
5232 | 5233 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFUdA | 265.692380037613 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.21 | 9.2 | Baseline | Baseline | 1476655.0675 | 25061.309 | 58.9217 | 232859.384 | 4820.2909 | 48.3082 | 6.3414 | 5.1991 | 233138.2506 | 4822.7901 | 0.99173 | 0.99503 | 9.2297 | 9.2278 | 0.21423 | 0.019731 | 9.0699 | 9.0781 | 9.3873 | 9.4255 | 0.31745 | 0.34745 | 0.10286 | 0.082179 | 115.2228 | 172.6938 | 0.0021549 | 0.0011195 | FALSE | 1.0002 | FALSE | [M-H]- | 66.4231 | 9.1351 | 9.3243 | 0.18927 | 9.1474 | 9.3155 | 0.16814 | 1580.7013 | 1.0818 | 1.2361 | 1.2053 | 0 | 18 | 8 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1476655.0675 | No Control Sample | 0 | off | off | 1476655.0675 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9576 | 1.27 | 0.71494 | 518.9649 | -4.5384 | -2.3553 | 0 | 0 | 0 | 0 | 0 | 0.81763 | FALSE | 100 | 58.9217
5234 | 5235 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFDoA | 189.643750121802 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.47 | 9.47 | Baseline | Baseline | 1405481.3738 | 28639.0836 | 49.0756 | 227860.6157 | 6098.3885 | 37.3641 | 6.1682 | 4.6962 | 227971.9399 | 6098.9934 | 0.99451 | 0.96635 | 9.494 | 9.4947 | 0.25361 | 0.024017 | 9.338 | 9.3445 | 9.7063 | 9.6865 | 0.36827 | 0.342 | 0.10033 | 0.074021 | 118.819 | 107.5215 | 0.0011025 | 0.00023011 | FALSE | 0.99993 | FALSE | [M-H]- | 47.4109 | 9.4096 | 9.5831 | 0.17354 | 9.4183 | 9.572 | 0.15369 | 682.1581 | 1.152 | 1.3088 | 1.767 | 0 | 26 | 9 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1405481.3738 | No Control Sample | 0 | off | off | 1405481.3738 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9543 | 1.0881 | 0.66697 | 568.9627 | -1.2238 | -0.69629 | 0 | 0 | 0 | 0 | 0 | 1.0165 | FALSE | 100 | 49.0756
5236 | 5237 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFTrDA | 230.643155785931 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.68 | 9.47 | Baseline | Baseline | 1245983.7902 | 28639.0836 | 43.5064 | 214970.16 | 6098.3885 | 35.2503 | 5.7961 | 4.6962 | 215107.3328 | 6098.9934 | 0.99256 | 0.96635 | 9.7118 | 9.4947 | 0.32858 | 0.031807 | 9.544 | 9.3445 | 9.8901 | 9.6865 | 0.34608 | 0.342 | 0.095279 | 0.074021 | 265.6883 | 107.5215 | 0.0012383 | 0.00023011 | FALSE | 1.0229 | FALSE | [M-H]- | 57.6608 | 9.6309 | 9.789 | 0.15813 | 9.638 | 9.7805 | 0.14255 | 769.189 | 1.0369 | 1.0611 | 1.3384 | 0 | 31 | 10 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1245983.7902 | No Control Sample | 0 | off | off | 1245983.7902 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9509 | 0.6564 | 0.43516 | 618.9585 | -5.1784 | -3.2052 | 0 | 0 | 0 | 0 | 0 | 2.5462 | FALSE | 100 | 43.5064
5238 | 5239 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFTeDA | 252.85739989395 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.87 | 9.86 | Baseline | Baseline | 1040611.1201 | 11607.5131 | 89.6498 | 209319.1263 | 2731.5386 | 76.6305 | 4.9714 | 4.2494 | 209531.603 | 2734.6908 | 0.99138 | 0.98944 | 9.8857 | 9.8899 | 0.15888 | 0.015682 | 9.7534 | 9.7944 | 10.0592 | 9.998 | 0.30572 | 0.20358 | 0.080136 | 0.071673 | 544.9005 | 513.051 | 0.0023176 | 0.0025436 | FALSE | 0.99957 | FALSE | [M-H]- | 63.2143 | 9.8223 | 9.9593 | 0.13701 | 9.8269 | 9.9517 | 0.12479 | 1586.794 | 1.1602 | 1.2932 | 0.94257 | 0 | 36 | 10 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1040611.1201 | No Control Sample | 0 | off | off | 1040611.1201 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9482 | 1.2715 | 0.90648 | 668.9563 | -2.3254 | -1.5556 | 0 | 0 | 0 | 0 | 0 | 1.2828 | FALSE | 100 | 89.6498
5245 | 5246 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFBS | 403.608033700442 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 3.5 | 3.5 | Baseline | Manual | 2794669.9253 | 26159.942 | 106.8301 | 236116.14 | 2132.0728 | 110.7449 | 11.836 | 12.2697 | 236443.4654 | 2139.1462 | 0.99389 | 0.98771 | 3.8135 | 3.813 | 8.9576 | 0.31352 | 3.4873 | 3.5734 | 4.2145 | 4.1384 | 0.72718 | 0.56505 | 0.18322 | 0.18641 | 792.5174 | 485.8057 | 0.0022289 | 0.00096276 | FALSE | 1.0001 | TRUE | [M-H]- | 100.902 | 3.624 | 4.0378 | 0.41389 | 3.6516 | 4.0132 | 0.36165 | 723.7251 | 1.0708 | 1.1834 | 4.3657 | 0 | 114 | 30 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2794669.9253 | No Control Sample | 0 | off | off | 2794669.9253 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9435 | 1.5774 | 0.47156 | 79.9562 | -12.7988 | -1.0234 | 0 | 0 | 0 | 0 | 0 | 0.96279 | FALSE | 100 | 106.8301
5247 | 5248 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFPeS | 304.525203832962 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 6.38 | 3.5 | Baseline | Manual | 2123118.1274 | 26159.942 | 81.1591 | 309760.7725 | 2132.0728 | 145.2862 | 6.8541 | 12.2697 | 310086.7255 | 2139.1462 | 0.99105 | 0.98771 | 6.3759 | 3.813 | 0.064126 | 0.0040912 | 6.2078 | 3.5734 | 6.5787 | 4.1384 | 0.37087 | 0.56505 | 0.11016 | 0.18641 | 219.535 | 485.8057 | 0.002263 | 0.00096276 | FALSE | 1.6722 | FALSE | [M-H]- | 76.1313 | 6.2791 | 6.4792 | 0.20006 | 6.2896 | 6.4661 | 0.17657 | 1890.1393 | 1.1267 | 1.253 | 4.1703 | 0 | 41 | 14 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2123118.1274 | No Control Sample | 0 | off | off | 2123118.1274 | 0 | off | 0 | off | 0 | off | 0 | off | 348.94 | 0.68604 | 0.23939 | 79.9559 | -16.0352 | -1.2821 | 0 | 0 | 0 | 0 | 0 | 1.3793 | FALSE | 100 | 81.1591
5249 | 5250 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | L-PFHxS | 315.415476364356 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 324.4 | 4 | 81.1 | 0 | 7.34 | 7.35 | Baseline | Baseline | 1347482.1171 | 22870.54 | 58.9178 | 293240.8059 | 5045.3522 | 58.121 | 4.5951 | 4.533 | 298439.0218 | 5061.6101 | 0.6187 | 0.98068 | 7.3466 | 7.3465 | 0.090471 | 0.0066405 | 7.261 | 7.2272 | 7.4826 | 7.4761 | 0.22158 | 0.24893 | 0.073455 | 0.070969 | 65.5673 | 138.3072 | 0.019179 | 0.0024853 | FALSE | 1 | TRUE | [M-H]- | 97.2304 | 7.282 | 7.4207 | 0.13866 | 7.289 | 7.4122 | 0.1232 | -25381.2781 | 1.1347 | 1.2754 | 6.9692 | 0 | 28 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1347482.1171 | No Control Sample | 0 | off | off | 1347482.1171 | 0 | off | 0 | off | 0 | off | 0 | off | 398.937 | 1.0158 | 0.40524 | 79.956 | -15.4018 | -1.2315 | 0 | 0 | 0 | 0 | 0 | 1.5531 | FALSE | 81.1 | 58.9178
5251 | 5252 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | Br-PFHxS | 126.262064114943 | Qualifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 75.6 | 4 | 18.9 | 0 | 7.16 | 7.35 | Manual | Baseline | 732852.7048 | 22870.54 | 32.0435 | 134759.4735 | 5045.3522 | 26.7096 | 5.4382 | 4.533 | 298439.0218 | 5061.6101 | 0.33666 | 0.98068 | 7.1637 | 7.3465 | 0.052156 | 0.0037344 | 7.0476 | 7.2272 | 7.2607 | 7.4761 | 0.21304 | 0.24893 | 0.086959 | 0.070969 | 30.1316 | 138.3072 | 0.012527 | 0.0024853 | TRUE | 0.97513 | FALSE | [M-H]- | 167.0133 | 7.079 | 7.2544 | 0.17538 | 7.0874 | 7.2409 | 0.15348 | 7924.0647 | 0.95675 | 0.84344 | 0.54387 | 0 | 25 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 732852.7048 | No Control Sample | 0 | off | off | 1347482.1171 | 0 | off | 0 | off | 0 | off | 0 | off | 398.937 | 1.0158 | 0.40524 | 79.9567 | -6.0687 | -0.48524 | 0 | 0 | 0 | 0 | 0 | 1.5531 | FALSE | 18.9 | 32.0435
5253 | 5254 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFHpS | 398.851606126846 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 400 | 4 | 100 | 0 | 7.98 | 7.35 | Baseline | Baseline | 1388305.1909 | 22870.54 | 60.7028 | 275116.541 | 5045.3522 | 54.5287 | 5.0462 | 4.533 | 275407.152 | 5061.6101 | 0.99142 | 0.98068 | 7.9887 | 7.3465 | 0.10919 | 0.0087135 | 7.8588 | 7.2272 | 8.3169 | 7.4761 | 0.4581 | 0.24893 | 0.079511 | 0.070969 | 120.2805 | 138.3072 | 0.00026612 | 0.0024853 | FALSE | 1.0874 | TRUE | [M-H]- | 99.7129 | 7.9194 | 8.0788 | 0.15939 | 7.927 | 8.0665 | 0.1395 | -159.8235 | 1.1405 | 1.2384 | 8.8229 | 0 | 33 | 8 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1388305.1909 | No Control Sample | 0 | off | off | 1388305.1909 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9334 | 0.051733 | 0.023225 | 79.9559 | -16.1707 | -1.293 | 0 | 0 | 0 | 0 | 0 | 1.8841 | FALSE | 100 | 60.7028
5255 | 5256 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | L-PFOS | 309.680751877467 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 315.2 | 4 | 78.8 | 0 | 8.46 | 8.45 | Manual | Baseline | 1155255.3317 | 14351.8684 | 80.4951 | 214142.8051 | 2417.6248 | 88.5757 | 5.3948 | 5.9364 | 216659.1476 | 2417.6293 | 0.50677 | 0.97397 | 8.4752 | 8.472 | 0.18021 | 0.015245 | 8.4004 | 8.3081 | 8.6044 | 8.6837 | 0.20405 | 0.37557 | 0.086694 | 0.09394 | 31.2101 | 69.7996 | 0.010168 | 3.3475e-06 | TRUE | 1.0004 | TRUE | [M-H]- | 98.249 | 8.3895 | 8.559 | 0.16945 | 8.3895 | 8.5494 | 0.15986 | -10671.3715 | 1.0449 | 0.97158 | 7.1949 | 0 | 13 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1155255.3317 | No Control Sample | 0 | off | off | 1155255.3317 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9309 | 1.3138 | 0.65547 | 79.956 | -15.6933 | -1.2548 | 0 | 0 | 0 | 0 | 0 | 0.77558 | FALSE | 78.8 | 80.4951
5257 | 5258 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | Br-PFOS | 125.349133150616 | Qualifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 84.4 | 4 | 21.1 | 0 | 8.3 | 8.45 | Manual | Baseline | 1001076.2089 | 14351.8684 | 69.7523 | 159601.7038 | 2417.6248 | 66.0159 | 6.2723 | 5.9364 | 216659.1476 | 2417.6293 | 0.43925 | 0.97397 | 8.3244 | 8.472 | 0.29403 | 0.024405 | 8.2073 | 8.3081 | 8.402 | 8.6837 | 0.19474 | 0.37557 | 0.1039 | 0.09394 | 23.261 | 69.7996 | 0.012341 | 3.3475e-06 | TRUE | 0.98257 | FALSE | [M-H]- | 148.5179 | 8.2338 | 8.4079 | 0.17407 | 8.2431 | 8.4079 | 0.16474 | 10114.138 | 0.86013 | 0.79341 | 0.86654 | 0 | 14 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1001076.2089 | No Control Sample | 0 | off | off | 1155255.3317 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9309 | 1.3138 | 0.65547 | 79.9563 | -13.0316 | -1.042 | 0 | 0 | 0 | 0 | 0 | 0.77558 | FALSE | 21.1 | 69.7523
5259 | 5260 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFNS | 390.259605924967 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 8.85 | 8.45 | Baseline | Baseline | 1241429.8218 | 14351.8684 | 86.4995 | 223683.7788 | 2417.6248 | 92.5221 | 5.5499 | 5.9364 | 223711.3211 | 2417.6293 | 0.99681 | 0.97397 | 8.8728 | 8.472 | 0.25796 | 0.02283 | 8.6276 | 8.3081 | 9.2066 | 8.6837 | 0.57897 | 0.37557 | 0.087705 | 0.09394 | 115.539 | 69.7996 | 5.0805e-05 | 3.3475e-06 | FALSE | 1.0473 | TRUE | [M-H]- | 97.5649 | 8.7873 | 8.9698 | 0.18246 | 8.7965 | 8.9546 | 0.15811 | -19.6286 | 1.0576 | 1.0503 | 7.1962 | 0 | 31 | 7 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Fail | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1241429.8218 | No Control Sample | 0 | off | off | 1241429.8218 | 0 | off | 0 | off | 0 | off | 0 | off | 548.927 | -0.034022 | -0.018675 | 79.9561 | -12.2299 | -0.97786 | 0 | 0 | 0 | 0 | 0 | 1.4323 | FALSE | 100 | 86.4995
5261 | 5262 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFDS | 389.278357536666 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 400 | 4 | 100 | 0 | 9.19 | 8.45 | Baseline | Baseline | 1272309.9738 | 14351.8684 | 88.6512 | 245177.1404 | 2417.6248 | 101.4124 | 5.1893 | 5.9364 | 245302.3449 | 2417.6293 | 0.99259 | 0.97397 | 9.1992 | 8.472 | 0.10008 | 0.0091974 | 9.037 | 8.3081 | 9.3917 | 8.6837 | 0.35465 | 0.37557 | 0.08099 | 0.09394 | 160.1451 | 69.7996 | 0.00091863 | 3.3475e-06 | FALSE | 1.0858 | TRUE | [M-H]- | 97.3196 | 9.1211 | 9.2893 | 0.16827 | 9.1323 | 9.2783 | 0.146 | 635.072 | 0.97392 | 0.94157 | 6.0765 | 0 | 20 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1272309.9738 | No Control Sample | 0 | off | off | 1272309.9738 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9237 | -0.29834 | -0.17868 | 79.956 | -11.7507 | -0.93955 | 0 | 0 | 0 | 0 | 0 | 0.72222 | FALSE | 100 | 88.6512
5265 | 5266 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FBSA | 507.165840146034 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 4.6 | 8.86 | Manual | Baseline | 4732731.8483 | 35390.6792 | 133.7282 | 351419.5627 | 7203.3193 | 48.7858 | 13.4675 | 4.9131 | 351844.3179 | 7206.5337 | 0.99541 | 0.98414 | 4.9387 | 9.0171 | 7.3622 | 0.33866 | 4.6067 | 8.8818 | 5.3209 | 9.2131 | 0.7142 | 0.33127 | 0.20825 | 0.076194 | 460.3968 | 100.6145 | 0.00078375 | 0.00073895 | TRUE | 0.5477 | FALSE | [M-H]- | 126.7915 | 4.723 | 5.1618 | 0.43882 | 4.7532 | 5.133 | 0.3798 | -385.642 | 1.0429 | 1.1079 | 2.503 | 0 | 132 | 40 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4732731.8483 | No Control Sample | 0 | off | off | 4732731.8483 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9596 | 2.1748 | 0.64801 | 77.9641 | -15.7289 | -1.2263 | 0 | 0 | 0 | 0 | 0 | 0.73186 | FALSE | 100 | 133.7282
5267 | 5268 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FPeSA | 215.48402467368 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 7 | 8.86 | Baseline | Baseline | 3122650.0533 | 35390.6792 | 88.2337 | 373414.1607 | 7203.3193 | 51.8392 | 8.3624 | 4.9131 | 373784.3084 | 7206.5337 | 0.99091 | 0.98414 | 7.1012 | 9.0171 | 1.4452 | 0.10116 | 6.9218 | 8.8818 | 7.4161 | 9.2131 | 0.49432 | 0.33127 | 0.13616 | 0.076194 | 104.7551 | 100.6145 | 0.0022172 | 0.00073895 | FALSE | 0.78752 | FALSE | [M-H]- | 53.871 | 6.9943 | 7.2081 | 0.21382 | 7.0042 | 7.1957 | 0.19154 | 1674.9286 | 0.88533 | 0.72703 | 3.1216 | 0 | 56 | 17 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3122650.0533 | No Control Sample | 0 | off | off | 3122650.0533 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9558 | 0.1737 | 0.060439 | 77.9637 | -20.7023 | -1.6141 | 0 | 0 | 0 | 0 | 0 | 0.96856 | FALSE | 100 | 88.2337
5269 | 5270 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FHxSA | 266.48100977155 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 7.86 | 8.86 | Baseline | Baseline | 1927054.0268 | 35390.6792 | 54.4509 | 305807.586 | 7203.3193 | 42.4537 | 6.3015 | 4.9131 | 305970.0609 | 7206.5337 | 0.99484 | 0.98414 | 7.993 | 9.0171 | 1.6919 | 0.13298 | 7.7625 | 8.8818 | 8.3687 | 9.2131 | 0.60622 | 0.33127 | 0.10163 | 0.076194 | 87.6498 | 100.6145 | 0.0013426 | 0.00073895 | FALSE | 0.88642 | FALSE | [M-H]- | 66.6203 | 7.913 | 8.0917 | 0.17877 | 7.9221 | 8.0797 | 0.15764 | 677.2566 | 1.1981 | 1.4066 | 6.5953 | 0 | 43 | 10 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1927054.0268 | No Control Sample | 0 | off | off | 1927054.0268 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9528 | 0.42559 | 0.16936 | 77.9639 | -19.2159 | -1.4982 | 0 | 0 | 0 | 0 | 0 | 1.4575 | FALSE | 100 | 54.4509
5271 | 5272 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FHpSA | 162.516739059021 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 8.58 | 8.86 | Baseline | Baseline | 2434239.4298 | 35390.6792 | 68.7819 | 363531.8079 | 7203.3193 | 50.4673 | 6.6961 | 4.9131 | 366357.2589 | 7206.5337 | 0.96864 | 0.98414 | 8.5792 | 9.0171 | 0.0089011 | 0.00076371 | 8.448 | 8.8818 | 8.7482 | 9.2131 | 0.30025 | 0.33127 | 0.11083 | 0.076194 | 51.5569 | 100.6145 | 0.0046935 | 0.00073895 | FALSE | 0.95144 | FALSE | [M-H]- | 40.6292 | 8.4869 | 8.6736 | 0.1867 | 8.4947 | 8.6652 | 0.17047 | -5682.6607 | 1.0512 | 1.1042 | 4.0876 | 0 | 18 | 9 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2434239.4298 | No Control Sample | 0 | off | off | 2434239.4298 | 0 | off | 0 | off | 0 | off | 0 | off | 447.949 | -0.93223 | -0.41759 | 77.9637 | -22.079 | -1.7214 | 0 | 0 | 0 | 0 | 0 | 0.68895 | FALSE | 100 | 68.7819
5273 | 5274 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FOSA | 218.295866557247 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 8.88 | 8.86 | Baseline | Baseline | 1757842.1805 | 35390.6792 | 49.6696 | 278828.4406 | 7203.3193 | 38.7083 | 6.3044 | 4.9131 | 279126.1296 | 7206.5337 | 0.98826 | 0.98414 | 9.0183 | 9.0171 | 1.5571 | 0.13827 | 8.8103 | 8.8818 | 9.2252 | 9.2131 | 0.41483 | 0.33127 | 0.10242 | 0.076194 | 91.5283 | 100.6145 | 0.0020826 | 0.00073895 | FALSE | 1.0001 | FALSE | [M-H]- | 54.574 | 8.9243 | 9.1132 | 0.18887 | 8.9352 | 9.1031 | 0.16793 | 1399.8055 | 1.0322 | 1.0821 | 5.6347 | 0 | 22 | 7 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1757842.1805 | No Control Sample | 0 | off | off | 1757842.1805 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9466 | 0.87836 | 0.43738 | 77.9639 | -19.4919 | -1.5197 | 0 | 0 | 0 | 0 | 0 | 0.80004 | FALSE | 100 | 49.6696
5292 | 5293 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 4:2 FTS | 312.039007510679 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 5.94 | 5.92 | Baseline | Baseline | 1505761.7889 | 9503.8182 | 158.4376 | 230799.3249 | 1500.2921 | 153.8363 | 6.5241 | 6.3346 | 230939.0462 | 1517.0947 | 0.99557 | 0.95918 | 5.9335 | 5.9298 | 0.11007 | 0.0065384 | 5.6853 | 5.8112 | 6.1448 | 6.0835 | 0.45943 | 0.2723 | 0.10419 | 0.1007 | 201.5455 | 133.1484 | 0.00112 | 0.021486 | FALSE | 1.0006 | FALSE | [M-H]- | 78.0098 | 5.8286 | 6.0388 | 0.21017 | 5.8438 | 6.0274 | 0.18362 | 562.6565 | 1.0208 | 1.0907 | 2.2322 | 0 | 58 | 14 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1505761.7889 | No Control Sample | 0 | off | off | 1505761.7889 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9751 | 2.373 | 0.77591 | 306.9666 | -3.171 | -0.97339 | 0 | 0 | 0 | 0 | 0 | 1.2923 | FALSE | 100 | 158.4376
5294 | 5295 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 6:2 FTS | 254.611018339461 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 7.92 | 7.92 | Baseline | Baseline | 991797.5069 | 6420.2702 | 154.4791 | 213067.2339 | 1568.0831 | 135.8775 | 4.6549 | 4.0943 | 213079.0008 | 1571.75 | 0.99888 | 0.98013 | 7.9262 | 7.9282 | 0.078853 | 0.0062451 | 7.7359 | 7.7707 | 8.4515 | 8.0431 | 0.71555 | 0.27242 | 0.072238 | 0.06183 | 100.3535 | 111.3901 | 0.00020982 | 0.0040097 | FALSE | 0.99976 | FALSE | [M-H]- | 63.6528 | 7.8494 | 7.9999 | 0.15052 | 7.8608 | 7.9891 | 0.12824 | 62.4763 | 1.026 | 1.0708 | 2.6749 | 0 | 49 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 991797.5069 | No Control Sample | 0 | off | off | 991797.5069 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9689 | 2.4215 | 1.0339 | 406.9598 | -4.2785 | -1.7412 | 0 | 0 | 0 | 0 | 0 | 0.46452 | FALSE | 100 | 154.4791
5296 | 5297 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 8:2 FTS | 249.648327520374 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 8.88 | 8.89 | Baseline | Baseline | 887078.8553 | 7242.0241 | 122.4905 | 155927.5169 | 1395.9065 | 111.7034 | 5.689 | 5.188 | 156070.2576 | 1397.2574 | 0.99379 | 0.99369 | 8.8864 | 8.8849 | 0.072584 | 0.0064455 | 8.7144 | 8.7226 | 9.0751 | 9.0275 | 0.36068 | 0.3049 | 0.088442 | 0.078418 | 97.7425 | 79.4726 | 0.0017989 | 0.0016933 | FALSE | 1.0002 | FALSE | [M-H]- | 62.4121 | 8.7963 | 8.9798 | 0.18347 | 8.8071 | 8.9701 | 0.16299 | 777.6696 | 0.92207 | 0.83734 | 2.0184 | 0 | 19 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 887078.8553 | No Control Sample | 0 | off | off | 887078.8553 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9632 | 3.2369 | 1.7057 | 506.9541 | -3.0814 | -1.5622 | 0 | 0 | 0 | 0 | 0 | 1.614 | FALSE | 100 | 122.4905
5298 | 5299 | PFAS CS9 100ng/ml | 38 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 39) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 10:2 FTS | 208.167979254411 | Quantifiers | 10/27/2023 05:22:02 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.5 | 8.89 | Baseline | Baseline | 770028.7357 | 7242.0241 | 106.3278 | 146484.8307 | 1395.9065 | 104.9389 | 5.2567 | 5.188 | 146524.5683 | 1397.2574 | 0.9948 | 0.99369 | 9.5127 | 8.8849 | 0.13406 | 0.012736 | 9.3595 | 8.7226 | 9.7206 | 9.0275 | 0.36115 | 0.3049 | 0.084487 | 0.078418 | 296.9499 | 79.4726 | 0.00061651 | 0.0016933 | FALSE | 1.0707 | FALSE | [M-H]- | 52.042 | 9.4304 | 9.5912 | 0.16081 | 9.4411 | 9.5827 | 0.14166 | 250.059 | 0.91908 | 0.84444 | 2.4624 | 0 | 26 | 8 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 770028.7357 | No Control Sample | 0 | off | off | 770028.7357 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9564 | 1.9774 | 1.2397 | 606.9483 | 0.5802 | 0.35215 | 0 | 0 | 0 | 0 | 0 | 19.6406 | FALSE | 100 | 106.3278
5359 | 5360 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFBA | 300.0708215704 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 1.49 | 1.48 | Manual | Baseline | 3156336.8227 | 58640.5513 | 53.8252 | 283866.125 | 6566.7641 | 43.2277 | 11.1191 | 8.9299 | 284494.1067 | 6579.0137 | 0.99211 | 0.98966 | 1.4774 | 1.468 | 0.8456 | 0.012599 | 1.3263 | 1.3245 | 1.766 | 1.707 | 0.43962 | 0.38245 | 0.18427 | 0.14101 | 143.0016 | 150.2246 | 0.00083518 | 0.0037813 | TRUE | 1.0064 | FALSE | [M-H]- | 75.0177 | 1.3589 | 1.668 | 0.30906 | 1.367 | 1.6488 | 0.28175 | -539.2759 | 1.4312 | 1.8215 | 0.91876 | 0 | 92 | 40 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3156336.8227 | No Control Sample | 0 | off | off | 3156336.8227 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9803 | 5.3473 | 1.1389 | 168.9874 | -9.5712 | -1.6174 | 0 | 0 | 0 | 0 | 0 | 0.3442 | FALSE | 100 | 53.8252
5361 | 5362 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFPeA | 348.183297709524 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 3.18 | 3.21 | Manual | Manual | 3734748.3995 | 47962.8944 | 77.8675 | 275661.2689 | 4378.0852 | 62.9639 | 13.5483 | 10.9552 | 276215.714 | 4389.1841 | 0.99204 | 0.98723 | 3.2123 | 3.2077 | 1.0167 | 0.032332 | 2.962 | 2.972 | 3.5835 | 3.5071 | 0.62149 | 0.53514 | 0.20757 | 0.17341 | 723.466 | 387.644 | 0.00075908 | 0.0023544 | TRUE | 1.0015 | FALSE | [M-H]- | 87.0458 | 3.0324 | 3.4378 | 0.40542 | 3.052 | 3.4103 | 0.35837 | -336.6874 | 1.3761 | 1.8064 | 1.0949 | 0 | 111 | 39 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3734748.3995 | No Control Sample | 0 | off | off | 3734748.3995 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9769 | 3.4311 | 0.9023 | 218.9843 | -7.0211 | -1.5375 | 0 | 0 | 0 | 0 | 0 | 0.7445 | FALSE | 100 | 77.8675
5363 | 5364 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFHxA | 184.965617939143 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 1545.1773 | 6.13 | 6.12 | Baseline | Baseline | 3149744.8815 | 73170.5521 | 43.0466 | 303198.5838 | 10998.156 | 27.5681 | 10.3884 | 6.653 | 303570.1081 | 11001.4408 | 0.99232 | 0.996 | 6.1447 | 6.1405 | 0.23901 | 0.014651 | 5.9455 | 5.9645 | 6.4021 | 6.3592 | 0.45658 | 0.39467 | 0.17202 | 0.10592 | 105.4737 | 193.8335 | 0.0025001 | 0.00063686 | FALSE | 1.0007 | FALSE | [M-H]- | 46.2414 | 6.0157 | 6.2773 | 0.26158 | 6.0265 | 6.2623 | 0.23585 | 1660.2457 | 0.97337 | 0.90886 | 1.5372 | 0 | 51 | 21 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3149744.8815 | No Control Sample | 203843.5852 | off | off | 3149744.8815 | 203843.5852 | off | 0.0025 | off | 400 | off | 6180.7093 | off | 312.9734 | 1.9674 | 0.61573 | 268.9808 | -5.9603 | -1.6032 | 0 | 0 | 0 | 0 | 0 | 0.87805 | FALSE | 100 | 43.0466
5365 | 5366 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFHpA | 139.368734431711 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 993.3108 | 7.26 | 7.27 | Baseline | Baseline | 2268014.3879 | 65035.7132 | 34.8734 | 291019.3794 | 13980.6435 | 20.8159 | 7.7933 | 4.6518 | 291470.6057 | 13992.8759 | 0.98839 | 0.9934 | 7.2865 | 7.2858 | 0.36512 | 0.026508 | 7.1372 | 7.1489 | 7.5716 | 7.413 | 0.43438 | 0.26412 | 0.12917 | 0.072236 | 103.5173 | 224.1052 | 0.0034082 | 0.0016238 | FALSE | 1.0001 | FALSE | [M-H]- | 34.8422 | 7.1908 | 7.3829 | 0.19215 | 7.198 | 7.3732 | 0.17527 | 2283.3642 | 0.99308 | 0.95743 | 1.8791 | 0 | 49 | 17 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2268014.3879 | No Control Sample | 228328.7703 | off | off | 2268014.3879 | 228328.7703 | off | 0.0025 | off | 400 | off | 3973.2433 | off | 362.97 | 0.99034 | 0.35946 | 318.9773 | -5.6826 | -1.8126 | 0 | 0 | 0 | 0 | 0 | 0.37031 | FALSE | 100 | 34.8734
5367 | 5368 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFOA | 187.796300753659 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 1015.8672 | 7.96 | 7.94 | Baseline | Baseline | 1945135.8987 | 42708.4881 | 45.5445 | 278054.8246 | 9531.2333 | 29.173 | 6.9955 | 4.4809 | 279010.9203 | 9539.2533 | 0.98368 | 0.99153 | 7.9783 | 7.973 | 0.23004 | 0.018311 | 7.8451 | 7.8516 | 8.3672 | 8.2117 | 0.52203 | 0.36008 | 0.11438 | 0.068819 | 68.5993 | 122.5211 | 0.0023943 | 0.00087769 | FALSE | 1.0007 | FALSE | [M-H]- | 46.9491 | 7.8942 | 8.0868 | 0.19265 | 7.9016 | 8.0743 | 0.17269 | -1275.2784 | 1.0211 | 0.98804 | 2.3907 | 0 | 37 | 11 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1945135.8987 | No Control Sample | 191475.4056 | off | off | 1945135.8987 | 191475.4056 | off | 0.0025 | off | 400 | off | 4063.4689 | off | 412.9672 | 1.9737 | 0.81505 | 368.9742 | -5.0513 | -1.8638 | 0 | 0 | 0 | 0 | 0 | 5.8839 | FALSE | 100 | 45.5445
5369 | 5370 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFNA | 189.952750503227 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 1198.2688 | 8.46 | 8.47 | Baseline | Baseline | 1906965.6213 | 43214.9447 | 44.1275 | 255295.2577 | 7868.2417 | 32.4463 | 7.4696 | 5.4923 | 255544.0456 | 7868.2417 | 0.9938 | 0.99944 | 8.4984 | 8.4919 | 0.45432 | 0.038436 | 8.3325 | 8.2769 | 8.7457 | 8.8156 | 0.41315 | 0.53865 | 0.12307 | 0.08704 | 71.8717 | 96.141 | 0.0018583 | 1.5491e-11 | FALSE | 1.0008 | FALSE | [M-H]- | 47.4882 | 8.3902 | 8.5965 | 0.20634 | 8.3991 | 8.5864 | 0.18733 | 1148.3138 | 0.87924 | 0.72721 | 1.6199 | 0 | 24 | 8 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1906965.6213 | No Control Sample | 159143.3885 | off | off | 1906965.6213 | 159143.3885 | off | 0.0025 | off | 400 | off | 4793.0753 | off | 462.9642 | 2.0291 | 0.9394 | 418.971 | -5.1375 | -2.1525 | 0 | 0 | 0 | 0 | 0 | 40.3117 | FALSE | 100 | 44.1275
5371 | 5372 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFDA | 202.564344649739 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 8.87 | 8.86 | Baseline | Baseline | 1681165.7622 | 33432.4184 | 50.2855 | 234716.1725 | 6098.3463 | 38.4885 | 7.1625 | 5.4822 | 234778.9898 | 6099.6019 | 0.99769 | 0.99847 | 8.9089 | 8.9074 | 0.43904 | 0.038943 | 8.6916 | 8.718 | 9.2207 | 9.1714 | 0.52903 | 0.45337 | 0.11583 | 0.084726 | 63.6459 | 83.9058 | 0.00059859 | 0.000513 | FALSE | 1.0002 | FALSE | [M-H]- | 50.6411 | 8.8062 | 9.0117 | 0.20543 | 8.8148 | 9.003 | 0.18826 | 265.5748 | 0.99042 | 0.97822 | 1.3027 | 0 | 28 | 7 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Fail | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1681165.7622 | No Control Sample | 0 | off | off | 1681165.7622 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9613 | 2.5151 | 1.2902 | 468.968 | -4.2923 | -2.0129 | 0 | 0 | 0 | 0 | 0 | 0.69854 | FALSE | 100 | 50.2855
5373 | 5374 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFUdA | 227.186522664983 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.21 | 9.2 | Baseline | Baseline | 1484871.6452 | 29472.0361 | 50.3824 | 232747.6045 | 5923.4828 | 39.2924 | 6.3798 | 4.9755 | 233068.6033 | 5925.2018 | 0.9908 | 0.99501 | 9.2418 | 9.2434 | 0.34573 | 0.031842 | 9.0878 | 9.0775 | 9.3859 | 9.3757 | 0.2981 | 0.2982 | 0.10424 | 0.078127 | 120.0006 | 155.6841 | 0.0023091 | 0.00050217 | FALSE | 0.99983 | FALSE | [M-H]- | 56.7966 | 9.149 | 9.3345 | 0.18553 | 9.1587 | 9.3272 | 0.16841 | 1802.8424 | 1.008 | 1.047 | 0.92043 | 0 | 17 | 7 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1484871.6452 | No Control Sample | 0 | off | off | 1484871.6452 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9579 | 1.9045 | 1.0721 | 518.9649 | -4.5562 | -2.3645 | 0 | 0 | 0 | 0 | 0 | 1.4473 | FALSE | 100 | 50.3824
5375 | 5376 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFDoA | 171.294517874991 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.47 | 9.47 | Baseline | Baseline | 1427818.1303 | 32210.8352 | 44.3273 | 230918.9632 | 7011.7329 | 32.9332 | 6.1832 | 4.5938 | 231045.0864 | 7014.981 | 0.99425 | 0.95905 | 9.5019 | 9.499 | 0.33633 | 0.031851 | 9.3534 | 9.3765 | 9.7043 | 9.6718 | 0.35088 | 0.2953 | 0.10075 | 0.072363 | 128.6765 | 103.5049 | 0.0012727 | 0.00097163 | FALSE | 1.0003 | FALSE | [M-H]- | 42.8236 | 9.4195 | 9.5932 | 0.17369 | 9.4247 | 9.5821 | 0.15743 | 837.5413 | 1.0895 | 1.1127 | 1.5991 | 0 | 25 | 9 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1427818.1303 | No Control Sample | 0 | off | off | 1427818.1303 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9541 | 0.70501 | 0.43214 | 568.9616 | -3.2072 | -1.8248 | 0 | 0 | 0 | 0 | 0 | 0.99114 | FALSE | 100 | 44.3273
5377 | 5378 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFTrDA | 209.550397614059 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.68 | 9.47 | Baseline | Baseline | 1273219.198 | 32210.8352 | 39.5277 | 219501.7637 | 7011.7329 | 31.3049 | 5.8005 | 4.5938 | 219671.6654 | 7014.981 | 0.99187 | 0.95905 | 9.7177 | 9.499 | 0.38982 | 0.037734 | 9.5424 | 9.3765 | 9.9163 | 9.6718 | 0.37393 | 0.2953 | 0.094182 | 0.072363 | 248.5331 | 103.5049 | 0.0015521 | 0.00097163 | FALSE | 1.023 | FALSE | [M-H]- | 52.3876 | 9.6366 | 9.7981 | 0.16151 | 9.6448 | 9.7882 | 0.14336 | 911.1003 | 1.0104 | 0.99935 | 1.3391 | 0 | 34 | 10 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1273219.198 | No Control Sample | 0 | off | off | 1273219.198 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9507 | 0.39953 | 0.26487 | 618.9586 | -5.0376 | -3.1181 | 0 | 0 | 0 | 0 | 0 | 3.1847 | FALSE | 100 | 39.5277
5379 | 5380 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFTeDA | 258.767694303075 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.87 | 9.86 | Baseline | Baseline | 1138753.6364 | 12412.1241 | 91.7453 | 215429.3363 | 2859.8262 | 75.3295 | 5.286 | 4.3402 | 215596.6357 | 2859.8262 | 0.99248 | 0.99471 | 9.8935 | 9.8905 | 0.23819 | 0.02351 | 9.7704 | 9.7431 | 10.0983 | 10.0442 | 0.32793 | 0.30108 | 0.087229 | 0.069601 | 442.4329 | 622.4825 | 0.0018803 | 1.044e-08 | FALSE | 1.0003 | FALSE | [M-H]- | 64.6919 | 9.8247 | 9.9659 | 0.14122 | 9.8314 | 9.9576 | 0.12615 | 1235.2499 | 1.1054 | 1.2088 | 1.0602 | 0 | 40 | 11 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1138753.6364 | No Control Sample | 0 | off | off | 1138753.6364 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9476 | 0.45691 | 0.32576 | 668.9557 | -3.2826 | -2.1959 | 0 | 0 | 0 | 0 | 0 | 2.8277 | FALSE | 100 | 91.7453
5386 | 5387 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFBS | 409.019802737013 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 3.5 | 3.5 | Baseline | Manual | 2938075.7341 | 27138.4293 | 108.2626 | 242716.6907 | 2203.4106 | 110.155 | 12.105 | 12.3166 | 243114.1011 | 2212.8204 | 0.99389 | 0.98168 | 3.8527 | 3.8551 | 10.078 | 0.35273 | 3.4832 | 3.603 | 4.2354 | 4.165 | 0.75218 | 0.56204 | 0.18733 | 0.20539 | 857.2989 | 512.6736 | 0.0030512 | 0.0020795 | FALSE | 0.99939 | TRUE | [M-H]- | 102.255 | 3.6615 | 4.0825 | 0.42099 | 3.6898 | 4.049 | 0.35919 | 984.5723 | 1.0041 | 0.98097 | 3.9235 | 0 | 118 | 31 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2938075.7341 | No Control Sample | 0 | off | off | 2938075.7341 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9435 | 1.5571 | 0.46548 | 79.9562 | -13.0636 | -1.0445 | 0 | 0 | 0 | 0 | 0 | 1.4313 | FALSE | 100 | 108.2626
5388 | 5389 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFPeS | 284.522105016726 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 6.38 | 3.5 | Baseline | Manual | 2057855.4303 | 27138.4293 | 75.8281 | 296533.2807 | 2203.4106 | 134.5792 | 6.9397 | 12.3166 | 296879.5233 | 2212.8204 | 0.99019 | 0.98168 | 6.3968 | 3.8551 | 0.26302 | 0.016781 | 6.2247 | 3.603 | 6.6094 | 4.165 | 0.38473 | 0.56204 | 0.11222 | 0.20539 | 190.8845 | 512.6736 | 0.0024887 | 0.0020795 | FALSE | 1.6593 | FALSE | [M-H]- | 71.1305 | 6.2978 | 6.4983 | 0.20055 | 6.3081 | 6.486 | 0.17784 | 1918.1717 | 1.0445 | 1.0767 | 4.6035 | 0 | 43 | 14 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2057855.4303 | No Control Sample | 0 | off | off | 2057855.4303 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9402 | 1.182 | 0.41245 | 79.9559 | -15.7193 | -1.2569 | 0 | 0 | 0 | 0 | 0 | 1.3616 | FALSE | 100 | 75.8281
5390 | 5391 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | L-PFHxS | 311.05378898539 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 324.4 | 4 | 81.1 | 0 | 7.34 | 7.35 | Baseline | Baseline | 1361899.3666 | 23439.3697 | 58.1031 | 296091.6379 | 5358.3655 | 55.2578 | 4.5996 | 4.3744 | 301303.7032 | 5377.1675 | 0.61795 | 0.97686 | 7.3635 | 7.3604 | 0.3198 | 0.023473 | 7.2729 | 7.2501 | 7.5043 | 7.4957 | 0.23142 | 0.24565 | 0.074438 | 0.069405 | 65.3546 | 197.0428 | 0.021166 | 0.0045323 | FALSE | 1.0004 | TRUE | [M-H]- | 95.8859 | 7.2969 | 7.4351 | 0.13823 | 7.3034 | 7.4264 | 0.12302 | -27081.2498 | 0.92974 | 0.81298 | 7.1091 | 0 | 29 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1361899.3666 | No Control Sample | 0 | off | off | 1361899.3666 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9363 | -0.84201 | -0.33591 | 79.9557 | -18.9545 | -1.5155 | 0 | 0 | 0 | 0 | 0 | 0.9115 | FALSE | 81.1 | 58.1031
5392 | 5393 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | Br-PFHxS | 124.090257949311 | Qualifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 75.6 | 4 | 18.9 | 0 | 7.16 | 7.35 | Manual | Baseline | 738160.8508 | 23439.3697 | 31.4924 | 130301.1644 | 5358.3655 | 24.3173 | 5.665 | 4.3744 | 301303.7032 | 5377.1675 | 0.33501 | 0.97686 | 7.1787 | 7.3604 | 0.26068 | 0.018665 | 7.059 | 7.2501 | 7.275 | 7.4957 | 0.21599 | 0.24565 | 0.092091 | 0.069405 | 28.7606 | 197.0428 | 0.014481 | 0.0045323 | TRUE | 0.9753 | FALSE | [M-H]- | 164.1406 | 7.092 | 7.2706 | 0.17862 | 7.1003 | 7.2549 | 0.15464 | 8736.1976 | 0.96608 | 0.83661 | 0.54201 | 0 | 26 | 12 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 738160.8508 | No Control Sample | 0 | off | off | 1361899.3666 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9363 | -0.84201 | -0.33591 | 79.9575 | 3.6347 | 0.29062 | 0 | 0 | 0 | 0 | 0 | 0.9115 | FALSE | 18.9 | 31.4924
5394 | 5395 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFHpS | 390.202580222338 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 7.98 | 7.35 | Baseline | Baseline | 1391980.8152 | 23439.3697 | 59.3864 | 269065.0811 | 5358.3655 | 50.214 | 5.1734 | 4.3744 | 269288.4819 | 5377.1675 | 0.99232 | 0.97686 | 8.0047 | 7.3604 | 0.30971 | 0.024715 | 7.8873 | 7.2501 | 8.3897 | 7.4957 | 0.50238 | 0.24565 | 0.082499 | 0.069405 | 126.0819 | 197.0428 | 0.00026933 | 0.0045323 | FALSE | 1.0875 | TRUE | [M-H]- | 97.5506 | 7.9334 | 8.0955 | 0.16213 | 7.9409 | 8.0828 | 0.14194 | -144.2495 | 1.1402 | 1.2328 | 9.4748 | 0 | 35 | 8 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1391980.8152 | No Control Sample | 0 | off | off | 1391980.8152 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9336 | 0.33504 | 0.15041 | 79.956 | -15.0531 | -1.2036 | 0 | 0 | 0 | 0 | 0 | 1.6341 | FALSE | 100 | 59.3864
5396 | 5397 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | L-PFOS | 314.57366989932 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 315.2 | 4 | 78.8 | 0 | 8.46 | 8.45 | Manual | Baseline | 1219113.8796 | 14909.6202 | 81.7669 | 219982.8398 | 2750.1057 | 79.9907 | 5.5419 | 5.4215 | 222411.6363 | 2750.1864 | 0.50549 | 0.96538 | 8.4938 | 8.4891 | 0.39964 | 0.033809 | 8.4148 | 8.3265 | 8.6023 | 8.6661 | 0.18752 | 0.3396 | 0.089791 | 0.085624 | 31.8149 | 73.4344 | 0.010161 | 6.4438e-05 | TRUE | 1.0006 | TRUE | [M-H]- | 99.8013 | 8.4089 | 8.5753 | 0.16646 | 8.4089 | 8.5651 | 0.15628 | -11920.1733 | 0.99934 | 0.87646 | 6.661 | 0 | 12 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1219113.8796 | No Control Sample | 0 | off | off | 1219113.8796 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9305 | 0.59158 | 0.29516 | 79.9558 | -18.4852 | -1.478 | 0 | 0 | 0 | 0 | 0 | 1.1344 | FALSE | 78.8 | 81.7669
5398 | 5399 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | Br-PFOS | 132.188011157934 | Qualifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 84.4 | 4 | 21.1 | 0 | 8.3 | 8.45 | Manual | Baseline | 1096720.627 | 14909.6202 | 73.5579 | 170841.2339 | 2750.1057 | 62.1217 | 6.4195 | 5.4215 | 222411.6363 | 2750.1864 | 0.45487 | 0.96538 | 8.3414 | 8.4891 | 0.49936 | 0.041447 | 8.229 | 8.3265 | 8.4233 | 8.6661 | 0.19428 | 0.3396 | 0.10455 | 0.085624 | 24.7078 | 73.4344 | 0.0020711 | 6.4438e-05 | TRUE | 0.98261 | FALSE | [M-H]- | 156.6209 | 8.2475 | 8.4297 | 0.18217 | 8.2587 | 8.4297 | 0.17104 | 1821.2316 | 1.029 | 1.2101 | 0.8996 | 0 | 13 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1096720.627 | No Control Sample | 0 | off | off | 1219113.8796 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9305 | 0.59158 | 0.29516 | 79.956 | -16.3893 | -1.3104 | 0 | 0 | 0 | 0 | 0 | 1.1344 | FALSE | 21.1 | 73.5579
5400 | 5401 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFNS | 381.76271406283 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 8.85 | 8.45 | Baseline | Baseline | 1261595.7569 | 14909.6202 | 84.6162 | 223358.4414 | 2750.1057 | 81.2181 | 5.6483 | 5.4215 | 223424.718 | 2750.1864 | 0.99634 | 0.96538 | 8.8864 | 8.4891 | 0.41145 | 0.036413 | 8.7323 | 8.3265 | 9.2445 | 8.6661 | 0.51217 | 0.3396 | 0.088339 | 0.085624 | 86.4547 | 73.4344 | 0.00043349 | 6.4438e-05 | FALSE | 1.0468 | TRUE | [M-H]- | 95.4407 | 8.8 | 8.9842 | 0.18418 | 8.8125 | 8.9719 | 0.15935 | -189.0466 | 0.98089 | 0.95839 | 7.3489 | 0 | 27 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Fail | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1261595.7569 | No Control Sample | 0 | off | off | 1261595.7569 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9274 | 0.69224 | 0.37999 | 79.9561 | -12.3626 | -0.98848 | 0 | 0 | 0 | 0 | 0 | 1.632 | FALSE | 100 | 84.6162
5402 | 5403 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFDS | 380.819376037297 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 400 | 4 | 100 | 0 | 9.19 | 8.45 | Baseline | Baseline | 1293033.7108 | 14909.6202 | 86.7248 | 252293.5551 | 2750.1057 | 91.7396 | 5.1251 | 5.4215 | 252439.3964 | 2750.1864 | 0.99214 | 0.96538 | 9.2123 | 8.4891 | 0.24253 | 0.022289 | 9.0736 | 8.3265 | 9.3903 | 8.6661 | 0.31668 | 0.3396 | 0.080759 | 0.085624 | 172.0153 | 73.4344 | 0.0011495 | 6.4438e-05 | FALSE | 1.0852 | TRUE | [M-H]- | 95.2048 | 9.134 | 9.3018 | 0.16783 | 9.143 | 9.2901 | 0.14714 | 915.7712 | 1.1586 | 1.3199 | 5.3256 | 0 | 18 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1293033.7108 | No Control Sample | 0 | off | off | 1293033.7108 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9244 | 0.97479 | 0.58383 | 79.9561 | -10.3463 | -0.82725 | 0 | 0 | 0 | 0 | 0 | 1.2478 | FALSE | 100 | 86.7248
5406 | 5407 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FBSA | 506.185654350721 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 4.6 | 8.86 | Manual | Baseline | 4959667.3321 | 37159.4868 | 133.4697 | 354150.3801 | 7719.8382 | 45.8754 | 14.0044 | 4.8135 | 354597.0137 | 7719.9308 | 0.99601 | 0.98077 | 4.9841 | 9.0313 | 8.3489 | 0.38405 | 4.654 | 8.8613 | 5.3648 | 9.1926 | 0.71075 | 0.33125 | 0.22039 | 0.074452 | 550.7088 | 69.4923 | 0 | 2.437e-05 | TRUE | 0.55187 | FALSE | [M-H]- | 126.5464 | 4.7671 | 5.2116 | 0.44447 | 4.8 | 5.1752 | 0.3752 | 0 | 0.91022 | 0.77585 | 2.3151 | 0 | 133 | 43 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4959667.3321 | No Control Sample | 0 | off | off | 4959667.3321 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9594 | 1.5813 | 0.47116 | 77.9641 | -15.2419 | -1.1883 | 0 | 0 | 0 | 0 | 0 | 1.4001 | FALSE | 100 | 133.4697
5408 | 5409 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FPeSA | 202.902611174673 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 7 | 8.86 | Baseline | Baseline | 3087284.7285 | 37159.4868 | 83.082 | 369873.3282 | 7719.8382 | 47.9121 | 8.3469 | 4.8135 | 371043.2507 | 7719.9308 | 0.9821 | 0.98077 | 7.1177 | 9.0313 | 1.6811 | 0.11767 | 6.938 | 8.8613 | 7.3048 | 9.1926 | 0.36675 | 0.33125 | 0.13675 | 0.074452 | 111.7432 | 69.4923 | 0.0068376 | 2.437e-05 | FALSE | 0.78811 | FALSE | [M-H]- | 50.7257 | 7.0114 | 7.2264 | 0.21495 | 7.021 | 7.214 | 0.19298 | 6895.8624 | 1.1916 | 1.393 | 2.7655 | 0 | 41 | 17 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3087284.7285 | No Control Sample | 0 | off | off | 3087284.7285 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9561 | 0.82572 | 0.28731 | 77.9637 | -20.1027 | -1.5673 | 0 | 0 | 0 | 0 | 0 | 4.9629 | FALSE | 100 | 83.082
5410 | 5411 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FHxSA | 248.73726989084 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 7.86 | 8.86 | Baseline | Baseline | 1888640.5017 | 37159.4868 | 50.8253 | 299102.5873 | 7719.8382 | 38.7447 | 6.3144 | 4.8135 | 299420.356 | 7719.9308 | 0.99202 | 0.98077 | 8.0094 | 9.0313 | 1.9012 | 0.14943 | 7.8593 | 8.8613 | 8.3698 | 9.1926 | 0.51052 | 0.33125 | 0.10147 | 0.074452 | 96.2425 | 69.4923 | 0.0007999 | 2.437e-05 | FALSE | 0.88686 | FALSE | [M-H]- | 62.1843 | 7.9274 | 8.1085 | 0.18109 | 7.9365 | 8.0963 | 0.15982 | 468.645 | 1.2162 | 1.4458 | 7.2054 | 0 | 36 | 9 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1888640.5017 | No Control Sample | 0 | off | off | 1888640.5017 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9526 | -0.059656 | -0.02374 | 77.9639 | -19.644 | -1.5315 | 0 | 0 | 0 | 0 | 0 | 0.85771 | FALSE | 100 | 50.8253
5412 | 5413 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FHpSA | 156.148677761728 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 8.58 | 8.86 | Baseline | Baseline | 2455750.9156 | 37159.4868 | 66.0868 | 363134.5319 | 7719.8382 | 47.0391 | 6.7626 | 4.8135 | 366034.571 | 7719.9308 | 0.96797 | 0.98077 | 8.5937 | 9.0313 | 0.1598 | 0.013711 | 8.4696 | 8.8613 | 8.7687 | 9.1926 | 0.2991 | 0.33125 | 0.11144 | 0.074452 | 62.053 | 69.4923 | 0.0053737 | 2.437e-05 | FALSE | 0.95155 | FALSE | [M-H]- | 39.0372 | 8.5033 | 8.6904 | 0.18717 | 8.5121 | 8.6815 | 0.16946 | -6524.1441 | 1.0434 | 1.0948 | 4.3176 | 0 | 18 | 9 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2455750.9156 | No Control Sample | 0 | off | off | 2455750.9156 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9495 | 0.21398 | 0.095852 | 77.9637 | -21.4813 | -1.6748 | 0 | 0 | 0 | 0 | 0 | 1.4143 | FALSE | 100 | 66.0868
5414 | 5415 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FOSA | 209.611911960853 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 8.88 | 8.86 | Baseline | Baseline | 1772275.1117 | 37159.4868 | 47.6937 | 284250.9354 | 7719.8382 | 36.8208 | 6.2349 | 4.8135 | 284563.7024 | 7719.9308 | 0.98802 | 0.98077 | 9.0325 | 9.0313 | 1.7174 | 0.1525 | 8.8514 | 8.8613 | 9.2436 | 9.1926 | 0.39225 | 0.33125 | 0.10011 | 0.074452 | 75.936 | 69.4923 | 0.002228 | 2.437e-05 | FALSE | 1.0001 | FALSE | [M-H]- | 52.403 | 8.938 | 9.1263 | 0.18837 | 8.948 | 9.1165 | 0.16843 | 1614.5312 | 0.98105 | 0.95996 | 6.0916 | 0 | 21 | 7 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1772275.1117 | No Control Sample | 0 | off | off | 1772275.1117 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9458 | -0.71436 | -0.35571 | 77.9637 | -21.203 | -1.6531 | 0 | 0 | 0 | 0 | 0 | 1.1132 | FALSE | 100 | 47.6937
5433 | 5434 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 4:2 FTS | 299.349464209687 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 5.94 | 5.92 | Baseline | Baseline | 1586343.0789 | 10436.8482 | 151.9945 | 236634.4584 | 1710.2616 | 138.3616 | 6.7038 | 6.1025 | 236720.5279 | 1715.7213 | 0.99656 | 0.98528 | 5.9583 | 5.9563 | 0.30792 | 0.01829 | 5.72 | 5.8025 | 6.18 | 6.1098 | 0.45995 | 0.30727 | 0.10523 | 0.09331 | 193.9407 | 159.7381 | 0.000701 | 0.0060348 | FALSE | 1.0003 | FALSE | [M-H]- | 74.8374 | 5.8532 | 6.0616 | 0.20841 | 5.8676 | 6.05 | 0.18236 | 360.6474 | 1.0188 | 1.0763 | 2.2339 | 0 | 57 | 14 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1586343.0789 | No Control Sample | 0 | off | off | 1586343.0789 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9748 | 1.6769 | 0.54829 | 306.9665 | -3.4988 | -1.074 | 0 | 0 | 0 | 0 | 0 | 0.65823 | FALSE | 100 | 151.9945
5435 | 5436 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 6:2 FTS | 269.316407272128 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 7.92 | 7.92 | Baseline | Baseline | 1000967.9485 | 6125.8287 | 163.4012 | 212473.2051 | 1391.8007 | 152.6607 | 4.711 | 4.4014 | 212503.0201 | 1394.9329 | 0.99782 | 0.97056 | 7.9404 | 7.9396 | 0.25726 | 0.020375 | 7.7673 | 7.8146 | 8.3197 | 8.0611 | 0.55238 | 0.24647 | 0.07258 | 0.069858 | 100.2452 | 99.9882 | 0.00045254 | 0.0043168 | FALSE | 1.0001 | FALSE | [M-H]- | 67.3291 | 7.8641 | 8.0179 | 0.15377 | 7.8752 | 8.0049 | 0.12968 | 174.0673 | 1.0673 | 1.1286 | 2.5142 | 0 | 40 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1000967.9485 | No Control Sample | 0 | off | off | 1000967.9485 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9691 | 2.7574 | 1.1773 | 406.9599 | -3.9042 | -1.5889 | 0 | 0 | 0 | 0 | 0 | 2.244 | FALSE | 100 | 163.4012
5437 | 5438 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 8:2 FTS | 207.653366014568 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 8.88 | 8.89 | Baseline | Baseline | 942423.7636 | 9249.8279 | 101.8855 | 170986.0427 | 1980.1467 | 86.3502 | 5.5117 | 4.6713 | 171001.5545 | 1980.8704 | 0.9986 | 0.99383 | 8.9015 | 8.9027 | 0.2425 | 0.021534 | 8.7332 | 8.7227 | 9.2268 | 9.0457 | 0.49355 | 0.32297 | 0.086067 | 0.070601 | 112.5273 | 113.2202 | 0.00025064 | 0.000646 | FALSE | 0.99987 | FALSE | [M-H]- | 51.9133 | 8.8132 | 8.9947 | 0.18156 | 8.8221 | 8.9815 | 0.15937 | 86.8319 | 1.1005 | 1.1661 | 2.1092 | 0 | 26 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 942423.7636 | No Control Sample | 0 | off | off | 942423.7636 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9631 | 2.926 | 1.5419 | 506.9548 | -1.8144 | -0.91981 | 0 | 0 | 0 | 0 | 0 | 0.47034 | FALSE | 100 | 101.8855
5439 | 5440 | PFAS CS9 100ng/ml | 39 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 40) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 10:2 FTS | 167.970171816319 | Quantifiers | 10/27/2023 05:38:39 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.5 | 8.89 | Baseline | Baseline | 793594.8834 | 9249.8279 | 85.7956 | 149791.6609 | 1980.1467 | 75.6467 | 5.298 | 4.6713 | 149841.0621 | 1980.8704 | 0.99489 | 0.99383 | 9.5201 | 8.9027 | 0.21199 | 0.020139 | 9.3748 | 8.7227 | 9.7073 | 9.0457 | 0.33247 | 0.32297 | 0.085393 | 0.070601 | 197.3755 | 113.2202 | 0.00076594 | 0.000646 | FALSE | 1.0694 | FALSE | [M-H]- | 41.9925 | 9.4411 | 9.6001 | 0.15903 | 9.4492 | 9.5916 | 0.14243 | 345.0929 | 1.0568 | 1.1213 | 2.158 | 0 | 24 | 8 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 793594.8834 | No Control Sample | 0 | off | off | 793594.8834 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9557 | 0.94779 | 0.59422 | 606.9479 | 0.011869 | 0.007204 | 0 | 0 | 0 | 0 | 0 | 1.9919 | FALSE | 100 | 85.7956
5500 | 5501 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFBA | 307.610905876418 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 1.49 | 1.48 | Manual | Baseline | 3170576.0521 | 57461.2297 | 55.1777 | 283774.1536 | 6673.2626 | 42.524 | 11.1729 | 8.6107 | 284188.4388 | 6824.1923 | 0.99271 | 0.938 | 1.4688 | 1.4593 | 1.4259 | 0.021246 | 1.3201 | 1.3397 | 1.7325 | 1.6531 | 0.41239 | 0.31343 | 0.18873 | 0.1398 | 138.5579 | 147.8459 | 0.0011127 | 0.0091709 | TRUE | 1.0065 | FALSE | [M-H]- | 76.9027 | 1.3492 | 1.6569 | 0.30779 | 1.3575 | 1.638 | 0.28052 | 765.6865 | 1.4976 | 1.9699 | 1.02 | 0 | 86 | 41 | C4HF7O2 | 212.9792 | 168.989 | Gray | Fail | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3170576.0521 | No Control Sample | 0 | off | off | 3170576.0521 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9802 | 4.8594 | 1.0349 | 168.9874 | -9.3623 | -1.5821 | 0 | 0 | 0 | 0 | 0 | 0.45732 | FALSE | 100 | 55.1777
5502 | 5503 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFPeA | 342.188446278184 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 3.18 | 3.21 | Manual | Manual | 3774214.6932 | 49318.8825 | 76.5268 | 273186.5923 | 4250.4774 | 64.272 | 13.8155 | 11.6031 | 273820.7614 | 4263.0726 | 0.99054 | 0.98718 | 3.1997 | 3.1937 | 0.62094 | 0.019746 | 2.9411 | 2.9773 | 3.5551 | 3.4838 | 0.61404 | 0.50656 | 0.22269 | 0.18401 | 710.727 | 285.7046 | 0.00068338 | 0 | TRUE | 1.0019 | FALSE | [M-H]- | 85.5471 | 3.0164 | 3.4226 | 0.40615 | 3.0372 | 3.3952 | 0.35803 | 304.0372 | 1.2062 | 1.4255 | 1.2234 | 0 | 111 | 43 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3774214.6932 | No Control Sample | 0 | off | off | 3774214.6932 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9772 | 4.4029 | 1.1579 | 218.9845 | -6.0709 | -1.3294 | 0 | 0 | 0 | 0 | 0 | 0.23748 | FALSE | 100 | 76.5268
5504 | 5505 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFHxA | 177.854755113613 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 1554.353 | 6.13 | 6.12 | Baseline | Baseline | 3168448.8139 | 76547.8814 | 41.3917 | 298467.2226 | 11401.4894 | 26.1779 | 10.6157 | 6.7138 | 298792.3412 | 11423.101 | 0.99216 | 0.99053 | 6.1354 | 6.1337 | 0.08859 | 0.0054306 | 5.9436 | 5.9707 | 6.3992 | 6.3015 | 0.45562 | 0.33083 | 0.17464 | 0.10901 | 106.8986 | 179.8711 | 0.0022617 | 0.0036675 | FALSE | 1.0003 | FALSE | [M-H]- | 44.4637 | 6.0081 | 6.2728 | 0.26472 | 6.0188 | 6.256 | 0.23719 | 1481.631 | 1.2903 | 1.5817 | 1.773 | 0 | 51 | 21 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3168448.8139 | No Control Sample | 203843.5852 | off | off | 3168448.8139 | 203843.5852 | off | 0.0025 | off | 400 | off | 6217.4119 | off | 312.9734 | 1.7588 | 0.55044 | 268.9809 | -5.5529 | -1.4936 | 0 | 0 | 0 | 0 | 0 | 0.4635 | FALSE | 100 | 41.3917
5506 | 5507 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFHpA | 134.959524251551 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 996.8278 | 7.26 | 7.27 | Baseline | Baseline | 2276044.7606 | 67398.2652 | 33.7701 | 290639.5203 | 14548.5987 | 19.9771 | 7.8312 | 4.6326 | 291613.3511 | 14571.0591 | 0.98048 | 0.9915 | 7.2768 | 7.2756 | 0.23138 | 0.016799 | 7.1268 | 7.1476 | 7.4336 | 7.3953 | 0.3068 | 0.24772 | 0.1291 | 0.072771 | 98.19 | 200.484 | 0.0065824 | 0.0030278 | FALSE | 1.0002 | FALSE | [M-H]- | 33.7399 | 7.1823 | 7.3716 | 0.18922 | 7.1898 | 7.3629 | 0.17317 | 6235.6304 | 1.1943 | 1.4126 | 2.029 | 0 | 37 | 17 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2276044.7606 | No Control Sample | 228328.7703 | off | off | 2276044.7606 | 228328.7703 | off | 0.0025 | off | 400 | off | 3987.3114 | off | 362.9706 | 2.6529 | 0.96291 | 318.9774 | -5.2016 | -1.6592 | 0 | 0 | 0 | 0 | 0 | 1.5764 | FALSE | 100 | 33.7701
5508 | 5509 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFOA | 168.809432747711 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 1003.2163 | 7.96 | 7.94 | Baseline | Baseline | 1920912.566 | 46920.45 | 40.9398 | 277008.0632 | 10565.4384 | 26.2183 | 6.9345 | 4.4409 | 277926.2223 | 10565.8898 | 0.9843 | 0.99676 | 7.9657 | 7.9639 | 0.072096 | 0.0057388 | 7.8301 | 7.8234 | 8.3301 | 8.2104 | 0.49995 | 0.38695 | 0.11286 | 0.067537 | 60.3162 | 97.4048 | 0.0023971 | 0.00011871 | FALSE | 1.0002 | FALSE | [M-H]- | 42.2024 | 7.881 | 8.0763 | 0.19534 | 7.889 | 8.0617 | 0.17272 | -1328.151 | 1.174 | 1.2985 | 2.4029 | 0 | 36 | 11 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1920912.566 | No Control Sample | 191475.4056 | off | off | 1920912.566 | 191475.4056 | off | 0.0025 | off | 400 | off | 4012.8654 | off | 412.9678 | 3.3518 | 1.3842 | 368.9743 | -4.9831 | -1.8386 | 0 | 0 | 0 | 0 | 0 | 0.24315 | FALSE | 100 | 40.9398
5510 | 5511 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFNA | 181.414125072193 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 1176.3851 | 8.46 | 8.47 | Baseline | Baseline | 1872139.0611 | 44422.571 | 42.1439 | 247348.2236 | 8230.8797 | 30.0512 | 7.5688 | 5.3971 | 247530.2914 | 8256.1354 | 0.99435 | 0.98586 | 8.4815 | 8.4763 | 0.25423 | 0.021508 | 8.3138 | 8.2756 | 8.7433 | 8.5976 | 0.42955 | 0.322 | 0.12454 | 0.08626 | 71.0846 | 142.8298 | 0.0016796 | 0.0047115 | FALSE | 1.0006 | FALSE | [M-H]- | 45.3535 | 8.3734 | 8.5827 | 0.20926 | 8.3848 | 8.5706 | 0.18581 | 967.1934 | 0.92707 | 0.83103 | 1.718 | 0 | 25 | 9 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1872139.0611 | No Control Sample | 159143.3885 | off | off | 1872139.0611 | 159143.3885 | off | 0.0025 | off | 400 | off | 4705.5403 | off | 462.9645 | 2.7484 | 1.2724 | 418.971 | -5.1927 | -2.1756 | 0 | 0 | 0 | 0 | 0 | 0.4388 | FALSE | 100 | 42.1439
5512 | 5513 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFDA | 194.145209562208 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 8.87 | 8.86 | Baseline | Baseline | 1743733.8721 | 36180.435 | 48.1955 | 237254.4801 | 6848.764 | 34.6419 | 7.3496 | 5.2828 | 237308.8734 | 6856.4104 | 0.99775 | 0.99287 | 8.8906 | 8.8928 | 0.23211 | 0.020588 | 8.6716 | 8.7343 | 9.2203 | 9.0575 | 0.54863 | 0.3232 | 0.11926 | 0.082299 | 79.392 | 102.1341 | 0.00053515 | 0.00049375 | FALSE | 0.99975 | FALSE | [M-H]- | 48.5363 | 8.7881 | 9.0024 | 0.21431 | 8.7992 | 8.9887 | 0.18951 | 231.4255 | 1.1018 | 1.1988 | 1.2066 | 0 | 29 | 8 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Fail | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1743733.8721 | No Control Sample | 0 | off | off | 1743733.8721 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9617 | 3.1479 | 1.6147 | 468.9682 | -3.8258 | -1.7942 | 0 | 0 | 0 | 0 | 0 | 0.38569 | FALSE | 100 | 48.1955
5514 | 5515 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFUdA | 244.827208246826 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.21 | 9.2 | Baseline | Baseline | 1502203.9107 | 27667.6962 | 54.2945 | 228964.5625 | 5425.4534 | 42.2019 | 6.5609 | 5.0996 | 229274.6953 | 5427.0815 | 0.99101 | 0.9958 | 9.2302 | 9.2286 | 0.21906 | 0.020175 | 9.0687 | 9.077 | 9.3856 | 9.3754 | 0.31688 | 0.29843 | 0.10735 | 0.080772 | 110.7412 | 92.2094 | 0.0022187 | 0.00057819 | FALSE | 1.0002 | FALSE | [M-H]- | 61.2068 | 9.1345 | 9.3273 | 0.19279 | 9.1463 | 9.3169 | 0.17052 | 1603.0906 | 0.90834 | 0.80873 | 1.3447 | 0 | 18 | 8 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Fail | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1502203.9107 | No Control Sample | 0 | off | off | 1502203.9107 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9585 | 2.9605 | 1.6666 | 518.9647 | -5.0402 | -2.6157 | 0 | 0 | 0 | 0 | 0 | 0.56776 | FALSE | 100 | 54.2945
5516 | 5517 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFDoA | 163.679877351467 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.47 | 9.47 | Baseline | Baseline | 1486875.6864 | 35103.6218 | 42.3568 | 228831.4996 | 7754.6515 | 29.5089 | 6.4977 | 4.5268 | 228986.8641 | 7775.3297 | 0.99391 | 0.95009 | 9.4954 | 9.4935 | 0.26777 | 0.025357 | 9.3531 | 9.3595 | 9.704 | 9.6164 | 0.35095 | 0.25688 | 0.10638 | 0.071235 | 101.4192 | 100.0148 | 0.0015391 | 0.0052376 | FALSE | 1.0002 | FALSE | [M-H]- | 40.92 | 9.4077 | 9.5858 | 0.17805 | 9.4162 | 9.5768 | 0.16053 | 1003.5183 | 0.9764 | 0.94224 | 1.8681 | 0 | 25 | 9 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1486875.6864 | No Control Sample | 0 | off | off | 1486875.6864 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9539 | 0.34493 | 0.21143 | 568.9612 | -3.7796 | -2.1505 | 0 | 0 | 0 | 0 | 0 | 2.7732 | FALSE | 100 | 42.3568
5518 | 5519 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFTrDA | 195.10109126273 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.68 | 9.47 | Baseline | Baseline | 1291886.406 | 35103.6218 | 36.8021 | 214127.0118 | 7754.6515 | 27.6127 | 6.0333 | 4.5268 | 214264.5549 | 7775.3297 | 0.99226 | 0.95009 | 9.7127 | 9.4935 | 0.33797 | 0.032716 | 9.5279 | 9.3595 | 9.8874 | 9.6164 | 0.35952 | 0.25688 | 0.098182 | 0.071235 | 238.0039 | 100.0148 | 0.0012339 | 0.0052376 | FALSE | 1.0231 | FALSE | [M-H]- | 48.7753 | 9.6297 | 9.7912 | 0.16144 | 9.6381 | 9.7829 | 0.14478 | 734.9335 | 1.0742 | 1.1687 | 1.4541 | 0 | 32 | 11 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Fail | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1291886.406 | No Control Sample | 0 | off | off | 1291886.406 | 0 | off | 0 | off | 0 | off | 0 | off | 662.951 | 0.80729 | 0.5352 | 618.9586 | -5.0062 | -3.0986 | 0 | 0 | 0 | 0 | 0 | 2.0875 | FALSE | 100 | 36.8021
5520 | 5521 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFTeDA | 238.724985993811 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.87 | 9.86 | Baseline | Baseline | 1125768.2784 | 13300.7925 | 84.6392 | 212584.5974 | 3571.7473 | 59.5184 | 5.2956 | 3.7239 | 212894.8936 | 3571.7634 | 0.98942 | 0.99598 | 9.8893 | 9.8868 | 0.19591 | 0.019336 | 9.7511 | 9.7724 | 10.0504 | 10.0101 | 0.29932 | 0.23772 | 0.086533 | 0.057818 | 506.4097 | 695.424 | 0.0031715 | 9.0216e-06 | FALSE | 1.0003 | FALSE | [M-H]- | 59.6812 | 9.8204 | 9.9612 | 0.14082 | 9.8257 | 9.9543 | 0.12861 | 2252.4696 | 1.2061 | 1.4229 | 0.94978 | 0 | 35 | 11 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Fail | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1125768.2784 | No Control Sample | 0 | off | off | 1125768.2784 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9475 | 0.29359 | 0.20932 | 668.9559 | -2.9676 | -1.9852 | 0 | 0 | 0 | 0 | 0 | 1.1782 | FALSE | 100 | 84.6392
5527 | 5528 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFBS | 388.980416773173 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 3.5 | 3.5 | Baseline | Manual | 2895264.2893 | 28120.7263 | 102.9584 | 225994.6665 | 2251.3696 | 100.381 | 12.8112 | 12.4905 | 226501.1831 | 2256.8235 | 0.99157 | 0.99106 | 3.8331 | 3.8352 | 9.5169 | 0.33309 | 3.4767 | 3.5706 | 4.2027 | 4.1705 | 0.72605 | 0.5999 | 0.20145 | 0.20356 | 766.5785 | 508.5759 | 0.0043966 | 0.00070877 | FALSE | 0.99945 | TRUE | [M-H]- | 97.2451 | 3.6394 | 4.0691 | 0.42966 | 3.6636 | 4.0391 | 0.37546 | 1368.52 | 1.0786 | 1.1459 | 4.2348 | 0 | 114 | 33 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2895264.2893 | No Control Sample | 0 | off | off | 2895264.2893 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9435 | 1.5689 | 0.46901 | 79.956 | -14.5378 | -1.1624 | 0 | 0 | 0 | 0 | 0 | 1.3925 | FALSE | 100 | 102.9584
5529 | 5530 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFPeS | 288.269884813072 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 6.38 | 3.5 | Baseline | Manual | 2160428.748 | 28120.7263 | 76.8269 | 290361.3519 | 2251.3696 | 128.971 | 7.4405 | 12.4905 | 290684.7157 | 2256.8235 | 0.99164 | 0.99106 | 6.3892 | 3.8352 | 0.14356 | 0.0091592 | 6.2125 | 3.5706 | 6.6156 | 4.1705 | 0.40318 | 0.5999 | 0.1185 | 0.20356 | 173.8818 | 508.5759 | 0.0021571 | 0.00070877 | FALSE | 1.6659 | FALSE | [M-H]- | 72.0675 | 6.2909 | 6.4934 | 0.20244 | 6.3006 | 6.4806 | 0.18006 | 1553.502 | 0.85585 | 0.65743 | 4.2192 | 0 | 45 | 15 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Fail | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2160428.748 | No Control Sample | 0 | off | off | 2160428.748 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9401 | 0.79252 | 0.27654 | 79.956 | -15.2576 | -1.22 | 0 | 0 | 0 | 0 | 0 | 1.0253 | FALSE | 100 | 76.8269
5531 | 5532 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | L-PFHxS | 316.571639595783 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 324.4 | 4 | 81.1 | 0 | 7.34 | 7.35 | Baseline | Baseline | 1391325.7519 | 23528.445 | 59.1338 | 313405.3782 | 5492.0843 | 57.0649 | 4.4394 | 4.2841 | 319134.9847 | 5496.5401 | 0.61239 | 0.9886 | 7.3539 | 7.3513 | 0.18933 | 0.013897 | 7.2633 | 7.2213 | 7.5031 | 7.5641 | 0.2398 | 0.3429 | 0.07174 | 0.066793 | 67.1204 | 202.579 | 0.021691 | 0.0010944 | FALSE | 1.0004 | TRUE | [M-H]- | 97.5868 | 7.2882 | 7.4251 | 0.13688 | 7.2952 | 7.4163 | 0.12108 | -28348.9607 | 1.0481 | 1.077 | 7.528 | 0 | 30 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1391325.7519 | No Control Sample | 0 | off | off | 1391325.7519 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9369 | 0.72823 | 0.29052 | 79.956 | -15.4339 | -1.2341 | 0 | 0 | 0 | 0 | 0 | 1.2036 | FALSE | 81.1 | 59.1338
5533 | 5534 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | Br-PFHxS | 130.581286713856 | Qualifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 75.6 | 4 | 18.9 | 0 | 7.16 | 7.35 | Manual | Baseline | 779725.1822 | 23528.445 | 33.1397 | 139381.496 | 5492.0843 | 25.3786 | 5.5942 | 4.2841 | 319134.9847 | 5496.5401 | 0.34331 | 0.9886 | 7.1697 | 7.3513 | 0.13586 | 0.0097277 | 7.0546 | 7.2213 | 7.2679 | 7.5641 | 0.2133 | 0.3429 | 0.089549 | 0.066793 | 29.8506 | 202.579 | 0.0058974 | 0.0010944 | TRUE | 0.9753 | FALSE | [M-H]- | 172.7266 | 7.0835 | 7.2715 | 0.18799 | 7.0921 | 7.2476 | 0.15549 | 3853.6483 | 1.0385 | 0.89932 | 0.56042 | 0 | 26 | 12 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 779725.1822 | No Control Sample | 0 | off | off | 1391325.7519 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9369 | 0.72823 | 0.29052 | 79.9568 | -4.9266 | -0.39392 | 0 | 0 | 0 | 0 | 0 | 1.2036 | FALSE | 18.9 | 33.1397
5535 | 5536 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFHpS | 395.483472813811 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 7.98 | 7.35 | Baseline | Baseline | 1416180.9496 | 23528.445 | 60.1902 | 271114.9819 | 5492.0843 | 49.3647 | 5.2235 | 4.2841 | 271387.3639 | 5496.5401 | 0.99232 | 0.9886 | 7.9912 | 7.3513 | 0.14054 | 0.011215 | 7.8724 | 7.2213 | 8.4898 | 7.5641 | 0.61733 | 0.3429 | 0.08298 | 0.066793 | 118.7458 | 202.579 | 0.00095189 | 0.0010944 | FALSE | 1.0871 | TRUE | [M-H]- | 98.8709 | 7.9207 | 8.081 | 0.16028 | 7.9282 | 8.0705 | 0.14226 | -418.0417 | 1.1663 | 1.324 | 9.2874 | 0 | 41 | 9 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1416180.9496 | No Control Sample | 0 | off | off | 1416180.9496 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9338 | 0.82328 | 0.3696 | 79.9561 | -14.2559 | -1.1399 | 0 | 0 | 0 | 0 | 0 | 1.466 | FALSE | 100 | 60.1902
5537 | 5538 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | L-PFOS | 306.198170245395 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 315.2 | 4 | 78.8 | 0 | 8.46 | 8.45 | Manual | Baseline | 1159339.6355 | 14566.4179 | 79.5899 | 205740.8129 | 2524.2989 | 81.5041 | 5.635 | 5.7705 | 208750.2141 | 2529.7718 | 0.47939 | 0.9566 | 8.4755 | 8.4702 | 0.18275 | 0.01546 | 8.3993 | 8.325 | 8.5862 | 8.6479 | 0.18687 | 0.32297 | 0.091164 | 0.091702 | 30.3027 | 77.1033 | 0.015296 | 0.0051524 | TRUE | 1.0006 | TRUE | [M-H]- | 97.1441 | 8.389 | 8.5595 | 0.17046 | 8.389 | 8.5514 | 0.16237 | -16840.2157 | 1.0548 | 1.0095 | 7.7824 | 0 | 12 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1159339.6355 | No Control Sample | 0 | off | off | 1159339.6355 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9307 | 0.96437 | 0.48115 | 79.9561 | -14.5959 | -1.1671 | 0 | 0 | 0 | 0 | 0 | 0.66186 | FALSE | 78.8 | 79.5899
5539 | 5540 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | Br-PFOS | 135.23658807554 | Qualifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 84.4 | 4 | 21.1 | 0 | 8.3 | 8.45 | Manual | Baseline | 1096186.2102 | 14566.4179 | 75.2543 | 180483.9806 | 2524.2989 | 71.4987 | 6.0736 | 5.7705 | 208750.2141 | 2529.7718 | 0.45342 | 0.9566 | 8.3253 | 8.4702 | 0.30475 | 0.025294 | 8.2108 | 8.325 | 8.4027 | 8.6479 | 0.19195 | 0.32297 | 0.10046 | 0.091702 | 26.5828 | 77.1033 | 0.029061 | 0.0051524 | TRUE | 0.98289 | FALSE | [M-H]- | 160.2329 | 8.24 | 8.4073 | 0.16732 | 8.2463 | 8.4073 | 0.16102 | 27325.1665 | 0.88524 | 0.82559 | 0.94553 | 0 | 13 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1096186.2102 | No Control Sample | 0 | off | off | 1159339.6355 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9307 | 0.96437 | 0.48115 | 79.9561 | -14.8241 | -1.1853 | 0 | 0 | 0 | 0 | 0 | 0.66186 | FALSE | 21.1 | 75.2543
5541 | 5542 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFNS | 398.819194028263 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 400 | 4 | 100 | 0 | 8.85 | 8.45 | Baseline | Baseline | 1287623.6579 | 14566.4179 | 88.3967 | 229972.4137 | 2524.2989 | 91.1035 | 5.599 | 5.7705 | 230022.1175 | 2529.7718 | 0.99632 | 0.9566 | 8.8696 | 8.4702 | 0.22153 | 0.019605 | 8.7114 | 8.325 | 9.2052 | 8.6479 | 0.49378 | 0.32297 | 0.088504 | 0.091702 | 117.1628 | 77.1033 | 0.0002132 | 0.0051524 | FALSE | 1.0471 | TRUE | [M-H]- | 99.7048 | 8.7876 | 8.9688 | 0.18121 | 8.7942 | 8.9535 | 0.15933 | -99.2948 | 1.1093 | 1.123 | 7.416 | 0 | 26 | 7 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1287623.6579 | No Control Sample | 0 | off | off | 1287623.6579 | 0 | off | 0 | off | 0 | off | 0 | off | 548.927 | -0.080014 | -0.043922 | 79.956 | -13.829 | -1.1057 | 0 | 0 | 0 | 0 | 0 | 1.7539 | FALSE | 100 | 88.3967
5543 | 5544 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | PFDS | 403.77664976258 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 400 | 4 | 100 | 0 | 9.19 | 8.45 | Baseline | Baseline | 1339424.3705 | 14566.4179 | 91.9529 | 255653.593 | 2524.2989 | 101.2771 | 5.2392 | 5.7705 | 255810.3493 | 2529.7718 | 0.99219 | 0.9566 | 9.1978 | 8.4702 | 0.084604 | 0.0077751 | 9.036 | 8.325 | 9.3733 | 8.6479 | 0.33728 | 0.32297 | 0.081579 | 0.091702 | 145.8703 | 77.1033 | 0.0011619 | 0.0051524 | FALSE | 1.0859 | TRUE | [M-H]- | 100.9442 | 9.1196 | 9.2904 | 0.17072 | 9.131 | 9.2782 | 0.14717 | 880.7187 | 0.98785 | 0.9619 | 6.3243 | 0 | 19 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1339424.3705 | No Control Sample | 0 | off | off | 1339424.3705 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9234 | -0.66576 | -0.39874 | 79.956 | -11.0705 | -0.88517 | 0 | 0 | 0 | 0 | 0 | 1.4294 | FALSE | 100 | 91.9529
5547 | 5548 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FBSA | 479.73841988534 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 4.6 | 8.86 | Manual | Baseline | 4919396.4088 | 38889.6749 | 126.4962 | 351643.799 | 8027.8612 | 43.8029 | 13.9897 | 4.8443 | 351919.4767 | 8028.931 | 0.99712 | 0.97954 | 4.9626 | 9.0186 | 7.8825 | 0.3626 | 4.6267 | 8.8571 | 5.3537 | 9.2117 | 0.72706 | 0.35458 | 0.22116 | 0.075769 | 456.6344 | 76.665 | 0.00055598 | 0.0002572 | TRUE | 0.55026 | FALSE | [M-H]- | 119.9346 | 4.7426 | 5.1912 | 0.44859 | 4.7709 | 5.158 | 0.38716 | 268.902 | 1.1772 | 1.3853 | 2.365 | 0 | 136 | 43 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Fail | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 4919396.4088 | No Control Sample | 0 | off | off | 4919396.4088 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9595 | 1.7428 | 0.51928 | 77.964 | -16.8465 | -1.3134 | 0 | 0 | 0 | 0 | 0 | 1.1832 | FALSE | 100 | 126.4962
5549 | 5550 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FPeSA | 199.083808918984 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 7 | 8.86 | Baseline | Baseline | 3170221.4101 | 38889.6749 | 81.5183 | 379276.7048 | 8027.8612 | 47.2451 | 8.3586 | 4.8443 | 379644.5706 | 8028.931 | 0.99037 | 0.97954 | 7.108 | 9.0186 | 1.5426 | 0.10798 | 6.9268 | 8.8571 | 7.4181 | 9.2117 | 0.49135 | 0.35458 | 0.13794 | 0.075769 | 113.45 | 76.665 | 0.0026616 | 0.0002572 | FALSE | 0.78815 | FALSE | [M-H]- | 49.771 | 7.0017 | 7.2167 | 0.21494 | 7.0113 | 7.204 | 0.19273 | 2054.5372 | 1.1998 | 1.4076 | 3.0214 | 0 | 56 | 17 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Fail | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 3170221.4101 | No Control Sample | 0 | off | off | 3170221.4101 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9559 | 0.35776 | 0.12448 | 77.9636 | -21.6493 | -1.6879 | 0 | 0 | 0 | 0 | 0 | 1.0715 | FALSE | 100 | 81.5183
5551 | 5552 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FHxSA | 241.728188146062 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 7.86 | 8.86 | Baseline | Baseline | 1920880.4536 | 38889.6749 | 49.3931 | 305522.2134 | 8027.8612 | 38.0577 | 6.2872 | 4.8443 | 305800.6093 | 8028.931 | 0.99199 | 0.97954 | 7.9964 | 9.0186 | 1.7353 | 0.1364 | 7.7779 | 8.8571 | 8.2991 | 9.2117 | 0.52123 | 0.35458 | 0.10186 | 0.075769 | 72.7092 | 76.665 | 0.0019507 | 0.0002572 | FALSE | 0.88665 | FALSE | [M-H]- | 60.432 | 7.9149 | 8.0943 | 0.17939 | 7.924 | 8.0832 | 0.15919 | 1143.4197 | 1.0528 | 1.0903 | 6.9979 | 0 | 38 | 10 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Fail | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1920880.4536 | No Control Sample | 0 | off | off | 1920880.4536 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9529 | 0.83604 | 0.3327 | 77.9639 | -19.2336 | -1.4996 | 0 | 0 | 0 | 0 | 0 | 1.4813 | FALSE | 100 | 49.3931
5553 | 5554 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FHpSA | 143.636209094396 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 8.58 | 8.86 | Baseline | Baseline | 2364147.3592 | 38889.6749 | 60.7911 | 354454.1703 | 8027.8612 | 44.153 | 6.6698 | 4.8443 | 357640.3297 | 8028.931 | 0.96334 | 0.97954 | 8.5815 | 9.0186 | 0.017057 | 0.0014635 | 8.4473 | 8.8571 | 8.7459 | 9.2117 | 0.29867 | 0.35458 | 0.10964 | 0.075769 | 51.4036 | 76.665 | 0.0051643 | 0.0002572 | FALSE | 0.95153 | FALSE | [M-H]- | 35.9091 | 8.4876 | 8.6739 | 0.18631 | 8.4946 | 8.6658 | 0.17118 | -6128.9521 | 1.069 | 1.1358 | 4.5867 | 0 | 18 | 8 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2364147.3592 | No Control Sample | 0 | off | off | 2364147.3592 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9496 | 0.52124 | 0.23349 | 77.9636 | -22.7399 | -1.7729 | 0 | 0 | 0 | 0 | 0 | 1.4985 | FALSE | 100 | 60.7911
5555 | 5556 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | FOSA | 206.678030015852 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 8.88 | 8.86 | Baseline | Baseline | 1828833.2015 | 38889.6749 | 47.0262 | 285715.0516 | 8027.8612 | 35.5904 | 6.4009 | 4.8443 | 285969.4225 | 8028.931 | 0.98857 | 0.97954 | 9.0173 | 9.0186 | 1.5461 | 0.13729 | 8.8268 | 8.8571 | 9.2432 | 9.2117 | 0.41648 | 0.35458 | 0.10453 | 0.075769 | 87.6318 | 76.665 | 0.001879 | 0.0002572 | FALSE | 0.99985 | FALSE | [M-H]- | 51.6695 | 8.9227 | 9.1135 | 0.1908 | 8.9321 | 9.1026 | 0.17047 | 1289.0031 | 1.0349 | 1.0595 | 6.0897 | 0 | 22 | 7 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1828833.2015 | No Control Sample | 0 | off | off | 1828833.2015 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9467 | 0.99766 | 0.49678 | 77.9637 | -21.7637 | -1.6968 | 0 | 0 | 0 | 0 | 0 | 31.3979 | FALSE | 100 | 47.0262
5574 | 5575 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 4:2 FTS | 316.485806595705 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 5.94 | 5.92 | Baseline | Baseline | 1609058.1828 | 10013.0929 | 160.6954 | 235119.4105 | 1712.3053 | 137.3116 | 6.8436 | 5.8477 | 235262.0725 | 1722.9318 | 0.99515 | 0.97573 | 5.9507 | 5.9486 | 0.18008 | 0.010697 | 5.726 | 5.8088 | 6.177 | 6.0887 | 0.45107 | 0.27982 | 0.10927 | 0.084688 | 167.8298 | 153.851 | 0.0010272 | 0.0014172 | FALSE | 1.0003 | FALSE | [M-H]- | 79.1215 | 5.8435 | 6.0543 | 0.21075 | 5.8593 | 6.0423 | 0.18304 | 535.4154 | 0.95816 | 0.94241 | 1.5838 | 0 | 56 | 15 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Fail | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1609058.1828 | No Control Sample | 0 | off | off | 1609058.1828 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9749 | 1.7267 | 0.56457 | 306.9667 | -3.0025 | -0.92168 | 0 | 0 | 0 | 0 | 0 | 9.7158 | FALSE | 100 | 160.6954
5576 | 5577 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 6:2 FTS | 245.162377968858 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 7.92 | 7.92 | Baseline | Baseline | 999416.7257 | 6718.9324 | 148.7464 | 210093.6253 | 1519.3681 | 138.277 | 4.757 | 4.4222 | 210148.7704 | 1519.3681 | 0.99653 | 0.98852 | 7.9288 | 7.928 | 0.11164 | 0.0088422 | 7.7514 | 7.7135 | 8.2559 | 8.1674 | 0.50448 | 0.45383 | 0.073959 | 0.066358 | 95.3616 | 89.9881 | 0.00071617 | 1.2024e-13 | FALSE | 1.0001 | FALSE | [M-H]- | 61.2906 | 7.8512 | 8.0026 | 0.15135 | 7.8624 | 7.9915 | 0.12912 | 298.253 | 0.90936 | 0.7914 | 2.1682 | 0 | 37 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Fail | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 999416.7257 | No Control Sample | 0 | off | off | 999416.7257 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9697 | 4.274 | 1.8249 | 406.9599 | -3.9587 | -1.6111 | 0 | 0 | 0 | 0 | 0 | 1.005 | FALSE | 100 | 148.7464
5578 | 5579 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 8:2 FTS | 243.169279834803 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | Accuracy, Total Width | 0 | 0 | 400 | 4 | 100 | 0 | 8.88 | 8.89 | Baseline | Baseline | 935417.8423 | 7840.1314 | 119.3115 | 165519.5648 | 1315.8738 | 125.7868 | 5.6514 | 5.9581 | 165535.1871 | 1315.8739 | 0.99861 | 0.99756 | 8.8842 | 8.8853 | 0.047801 | 0.0042447 | 8.7122 | 8.7019 | 9.2264 | 9.1594 | 0.51413 | 0.45753 | 0.087998 | 0.093912 | 94.2114 | 95.7815 | 0.00025383 | 4.5664e-08 | FALSE | 0.99988 | FALSE | [M-H]- | 60.7923 | 8.7947 | 8.9806 | 0.18591 | 8.8075 | 8.9689 | 0.1614 | 81.7169 | 0.96101 | 0.92305 | 2.068 | 0 | 27 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Fail | Pass | Fail | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 935417.8423 | No Control Sample | 0 | off | off | 935417.8423 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9635 | 3.7533 | 1.9778 | 506.9539 | -3.541 | -1.7951 | 0 | 0 | 0 | 0 | 0 | 1.248 | FALSE | 100 | 119.3115
5580 | 5581 | PFAS CS9 100ng/ml | 40 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 41) | CS9 | 0 | 0 | Standard | 0 | 0 | 0 | 10 | 1 | 10:2 FTS | 200.629288138653 | Quantifiers | 10/27/2023 05:55:18 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | Accuracy | 0 | 0 | 400 | 4 | 100 | 0 | 9.5 | 8.89 | Baseline | Baseline | 803434.9364 | 7840.1314 | 102.4772 | 150313.2452 | 1315.8738 | 114.2307 | 5.3451 | 5.9581 | 150371.5024 | 1315.8739 | 0.99484 | 0.99756 | 9.5115 | 8.8853 | 0.1206 | 0.011457 | 9.3745 | 8.7019 | 9.7297 | 9.1594 | 0.35518 | 0.45753 | 0.085595 | 0.093912 | 215.0604 | 95.7815 | 0.0010093 | 4.5664e-08 | FALSE | 1.0705 | FALSE | [M-H]- | 50.1573 | 9.4298 | 9.5925 | 0.16273 | 9.44 | 9.5827 | 0.14273 | 427.1346 | 1.128 | 1.3058 | 2.6798 | 0 | 26 | 8 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Fail | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 803434.9364 | No Control Sample | 0 | off | off | 803434.9364 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9557 | 0.97986 | 0.61433 | 606.9467 | -2.0051 | -1.217 | 0 | 0 | 0 | 0 | 0 | 2.5788 | FALSE | 100 | 102.4772
5641 | 5642 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFBA | 97.4267038931732 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 1.49 | 1.48 | Manual | Baseline | 1502472.0187 | 85973.9445 | 17.4759 | 195646.7011 | 11177.5847 | 17.5035 | 7.6795 | 7.6916 | 196079.8562 | 11198.451 | 0.98958 | 0.99008 | 1.4634 | 1.4639 | 1.7886 | 0.02665 | 1.3371 | 1.3097 | 1.723 | 1.6921 | 0.38583 | 0.38237 | 0.11206 | 0.11956 | 196.9863 | 189.1084 | 0.00096171 | 0.0044845 | TRUE | 0.9996 | TRUE | [M-H]- | 97.4267 | 1.3697 | 1.633 | 0.2633 | 1.3776 | 1.6096 | 0.23196 | 487.6604 | 1.4784 | 1.8601 | 1.1163 | 0 | 80 | 25 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1502472.0187 | No Control Sample | 0 | off | off | 1502472.0187 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9802 | 4.7928 | 1.0208 | 168.9878 | -6.9663 | -1.1772 | 0 | 0 | 0 | 0 | 0 | 0.40041 | FALSE | 25 | 17.4759
5643 | 5644 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFPeA | 103.2341749602 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 3.18 | 3.21 | Manual | Baseline | 1487755.1286 | 64440.6531 | 23.0872 | 150441.1468 | 6229.3782 | 24.1503 | 9.8893 | 10.3446 | 150745.8439 | 6286.7538 | 0.98943 | 0.96749 | 3.2059 | 3.2047 | 0.81481 | 0.025911 | 2.9764 | 2.9886 | 3.5088 | 3.4551 | 0.53242 | 0.4665 | 0.15191 | 0.15877 | 1042.3743 | 531.1819 | 0.000835 | 0.009586 | TRUE | 1.0004 | TRUE | [M-H]- | 103.2342 | 3.0525 | 3.4067 | 0.35426 | 3.0729 | 3.3731 | 0.30022 | 235.9375 | 1.177 | 1.308 | 1.2975 | 0 | 95 | 30 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1487755.1286 | No Control Sample | 0 | off | off | 1487755.1286 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9771 | 4.2967 | 1.1299 | 218.9855 | -1.5708 | -0.34398 | 0 | 0 | 0 | 0 | 0 | 0.21471 | FALSE | 25 | 23.0872
5645 | 5646 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFHxA | 85.3766725594973 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 1058.7184 | 6.13 | 6.12 | Baseline | Baseline | 2158129.6227 | 108615.1139 | 19.8695 | 286720.1766 | 17752.7547 | 16.1507 | 7.527 | 6.1182 | 286895.3694 | 17779.4224 | 0.99496 | 0.99113 | 6.1369 | 6.1342 | 0.11306 | 0.0069304 | 5.9459 | 5.9732 | 6.4147 | 6.3261 | 0.46888 | 0.3529 | 0.12194 | 0.093731 | 196.1938 | 220.7958 | 0.0012498 | 0.0022066 | FALSE | 1.0004 | FALSE | [M-H]- | 85.3767 | 6.0297 | 6.2439 | 0.21414 | 6.0415 | 6.232 | 0.1905 | 764.2503 | 0.94688 | 0.87978 | 0.86251 | 0 | 52 | 16 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Pass | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2158129.6227 | No Control Sample | 203843.5852 | off | off | 2158129.6227 | 203843.5852 | off | 0.01 | off | 100 | off | 1058.7184 | off | 312.9735 | 2.0626 | 0.64554 | 268.9811 | -4.7743 | -1.2842 | 0 | 0 | 0 | 0 | 0 | 1.4476 | FALSE | 25 | 19.8695
5647 | 5648 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFHpA | 66.3139625671642 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 778.1086 | 7.26 | 7.27 | Baseline | Baseline | 1776645.8763 | 107069.8885 | 16.5933 | 271555.1324 | 23693.8582 | 11.461 | 6.5425 | 4.5189 | 271682.6909 | 23717.6284 | 0.99579 | 0.99367 | 7.278 | 7.2772 | 0.24777 | 0.017988 | 7.1279 | 7.1396 | 7.5513 | 7.4054 | 0.42337 | 0.2658 | 0.10665 | 0.072921 | 125.8406 | 204.0925 | 0.00083388 | 0.0017688 | FALSE | 1.0001 | FALSE | [M-H]- | 66.314 | 7.1931 | 7.3617 | 0.16867 | 7.2018 | 7.3534 | 0.15152 | 534.8688 | 1.2009 | 1.4658 | 1.2299 | 0 | 48 | 15 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1776645.8763 | No Control Sample | 228328.7703 | off | off | 1776645.8763 | 228328.7703 | off | 0.01 | off | 100 | off | 778.1086 | off | 362.9704 | 2.0385 | 0.73991 | 318.9778 | -3.985 | -1.2711 | 0 | 0 | 0 | 0 | 0 | 1.0251 | FALSE | 25 | 16.5933
5649 | 5650 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFOA | 77.2731278165113 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 804.7111 | 7.96 | 7.94 | Baseline | Baseline | 1540823.8802 | 82219.6816 | 18.7403 | 274309.2656 | 18450.4964 | 14.8673 | 5.6171 | 4.4562 | 274591.5882 | 18450.9917 | 0.99344 | 0.99822 | 7.964 | 7.9659 | 0.04987 | 0.0039697 | 7.8147 | 7.8203 | 8.3128 | 8.2253 | 0.49815 | 0.405 | 0.090138 | 0.067766 | 87.5845 | 104.6265 | 0.00045295 | 5.7777e-05 | FALSE | 0.99976 | FALSE | [M-H]- | 77.2731 | 7.8914 | 8.0581 | 0.16665 | 7.8991 | 8.0459 | 0.14687 | -249.4185 | 1.18 | 1.3332 | 1.2078 | 0 | 36 | 10 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1540823.8802 | No Control Sample | 191475.4056 | off | off | 1540823.8802 | 191475.4056 | off | 0.01 | off | 100 | off | 804.7111 | off | 412.9674 | 2.2637 | 0.93484 | 368.9749 | -3.1678 | -1.1688 | 0 | 0 | 0 | 0 | 0 | 0.90412 | FALSE | 25 | 18.7403
5651 | 5652 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFNA | 81.2129588297356 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 886.4111 | 8.46 | 8.47 | Baseline | Baseline | 1410664.6557 | 74771.3352 | 18.8664 | 228854.12 | 14428.3299 | 15.8614 | 6.164 | 5.1823 | 228933.0145 | 14432.0886 | 0.99666 | 0.99685 | 8.4821 | 8.4791 | 0.26168 | 0.022138 | 8.3137 | 8.2589 | 8.7272 | 8.6461 | 0.41352 | 0.38723 | 0.099939 | 0.081625 | 137.219 | 149.5402 | 0.00073038 | 0.00045154 | FALSE | 1.0004 | FALSE | [M-H]- | 81.213 | 8.3881 | 8.5702 | 0.18211 | 8.3966 | 8.5598 | 0.16312 | 404.2143 | 0.92079 | 0.80516 | 0.71216 | 0 | 24 | 7 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1410664.6557 | No Control Sample | 159143.3885 | off | off | 1410664.6557 | 159143.3885 | off | 0.01 | off | 100 | off | 886.4111 | off | 462.9645 | 2.7383 | 1.2677 | 418.9719 | -3.0857 | -1.2928 | 0 | 0 | 0 | 0 | 0 | 4.6984 | FALSE | 25 | 18.8664
5653 | 5654 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFDA | 81.6344228072246 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy, Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 8.87 | 8.86 | Baseline | Baseline | 1208310.5384 | 59624.6025 | 20.2653 | 207121.7278 | 11722.4286 | 17.6688 | 5.8338 | 5.0864 | 207161.3895 | 11726.2891 | 0.99783 | 0.99667 | 8.8944 | 8.8947 | 0.2751 | 0.024401 | 8.7095 | 8.7368 | 9.2217 | 9.0771 | 0.51223 | 0.34035 | 0.092994 | 0.079438 | 125.9368 | 114.091 | 0.00051509 | 0.00064509 | FALSE | 0.99996 | FALSE | [M-H]- | 81.6344 | 8.8054 | 8.9902 | 0.18483 | 8.8139 | 8.9784 | 0.16454 | 208.2771 | 1.095 | 1.167 | 0.6337 | 0 | 27 | 7 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Fail | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1208310.5384 | No Control Sample | 0 | off | off | 1208310.5384 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9607 | 1.3605 | 0.69786 | 468.9677 | -4.8964 | -2.2963 | 0 | 0 | 0 | 0 | 0 | 3.64 | FALSE | 25 | 20.2653
5655 | 5656 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFUdA | 101.260981837135 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.21 | 9.2 | Baseline | Baseline | 1028399.4893 | 45795.5743 | 22.4563 | 189212.569 | 9538.5145 | 19.8367 | 5.4352 | 4.8011 | 189355.6878 | 9542.282 | 0.99452 | 0.99516 | 9.2324 | 9.2324 | 0.24367 | 0.022442 | 9.0698 | 9.0595 | 9.3879 | 9.3944 | 0.31813 | 0.33493 | 0.086131 | 0.074527 | 148.1374 | 190.3239 | 0.0014392 | 0.00032833 | FALSE | 1 | TRUE | [M-H]- | 101.261 | 9.1501 | 9.3181 | 0.16801 | 9.1601 | 9.3087 | 0.14862 | 856.0037 | 1.1581 | 1.3767 | 0.64288 | 0 | 18 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1028399.4893 | No Control Sample | 0 | off | off | 1028399.4893 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9588 | 3.5465 | 1.9965 | 518.9654 | -3.7425 | -1.9422 | 0 | 0 | 0 | 0 | 0 | 0.42263 | FALSE | 25 | 22.4563
5657 | 5658 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFDoA | 84.1032807207938 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 0 | 9.47 | 9.47 | Baseline | Baseline | 1062357.4041 | 48812.4121 | 21.7641 | 203050.9445 | 10760.8349 | 18.8694 | 5.232 | 4.5361 | 203105.5271 | 10762.6117 | 0.99712 | 0.98459 | 9.4987 | 9.4963 | 0.30298 | 0.028692 | 9.3553 | 9.3618 | 9.7186 | 9.7111 | 0.36325 | 0.34927 | 0.084691 | 0.072727 | 246.9408 | 206.4391 | 0.00051759 | 0.00041796 | FALSE | 1.0002 | FALSE | [M-H]- | 84.1033 | 9.4228 | 9.5771 | 0.15425 | 9.4289 | 9.5682 | 0.1393 | 289.3246 | 0.97647 | 0.90898 | 1.2002 | 0 | 26 | 7 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1062357.4041 | No Control Sample | 0 | off | off | 1062357.4041 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9542 | 0.91453 | 0.56057 | 568.9621 | -2.2798 | -1.2971 | 0 | 0 | 0 | 0 | 0 | 2.6063 | FALSE | 25 | 21.7641
5659 | 5660 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFTrDA | 100.232663889115 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.68 | 9.47 | Baseline | Baseline | 922894.899 | 48812.4121 | 18.907 | 199342.7977 | 10760.8349 | 18.5248 | 4.6297 | 4.5361 | 199441.8487 | 10762.6117 | 0.99483 | 0.98459 | 9.7138 | 9.4963 | 0.34893 | 0.033776 | 9.5451 | 9.3618 | 9.8827 | 9.7111 | 0.3376 | 0.34927 | 0.075801 | 0.072727 | 461.2774 | 206.4391 | 0.0010128 | 0.00041796 | FALSE | 1.0229 | TRUE | [M-H]- | 100.2327 | 9.6469 | 9.7842 | 0.1373 | 9.6539 | 9.7772 | 0.12325 | 598.0036 | 1.1224 | 1.2762 | 0.92816 | 0 | 30 | 8 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 922894.899 | No Control Sample | 0 | off | off | 922894.899 | 0 | off | 0 | off | 0 | off | 0 | off | 662.951 | 0.81344 | 0.53927 | 618.9591 | -4.2201 | -2.6121 | 0 | 0 | 0 | 0 | 0 | 1.9805 | FALSE | 25 | 18.907
5661 | 5662 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFTeDA | 101.739150809896 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 9.87 | 9.86 | Baseline | Baseline | 693769.1946 | 19233.2773 | 36.0713 | 179391.8343 | 5024.2498 | 35.7052 | 3.8673 | 3.8281 | 179576.598 | 5027.2313 | 0.9922 | 0.99328 | 9.8958 | 9.8948 | 0.26158 | 0.025818 | 9.7641 | 9.7959 | 10.0135 | 10.0144 | 0.24938 | 0.21845 | 0.061723 | 0.060383 | 964.9146 | 579.3749 | 0.0018458 | 0.0011811 | FALSE | 1.0001 | TRUE | [M-H]- | 101.7392 | 9.8369 | 9.9566 | 0.11975 | 9.8444 | 9.949 | 0.10459 | 1327.728 | 1.0624 | 1.1413 | 0.9461 | 0 | 28 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Fail | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 693769.1946 | No Control Sample | 0 | off | off | 693769.1946 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9478 | 0.70059 | 0.49949 | 668.956 | -2.9123 | -1.9482 | 0 | 0 | 0 | 0 | 0 | 3.332 | FALSE | 25 | 36.0713
5668 | 5669 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFBS | 101.267397822563 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 3.5 | 3.5 | Manual | Manual | 1003006.6871 | 37419.6943 | 26.8042 | 88349.7342 | 3325.242 | 26.5694 | 11.3527 | 11.2532 | 88477.297 | 3336.7313 | 0.99394 | 0.98699 | 3.8522 | 3.853 | 10.0638 | 0.35223 | 3.599 | 3.6055 | 4.2074 | 4.1769 | 0.60834 | 0.57141 | 0.17314 | 0.17906 | 899.2833 | 816.5499 | 0.00091038 | 0.00076017 | TRUE | 0.9998 | TRUE | [M-H]- | 101.2674 | 3.6729 | 4.0733 | 0.40044 | 3.6976 | 4.0396 | 0.34195 | -132.2157 | 1.1249 | 1.2313 | 5.0897 | 0 | 96 | 29 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1003006.6871 | No Control Sample | 0 | off | off | 1003006.6871 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9435 | 1.6222 | 0.48493 | 79.9565 | -9.3082 | -0.74426 | 0 | 0 | 0 | 0 | 0 | 0.86992 | FALSE | 25 | 26.8042
5670 | 5671 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFPeS | 95.9669338699318 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 6.38 | 3.5 | Baseline | Manual | 957053.0113 | 37419.6943 | 25.5762 | 173682.65 | 3325.242 | 52.2316 | 5.5104 | 11.2532 | 173810.4669 | 3336.7313 | 0.99274 | 0.98699 | 6.3881 | 3.853 | 0.1274 | 0.008128 | 6.2176 | 3.6055 | 6.6071 | 4.1769 | 0.38953 | 0.57141 | 0.08734 | 0.17906 | 312.059 | 816.5499 | 0.0016467 | 0.00076017 | FALSE | 1.658 | TRUE | [M-H]- | 95.9669 | 6.3029 | 6.4781 | 0.17519 | 6.3125 | 6.4664 | 0.15386 | 734.2417 | 1.0267 | 1.0323 | 3.6396 | 0 | 43 | 11 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 957053.0113 | No Control Sample | 0 | off | off | 957053.0113 | 0 | off | 0 | off | 0 | off | 0 | off | 348.94 | 0.59988 | 0.20932 | 79.9562 | -12.1707 | -0.97314 | 0 | 0 | 0 | 0 | 0 | 1.6147 | FALSE | 25 | 25.5762
5672 | 5673 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | L-PFHxS | 84.3352389640229 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 81.1 | 4 | 20.275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 677703.4293 | 43019.6637 | 15.7533 | 163891.2537 | 10305.2213 | 15.9037 | 4.1351 | 4.1746 | 165997.1209 | 10321.0741 | 0.72322 | 0.98576 | 7.3519 | 7.3514 | 0.16161 | 0.011862 | 7.265 | 7.232 | 7.4869 | 7.4886 | 0.22188 | 0.25668 | 0.0648 | 0.066365 | 98.0823 | 166.4716 | 0.013907 | 0.0024555 | FALSE | 1.0001 | TRUE | [M-H]- | 103.9892 | 7.2885 | 7.4221 | 0.13361 | 7.2961 | 7.413 | 0.11693 | -10271.9571 | 1.1406 | 1.2913 | 5.9297 | 0 | 28 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 677703.4293 | No Control Sample | 0 | off | off | 677703.4293 | 0 | off | 0 | off | 0 | off | 0 | off | 398.937 | 0.98554 | 0.39317 | 79.9562 | -12.0124 | -0.96048 | 0 | 0 | 0 | 0 | 0 | 1.2764 | FALSE | 20.275 | 15.7533
5674 | 5675 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | Br-PFHxS | 20.2056992367306 | Qualifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 18.9 | 4 | 4.725 | 0 | 7.16 | 7.35 | Manual | Baseline | 220601.4113 | 43019.6637 | 5.1279 | 40445.3965 | 10305.2213 | 3.9247 | 5.4543 | 4.1746 | 165997.1209 | 10321.0741 | 0.23549 | 0.98576 | 7.1693 | 7.3514 | 0.13042 | 0.009338 | 7.0539 | 7.232 | 7.2652 | 7.4886 | 0.21135 | 0.25668 | 0.086058 | 0.066365 | 24.2049 | 166.4716 | 0.034512 | 0.0024555 | TRUE | 0.97523 | TRUE | [M-H]- | 106.9085 | 7.0853 | 7.2732 | 0.18798 | 7.0933 | 7.2468 | 0.15349 | 6604.5413 | 1.0448 | 0.87482 | 0.32551 | 0 | 26 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 220601.4113 | No Control Sample | 0 | off | off | 677703.4293 | 0 | off | 0 | off | 0 | off | 0 | off | 398.937 | 0.98554 | 0.39317 | 79.9626 | 67.879 | 5.4274 | 0 | 0 | 0 | 0 | 0 | 1.2764 | FALSE | 4.725 | 5.1279
5676 | 5677 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFHpS | 101.562419010774 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7.98 | 7.35 | Baseline | Baseline | 664962.6169 | 43019.6637 | 15.4572 | 146567.2485 | 10305.2213 | 14.2226 | 4.5369 | 4.1746 | 146672.423 | 10321.0741 | 0.99309 | 0.98576 | 7.991 | 7.3514 | 0.13736 | 0.010962 | 7.8701 | 7.232 | 8.2373 | 7.4886 | 0.36722 | 0.25668 | 0.070853 | 0.066365 | 145.255 | 166.4716 | 0.00033818 | 0.0024555 | FALSE | 1.087 | TRUE | [M-H]- | 101.5624 | 7.9235 | 8.0736 | 0.15011 | 7.9325 | 8.0623 | 0.12983 | 134.9794 | 1.1749 | 1.3657 | 5.9303 | 0 | 27 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 664962.6169 | No Control Sample | 0 | off | off | 664962.6169 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9338 | 0.88017 | 0.39514 | 79.9562 | -12.2113 | -0.97638 | 0 | 0 | 0 | 0 | 0 | 1.6751 | FALSE | 25 | 15.4572
5678 | 5679 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | L-PFOS | 83.9918460290456 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 78.8 | 4 | 19.7 | 0 | 8.46 | 8.45 | Manual | Baseline | 537018.5961 | 24597.8343 | 21.8319 | 102649.5467 | 4786.121 | 21.4473 | 5.2316 | 5.1394 | 102951.1984 | 4798.7274 | 0.61829 | 0.97913 | 8.4779 | 8.4782 | 0.21128 | 0.017874 | 8.3991 | 8.3075 | 8.5924 | 8.6016 | 0.19335 | 0.29405 | 0.082877 | 0.081625 | 39.4436 | 120.2213 | 0.0046648 | 0.0041047 | TRUE | 0.99996 | TRUE | [M-H]- | 106.5886 | 8.3892 | 8.5596 | 0.17039 | 8.3892 | 8.5499 | 0.16068 | -2476.5005 | 1.0479 | 0.97636 | 6.292 | 0 | 12 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 537018.5961 | No Control Sample | 0 | off | off | 537018.5961 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9303 | 0.080436 | 0.040132 | 79.9563 | -12.6174 | -1.0089 | 0 | 0 | 0 | 0 | 0 | 0.67957 | FALSE | 19.7 | 21.8319
5680 | 5681 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | Br-PFOS | 21.5375661080417 | Qualifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 21.1 | 4 | 5.275 | 0 | 8.3 | 8.45 | Manual | Baseline | 294802.2938 | 24597.8343 | 11.9849 | 49693.8079 | 4786.121 | 10.3829 | 5.9324 | 5.1394 | 102951.1984 | 4798.7274 | 0.33953 | 0.97913 | 8.3274 | 8.4782 | 0.33018 | 0.027405 | 8.2159 | 8.3075 | 8.4008 | 8.6016 | 0.18487 | 0.29405 | 0.099012 | 0.081625 | 19.0951 | 120.2213 | 0.023542 | 0.0041047 | TRUE | 0.98221 | TRUE | [M-H]- | 102.0738 | 8.2415 | 8.4077 | 0.16613 | 8.2494 | 8.4077 | 0.15822 | 6327.929 | 0.89306 | 0.85916 | 0.54896 | 0 | 13 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 294802.2938 | No Control Sample | 0 | off | off | 537018.5961 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9303 | 0.080436 | 0.040132 | 79.9575 | 2.7146 | 0.21705 | 0 | 0 | 0 | 0 | 0 | 0.67957 | FALSE | 5.275 | 11.9849
5682 | 5683 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFNS | 106.925800185605 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 8.85 | 8.45 | Baseline | Baseline | 582961.0123 | 24597.8343 | 23.6997 | 113963.3874 | 4786.121 | 23.8112 | 5.1153 | 5.1394 | 114010.2165 | 4798.7274 | 0.99485 | 0.97913 | 8.8764 | 8.4782 | 0.2983 | 0.0264 | 8.7139 | 8.3075 | 9.2065 | 8.6016 | 0.49267 | 0.29405 | 0.080035 | 0.081625 | 130.1436 | 120.2213 | 0.00037633 | 0.0041047 | FALSE | 1.047 | TRUE | [M-H]- | 106.9258 | 8.7958 | 8.9679 | 0.17215 | 8.8051 | 8.9537 | 0.14869 | -87.0514 | 1.1037 | 1.1645 | 6.0615 | 0 | 26 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 582961.0123 | No Control Sample | 0 | off | off | 582961.0123 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9265 | -0.87379 | -0.47965 | 79.9561 | -12.0697 | -0.96506 | 0 | 0 | 0 | 0 | 0 | 2.3442 | FALSE | 25 | 23.6997
5684 | 5685 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFDS | 109.859156111719 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.19 | 8.45 | Baseline | Baseline | 615399.8139 | 24597.8343 | 25.0185 | 130703.3348 | 4786.121 | 27.3088 | 4.7084 | 5.1394 | 130757.095 | 4798.7274 | 0.99229 | 0.97913 | 9.2067 | 8.4782 | 0.18217 | 0.016741 | 9.037 | 8.3075 | 9.409 | 8.6016 | 0.37198 | 0.29405 | 0.073166 | 0.081625 | 166.1134 | 120.2213 | 0.00087859 | 0.0041047 | FALSE | 1.0859 | TRUE | [M-H]- | 109.8592 | 9.1324 | 9.2898 | 0.15732 | 9.14 | 9.2793 | 0.13932 | 308.7102 | 1.0494 | 1.0665 | 5.6945 | 0 | 21 | 5 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 615399.8139 | No Control Sample | 0 | off | off | 615399.8139 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9243 | 0.7197 | 0.43105 | 79.9562 | -8.5245 | -0.68159 | 0 | 0 | 0 | 0 | 0 | 1.1353 | FALSE | 25 | 25.0185
5688 | 5689 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FBSA | 109.794385715447 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 4.6 | 8.86 | Baseline | Baseline | 1538270.6166 | 53134.8678 | 28.9503 | 136482.4438 | 11306.4658 | 12.0712 | 11.2708 | 4.6995 | 137060.4561 | 11312.8718 | 0.98485 | 0.99211 | 4.9795 | 9.0224 | 8.2496 | 0.37948 | 4.6274 | 8.882 | 5.2873 | 9.1749 | 0.65988 | 0.29292 | 0.17384 | 0.071873 | 785.7719 | 140.7398 | 0.0067541 | 0.00072093 | FALSE | 0.5519 | TRUE | [M-H]- | 109.7944 | 4.7899 | 5.1839 | 0.39392 | 4.8171 | 5.1496 | 0.33252 | 1396.9478 | 0.976 | 0.90375 | 3.1977 | 0 | 119 | 34 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1538270.6166 | No Control Sample | 0 | off | off | 1538270.6166 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9595 | 1.7156 | 0.51118 | 77.9645 | -10.538 | -0.8216 | 0 | 0 | 0 | 0 | 0 | 1.4439 | FALSE | 25 | 28.9503
5690 | 5691 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FPeSA | 98.2841282518403 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7 | 8.86 | Baseline | Baseline | 2138367.451 | 53134.8678 | 40.2441 | 344295.0416 | 11306.4658 | 30.4512 | 6.2109 | 4.6995 | 344697.9182 | 11312.8718 | 0.99184 | 0.99211 | 7.1114 | 9.0224 | 1.591 | 0.11137 | 6.9372 | 8.882 | 7.297 | 9.1749 | 0.3598 | 0.29292 | 0.099986 | 0.071873 | 171.44 | 140.7398 | 0.0023185 | 0.00072093 | FALSE | 0.78819 | TRUE | [M-H]- | 98.2841 | 7.0207 | 7.1978 | 0.1771 | 7.0304 | 7.1869 | 0.15652 | 2218.5888 | 0.98552 | 0.95193 | 1.5071 | 0 | 40 | 13 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2138367.451 | No Control Sample | 0 | off | off | 2138367.451 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9561 | 0.85265 | 0.29668 | 77.964 | -16.7139 | -1.3031 | 0 | 0 | 0 | 0 | 0 | 0.50958 | FALSE | 25 | 40.2441
5692 | 5693 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FHxSA | 99.3588304445533 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7.86 | 8.86 | Baseline | Baseline | 1078759.9215 | 53134.8678 | 20.3023 | 217019.0733 | 11306.4658 | 19.1942 | 4.9708 | 4.6995 | 217129.0249 | 11312.8718 | 0.99472 | 0.99211 | 7.998 | 9.0224 | 1.756 | 0.13802 | 7.8551 | 8.882 | 8.3146 | 9.1749 | 0.45945 | 0.29292 | 0.078599 | 0.071873 | 118.8533 | 140.7398 | 0.00050601 | 0.00072093 | FALSE | 0.88646 | TRUE | [M-H]- | 99.3588 | 7.9249 | 8.0836 | 0.15868 | 7.9341 | 8.0736 | 0.1395 | 239.0108 | 1.0838 | 1.1797 | 3.1408 | 0 | 33 | 8 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1078759.9215 | No Control Sample | 0 | off | off | 1078759.9215 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9533 | 1.8319 | 0.72902 | 77.9642 | -15.2817 | -1.1914 | 0 | 0 | 0 | 0 | 0 | 2.2598 | FALSE | 25 | 20.3023
5694 | 5695 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FHpSA | 75.0285960897687 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 0 | 8.58 | 8.86 | Baseline | Baseline | 1687262.7568 | 53134.8678 | 31.7543 | 305180.8449 | 11306.4658 | 26.9917 | 5.5287 | 4.6995 | 306090.0718 | 11312.8718 | 0.98469 | 0.99211 | 8.5796 | 9.0224 | 0.0051902 | 0.00044532 | 8.448 | 8.882 | 8.8915 | 9.1749 | 0.44357 | 0.29292 | 0.089311 | 0.071873 | 101.3206 | 140.7398 | 0.0031016 | 0.00072093 | FALSE | 0.95092 | FALSE | [M-H]- | 75.0286 | 8.498 | 8.6659 | 0.16787 | 8.5094 | 8.6545 | 0.14517 | -2133.9175 | 1.0802 | 1.1869 | 1.9186 | 0 | 25 | 7 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1687262.7568 | No Control Sample | 0 | off | off | 1687262.7568 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9499 | 1.0352 | 0.46371 | 77.9641 | -17.2479 | -1.3447 | 0 | 0 | 0 | 0 | 0 | 1.0228 | FALSE | 25 | 31.7543
5696 | 5697 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FOSA | 97.5996375740649 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 8.88 | 8.86 | Baseline | Baseline | 1179976.2901 | 53134.8678 | 22.2072 | 227050.929 | 11306.4658 | 20.0815 | 5.197 | 4.6995 | 227151.2696 | 11312.8718 | 0.99191 | 0.99211 | 9.0237 | 9.0224 | 1.6183 | 0.1437 | 8.8312 | 8.882 | 9.2252 | 9.1749 | 0.39403 | 0.29292 | 0.081479 | 0.071873 | 114.1844 | 140.7398 | 0.0008463 | 0.00072093 | FALSE | 1.0001 | TRUE | [M-H]- | 97.5996 | 8.9383 | 9.1059 | 0.16753 | 8.9488 | 9.0959 | 0.14706 | 487.6574 | 1.0811 | 1.1956 | 3.183 | 0 | 21 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1179976.2901 | No Control Sample | 0 | off | off | 1179976.2901 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9467 | 0.96289 | 0.47947 | 77.9643 | -14.5329 | -1.1331 | 0 | 0 | 0 | 0 | 0 | 1.0763 | FALSE | 25 | 22.2072
5715 | 5716 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 4:2 FTS | 98.9000101113515 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 5.94 | 5.92 | Baseline | Baseline | 661199.1636 | 13166.9957 | 50.2164 | 110435.4071 | 2438.1482 | 45.2948 | 5.9872 | 5.4004 | 110521.6173 | 2441.5479 | 0.99482 | 0.99269 | 5.9491 | 5.9505 | 0.15311 | 0.0090945 | 5.734 | 5.7939 | 6.1272 | 6.1298 | 0.39327 | 0.33588 | 0.091226 | 0.083928 | 238.3306 | 174.0256 | 0.001299 | 0.0024289 | FALSE | 0.99977 | TRUE | [M-H]- | 98.9 | 5.8504 | 6.0482 | 0.1978 | 5.8651 | 6.0351 | 0.16998 | 364.7763 | 1.017 | 1.0602 | 2.7672 | 0 | 49 | 13 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 661199.1636 | No Control Sample | 0 | off | off | 661199.1636 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9752 | 2.7146 | 0.88762 | 306.9676 | 0.12983 | 0.039853 | 0 | 0 | 0 | 0 | 0 | 1.1154 | FALSE | 25 | 50.2164
5717 | 5718 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 6:2 FTS | 109.806365954991 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7.92 | 7.92 | Baseline | Baseline | 473867.2337 | 7112.7353 | 66.6224 | 115905.573 | 1787.7786 | 64.8322 | 4.0884 | 3.9785 | 115946.3512 | 1790.7964 | 0.99585 | 0.98848 | 7.9281 | 7.9268 | 0.10181 | 0.0080633 | 7.7774 | 7.7975 | 8.1289 | 8.0575 | 0.35158 | 0.26 | 0.06134 | 0.058756 | 104.0081 | 89.0494 | 0.00076043 | 0.0033652 | FALSE | 1.0002 | TRUE | [M-H]- | 109.8064 | 7.8566 | 7.9986 | 0.14195 | 7.8667 | 7.9871 | 0.12042 | 250.6886 | 0.93894 | 0.8378 | 2.0335 | 0 | 27 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 473867.2337 | No Control Sample | 0 | off | off | 473867.2337 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9696 | 4.0742 | 1.7395 | 406.9611 | -0.87881 | -0.35764 | 0 | 0 | 0 | 0 | 0 | 1.6777 | FALSE | 25 | 66.6224
5719 | 5720 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 8:2 FTS | 99.8432270363352 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 8.88 | 8.89 | Baseline | Baseline | 384582.0386 | 7850.4904 | 48.9883 | 71019.0922 | 1423.6286 | 49.886 | 5.4152 | 5.5144 | 71039.5159 | 1423.6286 | 0.997 | 1 | 8.8902 | 8.8926 | 0.11433 | 0.010152 | 8.7148 | 8.6686 | 9.1514 | 9.1605 | 0.4366 | 0.49182 | 0.084356 | 0.085654 | 115.7099 | 98.375 | 0.00068526 | 0 | FALSE | 0.99972 | TRUE | [M-H]- | 99.8432 | 8.8017 | 8.981 | 0.17929 | 8.815 | 8.9716 | 0.15657 | 111.4673 | 0.95151 | 0.93594 | 1.3123 | 0 | 23 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 384582.0386 | No Control Sample | 0 | off | off | 384582.0386 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9627 | 2.2639 | 1.193 | 506.9545 | -2.2702 | -1.1509 | 0 | 0 | 0 | 0 | 0 | 2.3891 | FALSE | 25 | 48.9883
5721 | 5722 | PFAS CS10 400ng/ml | 41 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 42) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 10:2 FTS | 103.575266254236 | Quantifiers | 10/27/2023 06:11:54 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.5 | 8.89 | Baseline | Baseline | 415322.9075 | 7850.4904 | 52.9041 | 92612.082 | 1423.6286 | 65.0535 | 4.4845 | 5.5144 | 92619.0877 | 1423.6286 | 0.9981 | 1 | 9.5164 | 8.8926 | 0.17254 | 0.016391 | 9.3432 | 8.6686 | 9.7329 | 9.1605 | 0.38972 | 0.49182 | 0.070921 | 0.085654 | 290.1777 | 98.375 | 0.00016239 | 0 | FALSE | 1.0701 | TRUE | [M-H]- | 103.5753 | 9.443 | 9.5875 | 0.14454 | 9.451 | 9.5784 | 0.1274 | 38.5911 | 0.95062 | 0.8722 | 2.4142 | 0 | 28 | 7 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 415322.9075 | No Control Sample | 0 | off | off | 415322.9075 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9563 | 1.8474 | 1.1583 | 606.9492 | 2.0856 | 1.2658 | 0 | 0 | 0 | 0 | 0 | 2.4082 | FALSE | 25 | 52.9041
5782 | 5783 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFBA | 100.242106833892 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 1.49 | 1.48 | Baseline | Baseline | 1568920.6088 | 87254.7848 | 17.9809 | 203635.7129 | 11014.7649 | 18.4875 | 7.7045 | 7.9216 | 203850.33 | 11025.8519 | 0.99288 | 0.99367 | 1.4541 | 1.4551 | 2.4101 | 0.03591 | 1.2821 | 1.3036 | 1.7238 | 1.686 | 0.44173 | 0.38243 | 0.12202 | 0.12249 | 268.7439 | 172.6694 | 0.0026514 | 0.0023858 | FALSE | 0.99932 | TRUE | [M-H]- | 100.2421 | 1.3611 | 1.6212 | 0.26008 | 1.3699 | 1.5944 | 0.22447 | 1222.2705 | 1.422 | 1.7143 | 1.0856 | 0 | 90 | 27 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1568920.6088 | No Control Sample | 0 | off | off | 1568920.6088 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9802 | 4.6655 | 0.99365 | 168.9881 | -5.3346 | -0.90148 | 0 | 0 | 0 | 0 | 0 | 0.31241 | FALSE | 25 | 17.9809
5784 | 5785 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFPeA | 105.254472275433 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 3.18 | 3.21 | Baseline | Baseline | 1496691.806 | 63583.4064 | 23.539 | 151077.4158 | 6527.8344 | 23.1436 | 9.9068 | 9.7404 | 151233.0684 | 6643.9946 | 0.99309 | 0.95047 | 3.1936 | 3.1908 | 0.42753 | 0.013595 | 2.915 | 3.011 | 3.5446 | 3.4611 | 0.6296 | 0.45007 | 0.15873 | 0.16313 | 1206.8882 | 447.0275 | 0.001768 | 0.021136 | FALSE | 1.0009 | TRUE | [M-H]- | 105.2545 | 3.0374 | 3.3866 | 0.34923 | 3.0584 | 3.3569 | 0.29851 | 424.2517 | 1.1821 | 1.3554 | 1.4146 | 0 | 112 | 31 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1496691.806 | No Control Sample | 0 | off | off | 1496691.806 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9771 | 4.0348 | 1.0611 | 218.9851 | -3.3963 | -0.74374 | 0 | 0 | 0 | 0 | 0 | 0.13865 | FALSE | 25 | 23.539
5786 | 5787 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFHxA | 76.4816661258644 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 1007.8768 | 6.13 | 6.12 | Baseline | Baseline | 2054492.186 | 115424.7956 | 17.7994 | 278166.5469 | 20276.4134 | 13.7187 | 7.3858 | 5.6926 | 278309.8207 | 20329.9272 | 0.99528 | 0.98782 | 6.1377 | 6.1316 | 0.12592 | 0.007719 | 5.9299 | 5.9572 | 6.4155 | 6.2692 | 0.48565 | 0.31202 | 0.12081 | 0.088519 | 211.8299 | 212.0257 | 0.0010253 | 0.0045892 | FALSE | 1.001 | FALSE | [M-H]- | 76.4817 | 6.025 | 6.2403 | 0.21522 | 6.0378 | 6.2274 | 0.18967 | 587.2671 | 1.0738 | 1.1679 | 1.1688 | 0 | 54 | 15 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2054492.186 | No Control Sample | 203843.5852 | off | off | 2054492.186 | 203843.5852 | off | 0.01 | off | 100 | off | 1007.8768 | off | 312.9735 | 2.0739 | 0.64906 | 268.9809 | -5.638 | -1.5165 | 0 | 0 | 0 | 0 | 0 | 1.3491 | FALSE | 25 | 17.7994
5788 | 5789 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFHpA | 66.1451298665001 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 780.2256 | 7.26 | 7.27 | Baseline | Baseline | 1781479.5534 | 107635.2259 | 16.5511 | 282386.545 | 23278.1023 | 12.131 | 6.3087 | 4.6239 | 282807.1744 | 23299.3961 | 0.99105 | 0.99353 | 7.2788 | 7.2775 | 0.25865 | 0.018778 | 7.1303 | 7.1512 | 7.4466 | 7.4083 | 0.3163 | 0.2571 | 0.10256 | 0.073103 | 153.0526 | 199.1539 | 0.003065 | 0.0012195 | FALSE | 1.0002 | FALSE | [M-H]- | 66.1451 | 7.1945 | 7.362 | 0.16743 | 7.2026 | 7.353 | 0.15038 | 2736.379 | 1.0502 | 1.0991 | 1.2537 | 0 | 38 | 14 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1781479.5534 | No Control Sample | 228328.7703 | off | off | 1781479.5534 | 228328.7703 | off | 0.01 | off | 100 | off | 780.2256 | off | 362.9705 | 2.485 | 0.90197 | 318.9778 | -4.0307 | -1.2857 | 0 | 0 | 0 | 0 | 0 | 1.48 | FALSE | 25 | 16.5511
5790 | 5791 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFOA | 75.4643468707537 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 777.3773 | 7.96 | 7.94 | Baseline | Baseline | 1488486.3853 | 81330.6664 | 18.3017 | 262798.1253 | 18993.0409 | 13.8365 | 5.664 | 4.2821 | 262957.979 | 19004.7808 | 0.99472 | 0.994 | 7.9634 | 7.9585 | 0.042455 | 0.0033794 | 7.7327 | 7.8363 | 8.2201 | 8.1179 | 0.48733 | 0.28158 | 0.0903 | 0.064729 | 86.3422 | 107.68 | 0.0011665 | 0.0013713 | FALSE | 1.0006 | FALSE | [M-H]- | 75.4643 | 7.887 | 8.0568 | 0.16978 | 7.8968 | 8.0446 | 0.14786 | 629.0457 | 1.0968 | 1.1872 | 1.0832 | 0 | 36 | 9 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1488486.3853 | No Control Sample | 191475.4056 | off | off | 1488486.3853 | 191475.4056 | off | 0.01 | off | 100 | off | 777.3773 | off | 412.9674 | 2.3376 | 0.96536 | 368.9746 | -3.9494 | -1.4572 | 0 | 0 | 0 | 0 | 0 | 0.18515 | FALSE | 25 | 18.3017
5792 | 5793 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFNA | 70.2945743057104 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 822.0795 | 8.46 | 8.47 | Baseline | Baseline | 1308285.1976 | 80115.64 | 16.33 | 215372.2303 | 15226.767 | 14.1443 | 6.0745 | 5.2615 | 215439.7274 | 15226.8134 | 0.99697 | 0.99845 | 8.4754 | 8.473 | 0.18248 | 0.015438 | 8.2999 | 8.261 | 8.7469 | 8.7176 | 0.44697 | 0.45652 | 0.098439 | 0.08381 | 127.2978 | 172.9906 | 0.0007445 | 6.0047e-06 | FALSE | 1.0003 | FALSE | [M-H]- | 70.2946 | 8.3843 | 8.5642 | 0.17987 | 8.3919 | 8.5544 | 0.16257 | 358.7418 | 1.0698 | 1.125 | 0.91164 | 0 | 26 | 7 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1308285.1976 | No Control Sample | 159143.3885 | off | off | 1308285.1976 | 159143.3885 | off | 0.01 | off | 100 | off | 822.0795 | off | 462.9642 | 2.0288 | 0.93927 | 418.9719 | -3.0496 | -1.2777 | 0 | 0 | 0 | 0 | 0 | 0.55165 | FALSE | 25 | 16.33
5794 | 5795 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFDA | 87.362948906743 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 8.87 | 8.86 | Baseline | Baseline | 1257696.4861 | 57992.0982 | 21.6874 | 214698.5213 | 10809.609 | 19.8618 | 5.858 | 5.3649 | 214728.7155 | 10813.2908 | 0.99833 | 0.99713 | 8.8858 | 8.8887 | 0.17836 | 0.015821 | 8.6929 | 8.7007 | 9.2046 | 9.0604 | 0.5117 | 0.35973 | 0.093437 | 0.084131 | 95.1285 | 129.0559 | 0.00034842 | 0.00066466 | FALSE | 0.99967 | FALSE | [M-H]- | 87.3629 | 8.7937 | 8.9806 | 0.18685 | 8.805 | 8.9697 | 0.16466 | 146.1905 | 0.95784 | 0.90864 | 0.58783 | 0 | 27 | 7 | C10HF19O2 | 512.96 | 468.97 | Gray | Pass | Pass | Fail | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1257696.4861 | No Control Sample | 0 | off | off | 1257696.4861 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9621 | 3.9843 | 2.0438 | 468.9689 | -2.3246 | -1.0902 | 0 | 0 | 0 | 0 | 0 | 1.5314 | FALSE | 25 | 21.6874
5796 | 5797 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFUdA | 92.4370837156688 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.21 | 9.2 | Baseline | Baseline | 1056963.7046 | 51560.5605 | 20.4995 | 198266.336 | 10330.6152 | 19.1921 | 5.331 | 4.991 | 198292.692 | 10335.2111 | 0.99803 | 0.99631 | 9.2226 | 9.223 | 0.1364 | 0.012563 | 9.0346 | 9.0242 | 9.4875 | 9.3633 | 0.4529 | 0.33903 | 0.085398 | 0.077971 | 225.1239 | 169.5529 | 0.00030296 | 0.00078084 | FALSE | 0.99995 | TRUE | [M-H]- | 92.4371 | 9.1382 | 9.3097 | 0.17155 | 9.15 | 9.2992 | 0.14924 | 132.6262 | 0.9881 | 0.99015 | 0.54395 | 0 | 26 | 7 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1056963.7046 | No Control Sample | 0 | off | off | 1056963.7046 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9587 | 3.2178 | 1.8115 | 518.966 | -2.5336 | -1.3149 | 0 | 0 | 0 | 0 | 0 | 3.0444 | FALSE | 25 | 20.4995
5798 | 5799 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFDoA | 79.4067283049614 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 0 | 9.47 | 9.47 | Baseline | Baseline | 1045351.1448 | 50871.8409 | 20.5487 | 207517.6099 | 11183.5393 | 18.5556 | 5.0374 | 4.5488 | 207585.8404 | 11191.6497 | 0.99646 | 0.98761 | 9.4866 | 9.4851 | 0.1749 | 0.016563 | 9.3409 | 9.3474 | 9.7089 | 9.6081 | 0.36798 | 0.26068 | 0.081145 | 0.072593 | 244.8902 | 219.542 | 0.00081153 | 0.0014138 | FALSE | 1.0002 | FALSE | [M-H]- | 79.4067 | 9.411 | 9.5641 | 0.15306 | 9.4185 | 9.5552 | 0.13671 | 457.648 | 0.98974 | 0.95856 | 0.8136 | 0 | 26 | 7 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1045351.1448 | No Control Sample | 0 | off | off | 1045351.1448 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9543 | 1.0297 | 0.63118 | 568.9625 | -1.5681 | -0.8922 | 0 | 0 | 0 | 0 | 0 | 3.3278 | FALSE | 25 | 20.5487
5800 | 5801 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFTrDA | 95.9296921914101 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.68 | 9.47 | Baseline | Baseline | 920541.1527 | 50871.8409 | 18.0953 | 193812.4202 | 11183.5393 | 17.3302 | 4.7496 | 4.5488 | 193883.8113 | 11191.6497 | 0.99551 | 0.98761 | 9.7039 | 9.4851 | 0.24695 | 0.023905 | 9.5333 | 9.3474 | 9.8921 | 9.6081 | 0.35878 | 0.26068 | 0.07648 | 0.072593 | 518.8228 | 219.542 | 0.00072236 | 0.0014138 | FALSE | 1.0231 | TRUE | [M-H]- | 95.9297 | 9.6343 | 9.7725 | 0.13817 | 9.6404 | 9.7642 | 0.12386 | 390.2163 | 0.91512 | 0.78484 | 1.0284 | 0 | 32 | 8 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 920541.1527 | No Control Sample | 0 | off | off | 920541.1527 | 0 | off | 0 | off | 0 | off | 0 | off | 662.951 | 0.79609 | 0.52777 | 618.9602 | -2.3946 | -1.4822 | 0 | 0 | 0 | 0 | 0 | 2.1884 | FALSE | 25 | 18.0953
5802 | 5803 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFTeDA | 100.871949891789 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.87 | 9.86 | Baseline | Baseline | 718423.4168 | 20087.9886 | 35.7638 | 174309.9766 | 5741.9536 | 30.3573 | 4.1215 | 3.4985 | 174508.2495 | 5749.5088 | 0.99219 | 0.98883 | 9.8811 | 9.8775 | 0.11244 | 0.011098 | 9.7545 | 9.7864 | 10.0073 | 9.9935 | 0.25282 | 0.20712 | 0.066035 | 0.054111 | 737.6233 | 536.7805 | 0.002194 | 0.0029495 | FALSE | 1.0004 | TRUE | [M-H]- | 100.8719 | 9.8246 | 9.9458 | 0.1212 | 9.8296 | 9.9377 | 0.10808 | 1512.7279 | 1.0317 | 1.0126 | 0.83658 | 0 | 28 | 9 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 718423.4168 | No Control Sample | 0 | off | off | 718423.4168 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9483 | 1.409 | 1.0045 | 668.9571 | -1.1309 | -0.75655 | 0 | 0 | 0 | 0 | 0 | 3.3577 | FALSE | 25 | 35.7638
5809 | 5810 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFBS | 100.513136105157 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 3.5 | 3.5 | Baseline | Manual | 985570.7439 | 37045.1226 | 26.6046 | 86945.5496 | 3131.6965 | 27.7631 | 11.3355 | 11.8291 | 87157.784 | 3144.7607 | 0.98929 | 0.98488 | 3.8349 | 3.8332 | 9.5681 | 0.33488 | 3.493 | 3.5862 | 4.1937 | 4.1387 | 0.70065 | 0.55253 | 0.17577 | 0.18396 | 956.9111 | 703.7789 | 0.0051214 | 0.00093822 | FALSE | 1.0004 | TRUE | [M-H]- | 100.5131 | 3.6574 | 4.0429 | 0.38551 | 3.6781 | 4.0196 | 0.34151 | 635.5296 | 1.1521 | 1.3293 | 5.6063 | 0 | 109 | 29 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 985570.7439 | No Control Sample | 0 | off | off | 985570.7439 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9436 | 2.1641 | 0.64695 | 79.9565 | -9.0475 | -0.72342 | 0 | 0 | 0 | 0 | 0 | 1.0558 | FALSE | 25 | 26.6046
5811 | 5812 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFPeS | 97.9549725901728 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 6.38 | 3.5 | Baseline | Manual | 967100.6246 | 37045.1226 | 26.106 | 179500.0147 | 3131.6965 | 57.3172 | 5.3877 | 11.8291 | 179644.0207 | 3144.7607 | 0.99221 | 0.98488 | 6.3874 | 3.8332 | 0.11581 | 0.007389 | 6.2279 | 3.5862 | 6.6082 | 4.1387 | 0.38032 | 0.55253 | 0.086301 | 0.18396 | 327.5151 | 703.7789 | 0.0018238 | 0.00093822 | FALSE | 1.6663 | TRUE | [M-H]- | 97.955 | 6.3022 | 6.4771 | 0.17496 | 6.3125 | 6.4652 | 0.15266 | 860.8066 | 1.0062 | 0.99281 | 3.6398 | 0 | 42 | 12 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 967100.6246 | No Control Sample | 0 | off | off | 967100.6246 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9401 | 0.84384 | 0.29445 | 79.956 | -14.4854 | -1.1582 | 0 | 0 | 0 | 0 | 0 | 1.0433 | FALSE | 25 | 26.106
5813 | 5814 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | L-PFHxS | 86.4948110707599 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 81.1 | 4 | 20.275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 683543.7816 | 42307.0447 | 16.1567 | 162623.8263 | 10293.1826 | 15.7992 | 4.2032 | 4.1102 | 164857.1448 | 10296.8307 | 0.72226 | 0.99282 | 7.3536 | 7.3518 | 0.18546 | 0.013613 | 7.2678 | 7.2254 | 7.4812 | 7.5189 | 0.21347 | 0.29348 | 0.066596 | 0.063769 | 102.2075 | 163.6527 | 0.015554 | 0.00080338 | FALSE | 1.0002 | TRUE | [M-H]- | 106.652 | 7.2897 | 7.4222 | 0.13248 | 7.2969 | 7.4135 | 0.11665 | -11849.1747 | 1.0987 | 1.1973 | 5.3645 | 0 | 27 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 683543.7816 | No Control Sample | 0 | off | off | 683543.7816 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9371 | 1.2213 | 0.48721 | 79.9561 | -13.624 | -1.0893 | 0 | 0 | 0 | 0 | 0 | 1.4972 | FALSE | 20.275 | 16.1567
5815 | 5816 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | Br-PFHxS | 20.8162445227274 | Qualifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 18.9 | 4 | 4.725 | 0 | 7.16 | 7.35 | Manual | Baseline | 223502.5395 | 42307.0447 | 5.2829 | 42017.9212 | 10293.1826 | 4.0821 | 5.3192 | 4.1102 | 164857.1448 | 10296.8307 | 0.23624 | 0.99282 | 7.1708 | 7.3518 | 0.15066 | 0.010787 | 7.0548 | 7.2254 | 7.2667 | 7.5189 | 0.21188 | 0.29348 | 0.084107 | 0.063769 | 26.4078 | 163.6527 | 0.017834 | 0.00080338 | TRUE | 0.97538 | TRUE | [M-H]- | 110.1389 | 7.0869 | 7.2678 | 0.18084 | 7.0946 | 7.2493 | 0.15469 | 3536.6357 | 1.1071 | 1.09 | 0.32698 | 0 | 25 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 223502.5395 | No Control Sample | 0 | off | off | 683543.7816 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9371 | 1.2213 | 0.48721 | 79.9607 | 43.6063 | 3.4866 | 0 | 0 | 0 | 0 | 0 | 1.4972 | FALSE | 4.725 | 5.2829
5817 | 5818 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFHpS | 103.625143901412 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7.98 | 7.35 | Baseline | Baseline | 667229.1628 | 42307.0447 | 15.7711 | 146935.4228 | 10293.1826 | 14.275 | 4.541 | 4.1102 | 147047.426 | 10296.8307 | 0.99258 | 0.99282 | 7.9891 | 7.3518 | 0.11385 | 0.0090854 | 7.8592 | 7.2254 | 8.2554 | 7.5189 | 0.3962 | 0.29348 | 0.071275 | 0.063769 | 126.4194 | 163.6527 | 7.4547e-05 | 0.00080338 | FALSE | 1.0867 | TRUE | [M-H]- | 103.6251 | 7.9213 | 8.0695 | 0.14825 | 7.9285 | 8.0605 | 0.13204 | 27.6467 | 1.0821 | 1.1542 | 6.254 | 0 | 29 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 667229.1628 | No Control Sample | 0 | off | off | 667229.1628 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9335 | 0.25435 | 0.11418 | 79.9561 | -13.4341 | -1.0742 | 0 | 0 | 0 | 0 | 0 | 1.5296 | FALSE | 25 | 15.7711
5819 | 5820 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | L-PFOS | 83.3061018966871 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 78.8 | 4 | 19.7 | 0 | 8.46 | 8.45 | Manual | Baseline | 517962.8144 | 23920.2901 | 21.6537 | 101704.2342 | 4551.5998 | 22.3447 | 5.0928 | 5.2554 | 102809.4162 | 4561.8638 | 0.59777 | 0.97891 | 8.47 | 8.4695 | 0.11848 | 0.010024 | 8.3905 | 8.3267 | 8.5821 | 8.6343 | 0.19164 | 0.30757 | 0.080721 | 0.084289 | 42.2252 | 128.4518 | 0.01016 | 0.0051008 | TRUE | 1.0001 | TRUE | [M-H]- | 105.7184 | 8.3728 | 8.5496 | 0.17682 | 8.3728 | 8.5412 | 0.16837 | -5391.6829 | 0.88484 | 0.68506 | 7.5852 | 0 | 13 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 517962.8144 | No Control Sample | 0 | off | off | 517962.8144 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9303 | 0.2342 | 0.11685 | 79.9562 | -13.4692 | -1.077 | 0 | 0 | 0 | 0 | 0 | 1.7977 | FALSE | 19.7 | 21.6537
5821 | 5822 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | Br-PFOS | 23.6260342840847 | Qualifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 21.1 | 4 | 5.275 | 0 | 8.3 | 8.45 | Manual | Baseline | 314481.1617 | 23920.2901 | 13.147 | 54307.2336 | 4551.5998 | 11.9315 | 5.7908 | 5.2554 | 102809.4162 | 4561.8638 | 0.36303 | 0.97891 | 8.3219 | 8.4695 | 0.26373 | 0.021889 | 8.2115 | 8.3267 | 8.3955 | 8.6343 | 0.18401 | 0.30757 | 0.094714 | 0.084289 | 22.5471 | 128.4518 | 0.0070273 | 0.0051008 | TRUE | 0.98257 | TRUE | [M-H]- | 111.9717 | 8.2322 | 8.4098 | 0.17756 | 8.2426 | 8.4098 | 0.16714 | 2073.9931 | 1.0227 | 1.188 | 0.60715 | 0 | 13 | 7 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 314481.1617 | No Control Sample | 0 | off | off | 517962.8144 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9303 | 0.2342 | 0.11685 | 79.9569 | -4.5279 | -0.36204 | 0 | 0 | 0 | 0 | 0 | 1.7977 | FALSE | 5.275 | 13.147
5823 | 5824 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFNS | 110.030436393989 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 8.85 | 8.45 | Baseline | Baseline | 583363.7081 | 23920.2901 | 24.3878 | 110163.1408 | 4551.5998 | 24.2032 | 5.2955 | 5.2554 | 110212.6719 | 4561.8638 | 0.99471 | 0.97891 | 8.8659 | 8.4695 | 0.17918 | 0.015857 | 8.6607 | 8.3267 | 9.1519 | 8.6343 | 0.4913 | 0.30757 | 0.082526 | 0.084289 | 102.3973 | 128.4518 | 8.7157e-05 | 0.0051008 | FALSE | 1.0468 | TRUE | [M-H]- | 110.0304 | 8.7852 | 8.9562 | 0.171 | 8.7958 | 8.9456 | 0.1498 | -19.5431 | 1.2282 | 1.536 | 6.9742 | 0 | 26 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 583363.7081 | No Control Sample | 0 | off | off | 583363.7081 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9272 | 0.39539 | 0.21704 | 79.9563 | -10.0337 | -0.80227 | 0 | 0 | 0 | 0 | 0 | 1.7115 | FALSE | 25 | 24.3878
5825 | 5826 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFDS | 116.326566436162 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 0 | 9.19 | 8.45 | Baseline | Baseline | 633679.3894 | 23920.2901 | 26.4913 | 134250.0424 | 4551.5998 | 29.4951 | 4.7201 | 5.2554 | 134323.7807 | 4561.8638 | 0.99295 | 0.97891 | 9.1942 | 8.4695 | 0.045567 | 0.0041876 | 9.0575 | 8.3267 | 9.3779 | 8.6343 | 0.32038 | 0.30757 | 0.073832 | 0.084289 | 232.6369 | 128.4518 | 0.00078753 | 0.0051008 | FALSE | 1.0856 | TRUE | [M-H]- | 116.3266 | 9.1196 | 9.2802 | 0.16064 | 9.1304 | 9.267 | 0.1366 | 329.9987 | 1.1363 | 1.281 | 5.7077 | 0 | 18 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 633679.3894 | No Control Sample | 0 | off | off | 633679.3894 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9241 | 0.45259 | 0.27107 | 79.9562 | -9.2483 | -0.73947 | 0 | 0 | 0 | 0 | 0 | 1.429 | FALSE | 25 | 26.4913
5829 | 5830 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FBSA | 105.672368108754 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 4.6 | 8.86 | Baseline | Baseline | 1544827.9053 | 55442.8643 | 27.8634 | 137422.8151 | 11853.5376 | 11.5934 | 11.2414 | 4.6773 | 137854.9745 | 11864.0196 | 0.98772 | 0.98964 | 4.9798 | 9.0114 | 8.257 | 0.37982 | 4.6505 | 8.8835 | 5.3025 | 9.1584 | 0.65195 | 0.27497 | 0.17418 | 0.073273 | 604.4015 | 153.6495 | 0.0038834 | 0.0016894 | FALSE | 0.55261 | TRUE | [M-H]- | 105.6724 | 4.7831 | 5.1782 | 0.39512 | 4.8117 | 5.1463 | 0.33462 | 818.5791 | 1.055 | 1.1089 | 3.0387 | 0 | 118 | 34 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1544827.9053 | No Control Sample | 0 | off | off | 1544827.9053 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9595 | 1.7568 | 0.52347 | 77.9644 | -11.1183 | -0.86684 | 0 | 0 | 0 | 0 | 0 | 1.2317 | FALSE | 25 | 27.8634
5831 | 5832 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FPeSA | 97.4038453296241 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7 | 8.86 | Baseline | Baseline | 2211266.58 | 55442.8643 | 39.8837 | 356629.9576 | 11853.5376 | 30.0864 | 6.2005 | 4.6773 | 356769.2432 | 11864.0196 | 0.99572 | 0.98964 | 7.1109 | 9.0114 | 1.5839 | 0.11087 | 6.9188 | 8.8835 | 7.3983 | 9.1584 | 0.47955 | 0.27497 | 0.10023 | 0.073273 | 181.2636 | 153.6495 | 0.00093066 | 0.0016894 | FALSE | 0.7891 | TRUE | [M-H]- | 97.4038 | 7.0232 | 7.1998 | 0.17653 | 7.0334 | 7.1886 | 0.15515 | 692.1128 | 1.0949 | 1.2038 | 1.6892 | 0 | 54 | 13 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2211266.58 | No Control Sample | 0 | off | off | 2211266.58 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9559 | 0.32635 | 0.11355 | 77.9638 | -18.7738 | -1.4637 | 0 | 0 | 0 | 0 | 0 | 1.0514 | FALSE | 25 | 39.8837
5833 | 5834 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FHxSA | 93.4710178011912 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7.86 | 8.86 | Baseline | Baseline | 1058915.6297 | 55442.8643 | 19.0992 | 214743.8228 | 11853.5376 | 18.1164 | 4.9311 | 4.6773 | 214901.2163 | 11864.0196 | 0.9934 | 0.98964 | 7.9955 | 9.0114 | 1.7241 | 0.13551 | 7.818 | 8.8835 | 8.2218 | 9.1584 | 0.40382 | 0.27497 | 0.077644 | 0.073273 | 143.1903 | 153.6495 | 0.0010698 | 0.0016894 | FALSE | 0.88727 | TRUE | [M-H]- | 93.471 | 7.9225 | 8.0797 | 0.15724 | 7.9323 | 8.0702 | 0.13796 | 568.9009 | 1.0079 | 1.0229 | 3.4632 | 0 | 30 | 8 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1058915.6297 | No Control Sample | 0 | off | off | 1058915.6297 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9533 | 1.9047 | 0.75799 | 77.9642 | -15.4812 | -1.207 | 0 | 0 | 0 | 0 | 0 | 2.3275 | FALSE | 25 | 19.0992
5835 | 5836 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FHpSA | 74.399459492131 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 0 | 8.58 | 8.86 | Baseline | Baseline | 1745788.9298 | 55442.8643 | 31.4881 | 328298.3304 | 11853.5376 | 27.6962 | 5.3177 | 4.6773 | 329364.295 | 11864.0196 | 0.98301 | 0.98964 | 8.5731 | 9.0114 | 0.080558 | 0.0069119 | 8.4501 | 8.8835 | 8.7309 | 9.1584 | 0.2808 | 0.27497 | 0.085153 | 0.073273 | 113.189 | 153.6495 | 0.0016023 | 0.0016894 | FALSE | 0.95136 | FALSE | [M-H]- | 74.3995 | 8.4932 | 8.656 | 0.16281 | 8.5008 | 8.6466 | 0.14576 | -1873.3729 | 0.96255 | 0.89505 | 1.728 | 0 | 17 | 7 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1745788.9298 | No Control Sample | 0 | off | off | 1745788.9298 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9498 | 0.87975 | 0.39409 | 77.9639 | -18.7533 | -1.4621 | 0 | 0 | 0 | 0 | 0 | 1.4023 | FALSE | 25 | 31.4881
5837 | 5838 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FOSA | 92.6840091779017 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 8.88 | 8.86 | Baseline | Baseline | 1169219.2099 | 55442.8643 | 21.0887 | 226962.8968 | 11853.5376 | 19.1473 | 5.1516 | 4.6773 | 227069.4301 | 11864.0196 | 0.99241 | 0.98964 | 9.0123 | 9.0114 | 1.4897 | 0.13229 | 8.8327 | 8.8835 | 9.2081 | 9.1584 | 0.37538 | 0.27497 | 0.081484 | 0.073273 | 127.557 | 153.6495 | 0.00077611 | 0.0016894 | FALSE | 1.0001 | TRUE | [M-H]- | 92.684 | 8.929 | 9.0947 | 0.16569 | 8.9386 | 9.0857 | 0.14712 | 469.2467 | 0.93309 | 0.85829 | 3.4632 | 0 | 20 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1169219.2099 | No Control Sample | 0 | off | off | 1169219.2099 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9464 | 0.46121 | 0.22966 | 77.9642 | -15.5804 | -1.2147 | 0 | 0 | 0 | 0 | 0 | 1.5514 | FALSE | 25 | 21.0887
5856 | 5857 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 4:2 FTS | 100.102346354722 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 5.94 | 5.92 | Baseline | Baseline | 662113.9548 | 13026.8442 | 50.8269 | 110292.9189 | 2226.2357 | 49.5423 | 6.0032 | 5.8515 | 110375.8921 | 2227.4537 | 0.99468 | 0.99443 | 5.9447 | 5.9377 | 0.079787 | 0.0047394 | 5.7192 | 5.7747 | 6.1187 | 6.0939 | 0.39953 | 0.31915 | 0.093219 | 0.092207 | 231.5091 | 182.5535 | 0.0012932 | 4.23e-05 | FALSE | 1.0012 | TRUE | [M-H]- | 100.1023 | 5.8442 | 6.0432 | 0.19903 | 5.8594 | 6.031 | 0.17158 | 356.9975 | 0.95803 | 0.93474 | 2.6798 | 0 | 50 | 14 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 662113.9548 | No Control Sample | 0 | off | off | 662113.9548 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9751 | 2.5737 | 0.84153 | 306.9673 | -0.93882 | -0.28819 | 0 | 0 | 0 | 0 | 0 | 1.5314 | FALSE | 25 | 50.8269
5858 | 5859 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 6:2 FTS | 96.9145477727498 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7.92 | 7.92 | Baseline | Baseline | 472429.9673 | 8034.4463 | 58.8006 | 119254.9018 | 2054.4599 | 58.0468 | 3.9615 | 3.9107 | 119276.038 | 2054.4599 | 0.99728 | 0.99765 | 7.9232 | 7.9252 | 0.040827 | 0.0032335 | 7.7507 | 7.7273 | 8.1935 | 8.1218 | 0.44278 | 0.39447 | 0.059848 | 0.058587 | 113.7309 | 114.9562 | 0.00042968 | 2.8946e-08 | FALSE | 0.99975 | TRUE | [M-H]- | 96.9145 | 7.8533 | 7.993 | 0.13969 | 7.864 | 7.9829 | 0.11894 | 115.7261 | 0.99851 | 1.0069 | 1.6988 | 0 | 33 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 472429.9673 | No Control Sample | 0 | off | off | 472429.9673 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9697 | 4.2568 | 1.8175 | 406.9612 | -0.85193 | -0.3467 | 0 | 0 | 0 | 0 | 0 | 2.7864 | FALSE | 25 | 58.8006
5860 | 5861 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 8:2 FTS | 99.1260195269677 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 8.88 | 8.89 | Baseline | Baseline | 384424.5204 | 7904.0525 | 48.6364 | 71880.8778 | 1423.9254 | 50.4808 | 5.3481 | 5.5509 | 71927.337 | 1423.9254 | 0.99463 | 0.99754 | 8.881 | 8.881 | 0.011728 | 0.0010414 | 8.7158 | 8.636 | 9.1143 | 9.1234 | 0.39848 | 0.48735 | 0.083567 | 0.088095 | 109.8162 | 123.7755 | 0.0015684 | 0 | FALSE | 1 | TRUE | [M-H]- | 99.126 | 8.7951 | 8.9742 | 0.17906 | 8.8035 | 8.9592 | 0.15573 | 282.9198 | 1.0931 | 1.1187 | 1.591 | 0 | 21 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 384424.5204 | No Control Sample | 0 | off | off | 384424.5204 | 0 | off | 0 | off | 0 | off | 0 | off | 526.964 | 4.7452 | 2.5006 | 506.9558 | 0.10834 | 0.054923 | 0 | 0 | 0 | 0 | 0 | 3.4788 | FALSE | 25 | 48.6364
5862 | 5863 | PFAS CS10 400ng/ml | 42 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 43) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 10:2 FTS | 102.482335015918 | Quantifiers | 10/27/2023 06:28:34 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.5 | 8.89 | Baseline | Baseline | 413744.1494 | 7904.0525 | 52.3458 | 91903.238 | 1423.9254 | 64.5422 | 4.502 | 5.5509 | 91921.8463 | 1423.9254 | 0.99604 | 0.99754 | 9.5048 | 8.881 | 0.050173 | 0.0047664 | 9.3624 | 8.636 | 9.7004 | 9.1234 | 0.33802 | 0.48735 | 0.071282 | 0.088095 | 306.9922 | 123.7755 | 0.00045133 | 0 | FALSE | 1.0702 | TRUE | [M-H]- | 102.4823 | 9.4298 | 9.5758 | 0.14599 | 9.4383 | 9.5671 | 0.1288 | 122.7107 | 0.94301 | 0.87048 | 1.44 | 0 | 24 | 7 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 413744.1494 | No Control Sample | 0 | off | off | 413744.1494 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9568 | 2.6566 | 1.6656 | 606.9479 | 0.038267 | 0.023226 | 0 | 0 | 0 | 0 | 0 | 3.3609 | FALSE | 25 | 52.3458
5923 | 5924 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFBA | 98.4384119844798 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 1.49 | 1.48 | Manual | Manual | 1508535.9692 | 85433.764 | 17.6574 | 200407.1828 | 11097.8405 | 18.0582 | 7.5274 | 7.6982 | 200619.3238 | 11132.3582 | 0.99356 | 0.98777 | 1.4518 | 1.4556 | 2.5633 | 0.038193 | 1.3316 | 1.3326 | 1.7046 | 1.6872 | 0.37297 | 0.35461 | 0.1145 | 0.11332 | 238.0185 | 164.2942 | 0 | 0.00093723 | TRUE | 0.99737 | TRUE | [M-H]- | 98.4384 | 1.3621 | 1.6219 | 0.25975 | 1.3709 | 1.5979 | 0.22704 | 0 | 1.5795 | 2.0907 | 1.0762 | 0 | 77 | 26 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1508535.9692 | No Control Sample | 0 | off | off | 1508535.9692 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9803 | 5.1751 | 1.1022 | 168.9879 | -6.5752 | -1.1111 | 0 | 0 | 0 | 0 | 0 | 0.44555 | FALSE | 25 | 17.6574
5925 | 5926 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFPeA | 98.3895404553714 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 3.18 | 3.21 | Baseline | Manual | 1466949.2918 | 66668.1092 | 22.0038 | 143328.6241 | 6426.4328 | 22.303 | 10.2349 | 10.374 | 143688.6398 | 6448.4401 | 0.98789 | 0.98536 | 3.1754 | 3.1732 | 0.14463 | 0.0045992 | 2.9009 | 2.963 | 3.4717 | 3.464 | 0.57082 | 0.50101 | 0.1572 | 0.1589 | 934.2371 | 501.1584 | 0.0047803 | 0.00096286 | FALSE | 1.0007 | TRUE | [M-H]- | 98.3895 | 3.0202 | 3.367 | 0.34683 | 3.0452 | 3.3352 | 0.28998 | 1200.2984 | 1.0935 | 1.1703 | 1.373 | 0 | 104 | 30 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1466949.2918 | No Control Sample | 0 | off | off | 1466949.2918 | 0 | off | 0 | off | 0 | off | 0 | off | 262.977 | 3.8988 | 1.0253 | 218.9848 | -4.4882 | -0.98286 | 0 | 0 | 0 | 0 | 0 | 0.76059 | FALSE | 25 | 22.0038
5927 | 5928 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFHxA | 79.4388840465803 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 1022.3423 | 6.13 | 6.12 | Baseline | Baseline | 2083979.2956 | 112722.9189 | 18.4876 | 273043.5934 | 18587.1274 | 14.6899 | 7.6324 | 6.0646 | 273222.0486 | 18615.5766 | 0.99466 | 0.99081 | 6.1182 | 6.1131 | 0.19189 | 0.011763 | 5.9185 | 5.9294 | 6.3742 | 6.2741 | 0.45565 | 0.34473 | 0.12424 | 0.09543 | 184.7125 | 225.7633 | 0.0015267 | 0.0028637 | FALSE | 1.0008 | FALSE | [M-H]- | 79.4389 | 6.0068 | 6.2245 | 0.21774 | 6.0191 | 6.2115 | 0.19241 | 914.8698 | 1.0389 | 1.0799 | 0.93181 | 0 | 51 | 16 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2083979.2956 | No Control Sample | 203843.5852 | off | off | 2083979.2956 | 203843.5852 | off | 0.01 | off | 100 | off | 1022.3423 | off | 312.9736 | 2.4143 | 0.7556 | 268.9812 | -4.5033 | -1.2113 | 0 | 0 | 0 | 0 | 0 | 1.3869 | FALSE | 25 | 18.4876
5929 | 5930 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFHpA | 66.1160614348863 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 788.7456 | 7.26 | 7.27 | Baseline | Baseline | 1800933.0663 | 108858.4273 | 16.5438 | 285690.0577 | 24896.9993 | 11.4749 | 6.3038 | 4.3724 | 285804.7074 | 24929.6757 | 0.99618 | 0.99136 | 7.2736 | 7.2734 | 0.18704 | 0.013579 | 7.1224 | 7.1341 | 7.5969 | 7.3925 | 0.47447 | 0.25835 | 0.10345 | 0.067671 | 126.8549 | 227.7469 | 0.00079024 | 0.002393 | FALSE | 1 | FALSE | [M-H]- | 66.1161 | 7.1892 | 7.3577 | 0.16853 | 7.1976 | 7.3485 | 0.15088 | 475.8289 | 0.98041 | 0.94698 | 1.1727 | 0 | 52 | 15 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1800933.0663 | No Control Sample | 228328.7703 | off | off | 1800933.0663 | 228328.7703 | off | 0.01 | off | 100 | off | 788.7456 | off | 362.9703 | 1.9002 | 0.68972 | 318.9779 | -3.8666 | -1.2334 | 0 | 0 | 0 | 0 | 0 | 0.9537 | FALSE | 25 | 16.5438
5931 | 5932 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFOA | 81.9994887938312 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy, Total Width | 0 | 0 | 100 | 4 | 25 | 784.7078 | 7.96 | 7.94 | Baseline | Baseline | 1502522.4799 | 75554.6356 | 19.8866 | 266755.44 | 16962.4113 | 15.7263 | 5.6326 | 4.4542 | 267080.0186 | 16978.6307 | 0.99269 | 0.99231 | 7.9648 | 7.9619 | 0.06025 | 0.0047959 | 7.8296 | 7.8361 | 8.3316 | 8.1636 | 0.50198 | 0.32755 | 0.09057 | 0.068527 | 82.7751 | 103.8016 | 0.00095325 | 0.0015584 | FALSE | 1.0004 | FALSE | [M-H]- | 81.9995 | 7.889 | 8.0582 | 0.16917 | 7.8982 | 8.0457 | 0.1475 | -506.5575 | 1.1297 | 1.2471 | 0.99438 | 0 | 36 | 9 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Fail | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1502522.4799 | No Control Sample | 191475.4056 | off | off | 1502522.4799 | 191475.4056 | off | 0.01 | off | 100 | off | 784.7078 | off | 412.9674 | 2.4128 | 0.99641 | 368.9752 | -2.5381 | -0.93651 | 0 | 0 | 0 | 0 | 0 | 1.3303 | FALSE | 25 | 19.8866
5933 | 5934 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFNA | 82.7837014506665 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 853.755 | 8.46 | 8.47 | Baseline | Baseline | 1358694.6901 | 70650.253 | 19.2313 | 219095.0256 | 13292.3963 | 16.4827 | 6.2014 | 5.3151 | 219156.9521 | 13319.0129 | 0.99715 | 0.99007 | 8.4805 | 8.479 | 0.24247 | 0.020513 | 8.2995 | 8.2921 | 8.7277 | 8.615 | 0.42823 | 0.32293 | 0.099999 | 0.085416 | 144.0575 | 133.6057 | 0.0006376 | 0.0033115 | FALSE | 1.0002 | FALSE | [M-H]- | 82.7837 | 8.3875 | 8.5693 | 0.18171 | 8.3947 | 8.5576 | 0.16292 | 326.2129 | 0.90367 | 0.74475 | 0.73558 | 0 | 25 | 7 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1358694.6901 | No Control Sample | 159143.3885 | off | off | 1358694.6901 | 159143.3885 | off | 0.01 | off | 100 | off | 853.755 | off | 462.9649 | 3.6032 | 1.6682 | 418.9714 | -4.2486 | -1.7801 | 0 | 0 | 0 | 0 | 0 | 1.0945 | FALSE | 25 | 19.2313
5935 | 5936 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFDA | 82.8720709220048 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy, Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 8.87 | 8.86 | Baseline | Baseline | 1167249.7341 | 56738.2389 | 20.5725 | 202173.5124 | 10392.4954 | 19.4538 | 5.7735 | 5.4595 | 202199.252 | 10403.3233 | 0.99855 | 0.99345 | 8.893 | 8.8909 | 0.25963 | 0.023029 | 8.6746 | 8.7187 | 9.2422 | 9.0778 | 0.56753 | 0.3591 | 0.092253 | 0.085238 | 120.5572 | 98.6256 | 0.00033245 | 0.0020451 | FALSE | 1.0002 | FALSE | [M-H]- | 82.8721 | 8.8003 | 8.9869 | 0.18653 | 8.8117 | 8.9732 | 0.1615 | 118.4288 | 1.0353 | 1.0524 | 0.57398 | 0 | 30 | 7 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Fail | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1167249.7341 | No Control Sample | 0 | off | off | 1167249.7341 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9624 | 4.5287 | 2.323 | 468.9688 | -2.5706 | -1.2055 | 0 | 0 | 0 | 0 | 0 | 2.5686 | FALSE | 25 | 20.5725
5937 | 5938 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFUdA | 100.826616843596 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.21 | 9.2 | Baseline | Baseline | 1032645.8091 | 46182.7705 | 22.36 | 189497.4537 | 9285.241 | 20.4085 | 5.4494 | 4.9738 | 189540.3067 | 9292.4926 | 0.99731 | 0.99399 | 9.2304 | 9.2276 | 0.22158 | 0.020408 | 9.0334 | 9.0787 | 9.5014 | 9.3615 | 0.46798 | 0.28285 | 0.086842 | 0.078185 | 173.4666 | 203.4772 | 0.00054059 | 0.0013827 | FALSE | 1.0003 | TRUE | [M-H]- | 100.8266 | 9.1452 | 9.3165 | 0.17134 | 9.1533 | 9.305 | 0.15177 | 218.896 | 1.0937 | 1.1608 | 0.56887 | 0 | 27 | 6 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1032645.8091 | No Control Sample | 0 | off | off | 1032645.8091 | 0 | off | 0 | off | 0 | off | 0 | off | 562.959 | 3.7754 | 2.1254 | 518.9662 | -2.0808 | -1.0799 | 0 | 0 | 0 | 0 | 0 | 1.5393 | FALSE | 25 | 22.36
5939 | 5940 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFDoA | 80.4972458353941 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 0 | 9.47 | 9.47 | Baseline | Baseline | 1046516.5396 | 50238.6121 | 20.8309 | 203491.2909 | 12012.7149 | 16.9397 | 5.1428 | 4.1821 | 203533.7195 | 12027.9077 | 0.99713 | 0.98119 | 9.4947 | 9.4912 | 0.26044 | 0.024664 | 9.3392 | 9.3624 | 9.7301 | 9.6203 | 0.39083 | 0.25785 | 0.083162 | 0.064454 | 216.4186 | 185.2103 | 0.00054479 | 0.0023006 | FALSE | 1.0004 | FALSE | [M-H]- | 80.4972 | 9.4181 | 9.5709 | 0.15282 | 9.4256 | 9.5616 | 0.13603 | 283.651 | 1.1197 | 1.237 | 1.029 | 0 | 28 | 8 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1046516.5396 | No Control Sample | 0 | off | off | 1046516.5396 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9542 | 0.94808 | 0.58113 | 568.9628 | -1.0187 | -0.57963 | 0 | 0 | 0 | 0 | 0 | 3.1763 | FALSE | 25 | 20.8309
5941 | 5942 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFTrDA | 95.4342473537058 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.68 | 9.47 | Baseline | Baseline | 904387.5782 | 50238.6121 | 18.0018 | 191839.3703 | 12012.7149 | 15.9697 | 4.7143 | 4.1821 | 191909.6785 | 12027.9077 | 0.99567 | 0.98119 | 9.7087 | 9.4912 | 0.29675 | 0.028726 | 9.5595 | 9.3624 | 9.9117 | 9.6203 | 0.35228 | 0.25785 | 0.076712 | 0.064454 | 482.8614 | 185.2103 | 0.00071822 | 0.0023006 | FALSE | 1.0229 | TRUE | [M-H]- | 95.4342 | 9.6399 | 9.7791 | 0.13919 | 9.648 | 9.7711 | 0.12314 | 391.1146 | 1.0197 | 1.0468 | 1.0277 | 0 | 32 | 9 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 904387.5782 | No Control Sample | 0 | off | off | 904387.5782 | 0 | off | 0 | off | 0 | off | 0 | off | 662.951 | 0.83896 | 0.55619 | 618.9591 | -4.1528 | -2.5704 | 0 | 0 | 0 | 0 | 0 | 2.3075 | FALSE | 25 | 18.0018
5943 | 5944 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFTeDA | 105.299201880121 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.87 | 9.86 | Baseline | Baseline | 669525.9811 | 17933.6523 | 37.3335 | 166317.1305 | 4672.0514 | 35.5983 | 4.0256 | 3.8385 | 166416.3447 | 4674.8395 | 0.99418 | 0.99297 | 9.8871 | 9.8861 | 0.17338 | 0.017112 | 9.7545 | 9.7864 | 10.0482 | 10.0078 | 0.29365 | 0.22143 | 0.064217 | 0.061489 | 874.9991 | 601.1954 | 0.0013117 | 0.0012744 | FALSE | 1.0001 | TRUE | [M-H]- | 105.2992 | 9.8273 | 9.9509 | 0.12352 | 9.8348 | 9.9432 | 0.10835 | 742.91 | 1.0947 | 1.2137 | 1.102 | 0 | 34 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 669525.9811 | No Control Sample | 0 | off | off | 669525.9811 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9477 | 0.60112 | 0.42857 | 668.956 | -2.8956 | -1.9371 | 0 | 0 | 0 | 0 | 0 | 2.7426 | FALSE | 25 | 37.3335
5950 | 5951 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFBS | 104.8738310451 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 3.5 | 3.5 | Baseline | Baseline | 960757.5397 | 34610.8867 | 27.7588 | 87080.9895 | 3183.9861 | 27.3497 | 11.0329 | 10.8703 | 87584.0081 | 3254.8085 | 0.97807 | 0.92936 | 3.8078 | 3.808 | 8.7945 | 0.30781 | 3.5184 | 3.5857 | 4.0874 | 4.0167 | 0.56897 | 0.43098 | 0.17486 | 0.1704 | 894.106 | 683.8183 | 0.0093204 | 0.02137 | FALSE | 0.99995 | TRUE | [M-H]- | 104.8738 | 3.6301 | 4.0132 | 0.38301 | 3.6562 | 3.9866 | 0.33043 | 1426.4924 | 1.0661 | 1.1515 | 5.3722 | 0 | 89 | 29 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 960757.5397 | No Control Sample | 0 | off | off | 960757.5397 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9435 | 1.5906 | 0.47551 | 79.9565 | -8.2837 | -0.66234 | 0 | 0 | 0 | 0 | 0 | 1.5765 | FALSE | 25 | 27.7588
5952 | 5953 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFPeS | 104.113428307867 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 6.38 | 3.5 | Baseline | Baseline | 960359.0036 | 34610.8867 | 27.7473 | 174800.1368 | 3183.9861 | 54.8998 | 5.494 | 10.8703 | 174904.8887 | 3254.8085 | 0.99387 | 0.92936 | 6.3705 | 3.808 | 0.14874 | 0.0094893 | 6.2047 | 3.5857 | 6.6103 | 4.0167 | 0.40558 | 0.43098 | 0.086016 | 0.1704 | 299.7433 | 683.8183 | 0.0014072 | 0.02137 | FALSE | 1.6729 | TRUE | [M-H]- | 104.1134 | 6.2849 | 6.4619 | 0.17707 | 6.2962 | 6.4499 | 0.15373 | 606.4853 | 0.98172 | 0.94885 | 4.23 | 0 | 45 | 12 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 960359.0036 | No Control Sample | 0 | off | off | 960359.0036 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9401 | 0.9018 | 0.31467 | 79.9562 | -13.0421 | -1.0428 | 0 | 0 | 0 | 0 | 0 | 1.3462 | FALSE | 25 | 27.7473
5954 | 5955 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | L-PFHxS | 88.7225618753606 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 81.1 | 4 | 20.275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 675090.0891 | 40734.6572 | 16.5729 | 163079.6581 | 9850.2069 | 16.556 | 4.1396 | 4.1354 | 165042.6523 | 9854.8558 | 0.72172 | 0.9915 | 7.3484 | 7.3482 | 0.11467 | 0.0084168 | 7.2605 | 7.2358 | 7.5008 | 7.4922 | 0.24035 | 0.25637 | 0.065027 | 0.063922 | 92.3949 | 161.5107 | 0.014143 | 0.00068981 | FALSE | 1 | TRUE | [M-H]- | 109.399 | 7.2863 | 7.4188 | 0.13249 | 7.2938 | 7.4098 | 0.11605 | -9596.3521 | 1.1786 | 1.3823 | 5.4157 | 0 | 30 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 675090.0891 | No Control Sample | 0 | off | off | 675090.0891 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9373 | 1.6262 | 0.64877 | 79.9563 | -11.8326 | -0.9461 | 0 | 0 | 0 | 0 | 0 | 1.6095 | FALSE | 20.275 | 16.5729
5956 | 5957 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | Br-PFHxS | 21.8081565549867 | Qualifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 18.9 | 4 | 4.725 | 0 | 7.16 | 7.35 | Manual | Baseline | 225450.0903 | 40734.6572 | 5.5346 | 40453.7311 | 9850.2069 | 4.1069 | 5.573 | 4.1354 | 165042.6523 | 9854.8558 | 0.24112 | 0.9915 | 7.1637 | 7.3482 | 0.051696 | 0.0037015 | 7.0506 | 7.2358 | 7.2628 | 7.4922 | 0.21225 | 0.25637 | 0.088479 | 0.063922 | 22.9196 | 161.5107 | 0.019059 | 0.00068981 | TRUE | 0.97489 | TRUE | [M-H]- | 115.3871 | 7.0788 | 7.2687 | 0.18988 | 7.0876 | 7.2423 | 0.15462 | 3632.5138 | 1.1611 | 1.1197 | 0.33396 | 0 | 26 | 12 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Fail | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 225450.0903 | No Control Sample | 0 | off | off | 675090.0891 | 0 | off | 0 | off | 0 | off | 0 | off | 398.9373 | 1.6262 | 0.64877 | 79.9633 | 76.6521 | 6.1289 | 0 | 0 | 0 | 0 | 0 | 1.6095 | FALSE | 4.725 | 5.5346
5958 | 5959 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFHpS | 108.521614254111 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7.98 | 7.35 | Baseline | Baseline | 672786.8506 | 40734.6572 | 16.5163 | 149488.1455 | 9850.2069 | 15.1761 | 4.5006 | 4.1354 | 149615.322 | 9854.8558 | 0.99362 | 0.9915 | 7.99 | 7.3482 | 0.1251 | 0.0099826 | 7.8722 | 7.2358 | 8.3021 | 7.4922 | 0.4299 | 0.25637 | 0.070549 | 0.063922 | 161.4598 | 161.5107 | 0.00071802 | 0.00068981 | FALSE | 1.0873 | TRUE | [M-H]- | 108.5216 | 7.923 | 8.0716 | 0.14859 | 7.9321 | 8.0617 | 0.12963 | -249.6771 | 1.1221 | 1.2702 | 6.4355 | 0 | 31 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 672786.8506 | No Control Sample | 0 | off | off | 672786.8506 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9342 | 1.7653 | 0.7925 | 79.9563 | -11.6405 | -0.93074 | 0 | 0 | 0 | 0 | 0 | 1.7218 | FALSE | 25 | 16.5163
5960 | 5961 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | L-PFOS | 84.091068207707 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 78.8 | 4 | 19.7 | 0 | 8.46 | 8.45 | Manual | Baseline | 516649.099 | 23636.898 | 21.8577 | 101244.7846 | 4509.4337 | 22.4518 | 5.103 | 5.2417 | 102014.0083 | 4532.2088 | 0.59492 | 0.97226 | 8.4757 | 8.4736 | 0.18541 | 0.015686 | 8.3982 | 8.2933 | 8.5922 | 8.6023 | 0.19402 | 0.30893 | 0.081265 | 0.082236 | 46.4882 | 104.8284 | 0.0046865 | 0.0079126 | TRUE | 1.0002 | TRUE | [M-H]- | 106.7146 | 8.3905 | 8.5572 | 0.16662 | 8.3905 | 8.5456 | 0.15505 | -2445.5247 | 1.0266 | 0.9106 | 6.1181 | 0 | 12 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 516649.099 | No Control Sample | 0 | off | off | 516649.099 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9312 | 1.9125 | 0.9542 | 79.9563 | -12.8673 | -1.0288 | 0 | 0 | 0 | 0 | 0 | 0.80809 | FALSE | 19.7 | 21.8577
5962 | 5963 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | Br-PFOS | 24.0773269126371 | Qualifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 21.1 | 4 | 5.275 | 0 | 8.3 | 8.45 | Manual | Baseline | 316691.2856 | 23636.898 | 13.3982 | 53912.2046 | 4509.4337 | 11.9554 | 5.8742 | 5.2417 | 102014.0083 | 4532.2088 | 0.36475 | 0.97226 | 8.3285 | 8.4736 | 0.34303 | 0.028472 | 8.2177 | 8.2933 | 8.399 | 8.6023 | 0.18131 | 0.30893 | 0.098384 | 0.082236 | 24.7547 | 104.8284 | 0.0070718 | 0.0079126 | TRUE | 0.98287 | TRUE | [M-H]- | 114.1106 | 8.2372 | 8.409 | 0.17179 | 8.2459 | 8.409 | 0.16304 | 2102.7507 | 0.86991 | 0.81173 | 0.61297 | 0 | 13 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 316691.2856 | No Control Sample | 0 | off | off | 516649.099 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9312 | 1.9125 | 0.9542 | 79.9569 | -5.3899 | -0.43096 | 0 | 0 | 0 | 0 | 0 | 0.80809 | FALSE | 5.275 | 13.3982
5964 | 5965 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFNS | 108.532739785923 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 8.85 | 8.45 | Baseline | Baseline | 568605.9198 | 23636.898 | 24.0559 | 111224.5891 | 4509.4337 | 24.6649 | 5.1122 | 5.2417 | 111246.8371 | 4532.2088 | 0.99611 | 0.97226 | 8.8733 | 8.4736 | 0.2634 | 0.023311 | 8.7143 | 8.2933 | 9.2075 | 8.6023 | 0.49318 | 0.30893 | 0.079529 | 0.082236 | 142.6591 | 104.8284 | 0.00029613 | 0.0079126 | FALSE | 1.0472 | TRUE | [M-H]- | 108.5327 | 8.7935 | 8.9631 | 0.16953 | 8.8009 | 8.9511 | 0.15017 | -66.7843 | 1.0483 | 1.0441 | 6.9724 | 0 | 26 | 5 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 568605.9198 | No Control Sample | 0 | off | off | 568605.9198 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9271 | 0.22269 | 0.12224 | 79.9563 | -9.8101 | -0.78439 | 0 | 0 | 0 | 0 | 0 | 1.7756 | FALSE | 25 | 24.0559
5966 | 5967 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFDS | 109.362509495819 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.19 | 8.45 | Baseline | Baseline | 588685.282 | 23636.898 | 24.9054 | 120808.2965 | 4509.4337 | 26.7901 | 4.8729 | 5.2417 | 120887.6629 | 4532.2088 | 0.99256 | 0.97226 | 9.2009 | 8.4736 | 0.11841 | 0.010882 | 9.0563 | 8.2933 | 9.3594 | 8.6023 | 0.30312 | 0.30893 | 0.076185 | 0.082236 | 201.7294 | 104.8284 | 0.0011488 | 0.0079126 | FALSE | 1.0858 | TRUE | [M-H]- | 109.3625 | 9.1257 | 9.2844 | 0.15877 | 9.1351 | 9.2745 | 0.13935 | 457.8772 | 0.95946 | 0.90143 | 6.2686 | 0 | 17 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Pass | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 588685.282 | No Control Sample | 0 | off | off | 588685.282 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9237 | -0.22224 | -0.1331 | 79.9562 | -8.2644 | -0.66079 | 0 | 0 | 0 | 0 | 0 | 1.5322 | FALSE | 25 | 24.9054
5970 | 5971 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FBSA | 109.543440638562 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 4.6 | 8.86 | Baseline | Baseline | 1560061.6117 | 54011.0184 | 28.8841 | 134257.267 | 11073.9139 | 12.1237 | 11.6199 | 4.8773 | 134638.4576 | 11076.7537 | 0.98873 | 0.99421 | 4.9407 | 9.017 | 7.4058 | 0.34067 | 4.6259 | 8.8211 | 5.2998 | 9.253 | 0.67388 | 0.43188 | 0.17333 | 0.075369 | 617.8759 | 112.3383 | 0.0012598 | 0.00052676 | FALSE | 0.54793 | TRUE | [M-H]- | 109.5434 | 4.7393 | 5.1446 | 0.4053 | 4.7721 | 5.1186 | 0.34645 | 250.9926 | 1.0728 | 1.2198 | 3.4125 | 0 | 122 | 34 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1560061.6117 | No Control Sample | 0 | off | off | 1560061.6117 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9594 | 1.3549 | 0.40371 | 77.9646 | -9.5558 | -0.74502 | 0 | 0 | 0 | 0 | 0 | 1.3794 | FALSE | 25 | 28.8841
5972 | 5973 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FPeSA | 95.1046314963052 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7 | 8.86 | Baseline | Baseline | 2103310.4232 | 54011.0184 | 38.9422 | 351871.6629 | 11073.9139 | 31.7748 | 5.9775 | 4.8773 | 352044.0133 | 11076.7537 | 0.99472 | 0.99421 | 7.1024 | 9.017 | 1.4634 | 0.10244 | 6.9207 | 8.8211 | 7.3498 | 9.253 | 0.42915 | 0.43188 | 0.096054 | 0.075369 | 199.4803 | 112.3383 | 0.0010707 | 0.00052676 | FALSE | 0.78767 | TRUE | [M-H]- | 95.1046 | 7.0135 | 7.191 | 0.17755 | 7.0243 | 7.1801 | 0.15581 | 877.9172 | 0.97014 | 0.9303 | 1.811 | 0 | 48 | 12 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2103310.4232 | No Control Sample | 0 | off | off | 2103310.4232 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9561 | 1.0292 | 0.35813 | 77.9638 | -18.7343 | -1.4606 | 0 | 0 | 0 | 0 | 0 | 0.80313 | FALSE | 25 | 38.9422
5974 | 5975 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FHxSA | 98.631271876859 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 7.86 | 8.86 | Baseline | Baseline | 1088518.2816 | 54011.0184 | 20.1536 | 221920.3287 | 11073.9139 | 20.0399 | 4.905 | 4.8773 | 222023.673 | 11076.7537 | 0.99457 | 0.99421 | 7.9956 | 9.017 | 1.7255 | 0.13563 | 7.7768 | 8.8211 | 8.3003 | 9.253 | 0.52347 | 0.43188 | 0.077175 | 0.075369 | 140.3535 | 112.3383 | 0.00098331 | 0.00052676 | FALSE | 0.88672 | TRUE | [M-H]- | 98.6313 | 7.9231 | 8.0797 | 0.1566 | 7.9329 | 8.0707 | 0.13784 | 416.8699 | 1.0193 | 1.0546 | 3.5887 | 0 | 38 | 8 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Fail | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1088518.2816 | No Control Sample | 0 | off | off | 1088518.2816 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9535 | 2.4017 | 0.95576 | 77.9643 | -14.2676 | -1.1124 | 0 | 0 | 0 | 0 | 0 | 0.62601 | FALSE | 25 | 20.1536
5976 | 5977 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FHpSA | 75.7450263220128 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 0 | 8.58 | 8.86 | Baseline | Baseline | 1731461.2748 | 54011.0184 | 32.0576 | 315025.6937 | 11073.9139 | 28.4475 | 5.4963 | 4.8773 | 316036.9224 | 11076.7537 | 0.98389 | 0.99421 | 8.5781 | 9.017 | 0.0218 | 0.0018705 | 8.4492 | 8.8211 | 8.749 | 9.253 | 0.2998 | 0.43188 | 0.089036 | 0.075369 | 91.1381 | 112.3383 | 0.002123 | 0.00052676 | FALSE | 0.95132 | FALSE | [M-H]- | 75.745 | 8.4964 | 8.6624 | 0.16597 | 8.5073 | 8.6527 | 0.14545 | -2230.8134 | 1.0363 | 1.1024 | 1.7791 | 0 | 18 | 7 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1731461.2748 | No Control Sample | 0 | off | off | 1731461.2748 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9494 | 0.0063398 | 0.0028399 | 77.9639 | -19.6014 | -1.5282 | 0 | 0 | 0 | 0 | 0 | 0.53242 | FALSE | 25 | 32.0576
5978 | 5979 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FOSA | 95.8143182728449 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 8.88 | 8.86 | Baseline | Baseline | 1177492.7763 | 54011.0184 | 21.801 | 234928.8667 | 11073.9139 | 21.2146 | 5.0121 | 4.8773 | 235058.3601 | 11076.7537 | 0.99192 | 0.99421 | 9.0196 | 9.017 | 1.572 | 0.13959 | 8.8318 | 8.8211 | 9.2066 | 9.253 | 0.37487 | 0.43188 | 0.078714 | 0.075369 | 143.295 | 112.3383 | 0.0010002 | 0.00052676 | FALSE | 1.0003 | TRUE | [M-H]- | 95.8143 | 8.9359 | 9.1016 | 0.16578 | 8.9441 | 9.0895 | 0.14538 | 626.8216 | 1.0269 | 1.0057 | 2.9387 | 0 | 20 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1177492.7763 | No Control Sample | 0 | off | off | 1177492.7763 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9465 | 0.62356 | 0.3105 | 77.964 | -17.3859 | -1.3555 | 0 | 0 | 0 | 0 | 0 | 1.3609 | FALSE | 25 | 21.801
5997 | 5998 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 4:2 FTS | 105.383430943423 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 5.94 | 5.92 | Baseline | Baseline | 648511.9964 | 12119.8269 | 53.5084 | 109132.9806 | 1993.0499 | 54.7568 | 5.9424 | 6.081 | 109186.4732 | 2005.5075 | 0.99581 | 0.97694 | 5.9246 | 5.9226 | 0.26003 | 0.015446 | 5.681 | 5.7849 | 6.1234 | 6.0986 | 0.4423 | 0.31367 | 0.092574 | 0.093225 | 235.3012 | 130.9375 | 0.00085846 | 0.011625 | FALSE | 1.0003 | TRUE | [M-H]- | 105.3834 | 5.8243 | 6.0255 | 0.20119 | 5.8383 | 6.0117 | 0.17337 | 211.8158 | 0.96568 | 0.92312 | 2.9166 | 0 | 56 | 14 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 648511.9964 | No Control Sample | 0 | off | off | 648511.9964 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9749 | 1.8811 | 0.61507 | 306.9671 | -1.5283 | -0.46913 | 0 | 0 | 0 | 0 | 0 | 0.5148 | FALSE | 25 | 53.5084
5999 | 6000 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 6:2 FTS | 96.577309906033 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7.92 | 7.92 | Baseline | Baseline | 461099.7414 | 7869.1398 | 58.596 | 111357.0417 | 1765.198 | 63.0847 | 4.1407 | 4.4579 | 111380.5222 | 1768.7015 | 0.99659 | 0.98503 | 7.9256 | 7.926 | 0.070115 | 0.0055531 | 7.765 | 7.7996 | 8.1928 | 8.0593 | 0.42785 | 0.25972 | 0.062871 | 0.069661 | 99.4412 | 95.1007 | 0.00053379 | 0.0038633 | FALSE | 0.99995 | TRUE | [M-H]- | 96.5773 | 7.8533 | 7.9963 | 0.14297 | 7.8644 | 7.9856 | 0.12121 | 138.9302 | 1.0277 | 1.074 | 1.518 | 0 | 32 | 8 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Pass | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 461099.7414 | No Control Sample | 0 | off | off | 461099.7414 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9702 | 5.4468 | 2.3256 | 406.9618 | 0.83573 | 0.34011 | 0 | 0 | 0 | 0 | 0 | 2.8723 | FALSE | 25 | 58.596
6001 | 6002 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 8:2 FTS | 94.9707208313498 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 8.88 | 8.89 | Baseline | Baseline | 394909.4133 | 8474.8916 | 46.5976 | 71830.1777 | 1597.5326 | 44.9632 | 5.4978 | 5.305 | 71870.0924 | 1601.5058 | 0.99531 | 0.99068 | 8.8863 | 8.8846 | 0.070895 | 0.0062955 | 8.7338 | 8.742 | 9.0943 | 9.0097 | 0.36048 | 0.26773 | 0.085055 | 0.082979 | 106.2299 | 103.8897 | 0.001294 | 0.0044122 | FALSE | 1.0002 | TRUE | [M-H]- | 94.9707 | 8.7979 | 8.9788 | 0.18092 | 8.809 | 8.9676 | 0.15862 | 257.848 | 1.1565 | 1.3607 | 1.5326 | 0 | 19 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 394909.4133 | No Control Sample | 0 | off | off | 394909.4133 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9635 | 3.8362 | 2.0215 | 506.9553 | -0.7733 | -0.39203 | 0 | 0 | 0 | 0 | 0 | 1.7764 | FALSE | 25 | 46.5976
6003 | 6004 | PFAS CS10 400ng/ml | 43 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 44) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 10:2 FTS | 94.998521863679 | Quantifiers | 10/27/2023 06:45:11 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.5 | 8.89 | Baseline | Baseline | 411229.2891 | 8474.8916 | 48.5233 | 92048.5236 | 1597.5326 | 57.6192 | 4.4675 | 5.305 | 92052.176 | 1601.5058 | 0.99804 | 0.99068 | 9.5111 | 8.8846 | 0.11694 | 0.011109 | 9.3606 | 8.742 | 9.6989 | 9.0097 | 0.33828 | 0.26773 | 0.069535 | 0.082979 | 275.5975 | 103.8897 | 8.4738e-05 | 0.0044122 | FALSE | 1.0705 | TRUE | [M-H]- | 94.9985 | 9.4369 | 9.5829 | 0.14605 | 9.4458 | 9.573 | 0.12723 | 23.0575 | 1.0677 | 1.139 | 1.7342 | 0 | 24 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 411229.2891 | No Control Sample | 0 | off | off | 411229.2891 | 0 | off | 0 | off | 0 | off | 0 | off | 626.9566 | 2.3156 | 1.4518 | 606.9491 | 1.9872 | 1.2061 | 0 | 0 | 0 | 0 | 0 | 1.9139 | FALSE | 25 | 48.5233
6064 | 6065 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFBA | 98.5406336853844 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 2 | 0 | 0 | 168.9790 - 168.9990 | 171.9889 - 172.0089 | IDA-13C4_PFBA | PFBA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 1.49 | 1.48 | Manual | Baseline | 1503742.3085 | 85073.9383 | 17.6757 | 185127.1852 | 11165.1585 | 16.5808 | 8.1228 | 7.6196 | 185752.2652 | 11219.0494 | 0.98743 | 0.97687 | 1.4645 | 1.4608 | 1.712 | 0.025509 | 1.3331 | 1.3197 | 1.7126 | 1.6623 | 0.37953 | 0.34263 | 0.11974 | 0.12223 | 210.7956 | 192.2448 | 0 | 0.012107 | TRUE | 1.0025 | TRUE | [M-H]- | 98.5406 | 1.3659 | 1.6286 | 0.26272 | 1.3734 | 1.6043 | 0.23086 | 0 | 1.1736 | 1.211 | 1.1206 | 0 | 79 | 27 | C4HF7O2 | 212.9792 | 168.989 | Gray | Pass | Pass | Pass | 0.02 | 118.3509 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1503742.3085 | No Control Sample | 0 | off | off | 1503742.3085 | 0 | off | 0 | off | 0 | off | 0 | off | 212.9802 | 4.7324 | 1.0079 | 168.988 | -5.8247 | -0.9843 | 0 | 0 | 0 | 0 | 0 | 0.14108 | FALSE | 25 | 17.6757
6066 | 6067 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFPeA | 104.035686447155 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 4 | 0 | 0 | 218.9758 - 218.9958 | 222.9895 - 223.0095 | IDA-13C5_PFPeA | PFPeA | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 3.18 | 3.21 | Manual | Baseline | 1484721.9363 | 63813.822 | 23.2665 | 151313.1752 | 6485.6773 | 23.3304 | 9.8122 | 9.8392 | 151611.6843 | 6594.8673 | 0.99006 | 0.95 | 3.1777 | 3.172 | 0.071542 | 0.002275 | 2.946 | 3.003 | 3.5044 | 3.4092 | 0.55845 | 0.40623 | 0.14911 | 0.15937 | 884.8509 | 525.4819 | 0.00075905 | 0.0085689 | TRUE | 1.0018 | TRUE | [M-H]- | 104.0357 | 3.0228 | 3.3687 | 0.34589 | 3.0425 | 3.3417 | 0.29922 | -205.6658 | 1.1828 | 1.3638 | 1.3302 | 0 | 100 | 29 | C5HF9O2 | 262.976 | 218.9858 | Gray | Pass | Pass | Fail | 0.02 | 91.3301 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1484721.9363 | No Control Sample | 0 | off | off | 1484721.9363 | 0 | off | 0 | off | 0 | off | 0 | off | 262.9771 | 4.1881 | 1.1014 | 218.9852 | -2.8393 | -0.62176 | 0 | 0 | 0 | 0 | 0 | 4.1724 | FALSE | 25 | 23.2665
6068 | 6069 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFHxA | 78.3674303008864 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 6 | 0 | 0 | 268.9724 - 268.9924 | 272.9861 - 273.0061 | IDA-13C5_PFHxA | PFHxA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 1016.0616 | 6.13 | 6.12 | Baseline | Baseline | 2071176.379 | 113562.1062 | 18.2383 | 273283.5702 | 19072.7353 | 14.3285 | 7.5789 | 5.9542 | 273448.4862 | 19088.1211 | 0.99483 | 0.99374 | 6.122 | 6.1157 | 0.1308 | 0.008018 | 5.9205 | 5.9395 | 6.4048 | 6.3072 | 0.48437 | 0.36772 | 0.12577 | 0.094125 | 206.0052 | 213.3055 | 0.0013415 | 0.0010651 | FALSE | 1.001 | FALSE | [M-H]- | 78.3674 | 6.0127 | 6.2296 | 0.21687 | 6.0237 | 6.2163 | 0.19253 | 756.9068 | 0.9742 | 0.91939 | 1.1786 | 0 | 54 | 16 | C6HF11O2 | 312.9728 | 268.9824 | Gray | Fail | Pass | Pass | 0.02 | 74.3543 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2071176.379 | No Control Sample | 203843.5852 | off | off | 2071176.379 | 203843.5852 | off | 0.01 | off | 100 | off | 1016.0616 | off | 312.9736 | 2.5337 | 0.79298 | 268.9809 | -5.6441 | -1.5182 | 0 | 0 | 0 | 0 | 0 | 1.2776 | FALSE | 25 | 18.2383
6070 | 6071 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFHpA | 66.6633356463489 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 8 | 0 | 0 | 318.9691 - 318.9891 | 321.9800 - 322.0000 | IDA-13C4_PFHpA | PFHpA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 785.0175 | 7.26 | 7.27 | Baseline | Baseline | 1792420.8404 | 107454.4491 | 16.6808 | 280686.7745 | 22821.7743 | 12.2991 | 6.3858 | 4.7084 | 281012.8022 | 22843.7098 | 0.9928 | 0.9936 | 7.2791 | 7.2777 | 0.26319 | 0.019108 | 7.1273 | 7.1299 | 7.4676 | 7.421 | 0.34032 | 0.2911 | 0.10457 | 0.075043 | 123.6892 | 211.0747 | 0.0021656 | 0.0019412 | FALSE | 1.0002 | FALSE | [M-H]- | 66.6633 | 7.1936 | 7.3635 | 0.16982 | 7.202 | 7.3544 | 0.15236 | 1786.1753 | 0.90777 | 0.78945 | 1.3392 | 0 | 41 | 15 | C7HF13O2 | 362.9696 | 318.9791 | Gray | Fail | Pass | Pass | 0.02 | 62.7 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1792420.8404 | No Control Sample | 228328.7703 | off | off | 1792420.8404 | 228328.7703 | off | 0.01 | off | 100 | off | 785.0175 | off | 362.9702 | 1.5978 | 0.57994 | 318.9781 | -3.2311 | -1.0307 | 0 | 0 | 0 | 0 | 0 | 0.99997 | FALSE | 25 | 16.6808
6072 | 6073 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFOA | 75.4412828243386 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 10 | 0 | 0 | 368.9661 - 368.9861 | 375.9903 - 376.0103 | IDA-13C8_PFOA | PFOA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 780.772 | 7.96 | 7.94 | Baseline | Baseline | 1494986.431 | 81710.801 | 18.2961 | 270363.1685 | 18263.9822 | 14.8031 | 5.5295 | 4.4739 | 270565.9236 | 18266.7579 | 0.99385 | 0.99763 | 7.9672 | 7.9648 | 0.090648 | 0.0072156 | 7.7375 | 7.8131 | 8.2072 | 8.2623 | 0.46975 | 0.44922 | 0.087625 | 0.068289 | 88.3607 | 93.8118 | 0.0013878 | 0.00045541 | FALSE | 1.0003 | FALSE | [M-H]- | 75.4413 | 7.8922 | 8.0612 | 0.16899 | 7.9015 | 8.0487 | 0.14725 | 798.7361 | 1.1238 | 1.2319 | 1.1868 | 0 | 35 | 9 | C8HF15O2 | 412.9664 | 368.9761 | Gray | Fail | Pass | Pass | 0.02 | 54.2041 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1494986.431 | No Control Sample | 191475.4056 | off | off | 1494986.431 | 191475.4056 | off | 0.01 | off | 100 | off | 780.772 | off | 412.9679 | 3.6269 | 1.4978 | 368.9747 | -3.7715 | -1.3916 | 0 | 0 | 0 | 0 | 0 | 0.27871 | FALSE | 25 | 18.2961
6074 | 6075 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFNA | 81.4719284442561 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 12 | 0 | 0 | 418.9632 - 418.9832 | 426.9909 - 427.0109 | IDA-13C9_PFNA | PFNA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 856.0445 | 8.46 | 8.47 | Baseline | Baseline | 1362338.1834 | 71980.2942 | 18.9265 | 219130.5506 | 13951.5229 | 15.7066 | 6.217 | 5.1593 | 219196.5326 | 13972.8239 | 0.99678 | 0.99215 | 8.4829 | 8.4778 | 0.27016 | 0.022856 | 8.303 | 8.3267 | 8.7322 | 8.6186 | 0.42923 | 0.29193 | 0.10035 | 0.082119 | 163.5699 | 164.3766 | 0.00073642 | 0.0026877 | FALSE | 1.0006 | FALSE | [M-H]- | 81.4719 | 8.3907 | 8.5704 | 0.1797 | 8.3972 | 8.5598 | 0.16257 | 375.9526 | 1.11 | 1.1855 | 0.85805 | 0 | 25 | 7 | C9HF17O2 | 462.9632 | 418.9732 | Gray | Fail | Pass | Pass | 0.02 | 47.7358 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1362338.1834 | No Control Sample | 159143.3885 | off | off | 1362338.1834 | 159143.3885 | off | 0.01 | off | 100 | off | 856.0445 | off | 462.9644 | 2.5696 | 1.1896 | 418.9714 | -4.3163 | -1.8084 | 0 | 0 | 0 | 0 | 0 | 2.2205 | FALSE | 25 | 18.9265
6076 | 6077 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFDA | 84.5444016069181 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 14 | 0 | 0 | 468.9600 - 468.9800 | 473.9767 - 473.9967 | IDA-13C6_PFDA | PFDA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 0 | 8.87 | 8.86 | Baseline | Baseline | 1214166.0391 | 57851.3485 | 20.9877 | 213003.0941 | 10849.4055 | 19.6327 | 5.7002 | 5.3322 | 213169.7489 | 10854.9384 | 0.99459 | 0.99543 | 8.8942 | 8.895 | 0.27227 | 0.024151 | 8.7331 | 8.7232 | 9.0929 | 9.0453 | 0.35978 | 0.32212 | 0.091116 | 0.083711 | 110.3606 | 110.4368 | 0.0015654 | 0.00089045 | FALSE | 0.9999 | FALSE | [M-H]- | 84.5444 | 8.8018 | 8.9883 | 0.18645 | 8.8145 | 8.9756 | 0.16116 | 926.7899 | 0.9997 | 0.99987 | 0.58443 | 0 | 19 | 7 | C10HF19O2 | 512.96 | 468.97 | Gray | Fail | Pass | Pass | 0.02 | 42.6467 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1214166.0391 | No Control Sample | 0 | off | off | 1214166.0391 | 0 | off | 0 | off | 0 | off | 0 | off | 512.9622 | 4.128 | 2.1175 | 468.9687 | -2.7789 | -1.3032 | 0 | 0 | 0 | 0 | 0 | 1.2133 | FALSE | 25 | 20.9877
6078 | 6079 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFUdA | 91.1904267202161 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 16 | 0 | 0 | 518.9573 - 518.9773 | 524.9764 - 524.9964 | IDA-13C7_PFUdA | PFUdA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.21 | 9.2 | Baseline | Baseline | 1037823.0437 | 51318.9607 | 20.223 | 201233.1953 | 10952.4573 | 18.3733 | 5.1573 | 4.6856 | 201399.0862 | 10957.7054 | 0.994 | 0.99632 | 9.23 | 9.2302 | 0.21767 | 0.020048 | 9.0752 | 9.0834 | 9.391 | 9.3808 | 0.31577 | 0.29735 | 0.082167 | 0.07269 | 194.4933 | 169.3603 | 0.0015439 | 0.00090307 | FALSE | 0.99998 | TRUE | [M-H]- | 91.1904 | 9.1483 | 9.3156 | 0.16735 | 9.157 | 9.3041 | 0.14711 | 983.8726 | 1.0508 | 1.0755 | 0.56081 | 0 | 18 | 7 | C11HF21O2 | 562.9568 | 518.9673 | Gray | Pass | Pass | Pass | 0.02 | 38.5381 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1037823.0437 | No Control Sample | 0 | off | off | 1037823.0437 | 0 | off | 0 | off | 0 | off | 0 | off | 562.9583 | 2.5529 | 1.4372 | 518.9655 | -3.4084 | -1.7688 | 0 | 0 | 0 | 0 | 0 | 0.45829 | FALSE | 25 | 20.223
6080 | 6081 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFDoA | 76.6466751235225 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 18 | 0 | 0 | 568.9534 - 568.9734 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFDoA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 0 | 9.47 | 9.47 | Baseline | Baseline | 1001013.5398 | 50468.3596 | 19.8345 | 205268.4691 | 11559.1798 | 17.758 | 4.8766 | 4.3661 | 205315.4316 | 11571.1795 | 0.99687 | 0.98076 | 9.4946 | 9.4971 | 0.25995 | 0.024617 | 9.3417 | 9.3649 | 9.7092 | 9.6227 | 0.36743 | 0.25772 | 0.078219 | 0.069066 | 208.1078 | 157.2788 | 0.00053776 | 0.0019826 | FALSE | 0.99974 | FALSE | [M-H]- | 76.6467 | 9.4213 | 9.5714 | 0.15016 | 9.4285 | 9.5609 | 0.13235 | 300.4237 | 1.1105 | 1.1928 | 1.2485 | 0 | 26 | 8 | C12HF23O2 | 612.9537 | 568.9634 | Gray | Fail | Pass | Pass | 0.02 | 35.1516 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1001013.5398 | No Control Sample | 0 | off | off | 1001013.5398 | 0 | off | 0 | off | 0 | off | 0 | off | 612.9543 | 0.9839 | 0.60309 | 568.9622 | -2.0958 | -1.1924 | 0 | 0 | 0 | 0 | 0 | 3.3151 | FALSE | 25 | 19.8345
6082 | 6083 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFTrDA | 93.9140743826987 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 20 | 0 | 0 | 618.9517 - 618.9717 | 569.9573 - 569.9773 | IDA-13C2_PFDoA | PFTrDA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.68 | 9.47 | Baseline | Baseline | 894051.5798 | 50468.3596 | 17.7151 | 194666.9015 | 11559.1798 | 16.8409 | 4.5927 | 4.3661 | 194719.4988 | 11571.1795 | 0.99627 | 0.98076 | 9.7135 | 9.4971 | 0.34593 | 0.033486 | 9.548 | 9.3649 | 9.8942 | 9.6227 | 0.34622 | 0.25772 | 0.074275 | 0.069066 | 537.499 | 157.2788 | 0.00055249 | 0.0019826 | FALSE | 1.0228 | TRUE | [M-H]- | 93.9141 | 9.6455 | 9.7818 | 0.13635 | 9.6525 | 9.7736 | 0.12111 | 310.6484 | 1.0551 | 1.1025 | 0.87184 | 0 | 31 | 8 | C13HF25O2 | 662.9505 | 618.9617 | Gray | Pass | Pass | Pass | 0.02 | 32.3122 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 894051.5798 | No Control Sample | 0 | off | off | 894051.5798 | 0 | off | 0 | off | 0 | off | 0 | off | 662.9509 | 0.70203 | 0.46541 | 618.9592 | -3.9935 | -2.4718 | 0 | 0 | 0 | 0 | 0 | 2.5966 | FALSE | 25 | 17.7151
6084 | 6085 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFTeDA | 100.670455128634 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 22 | 0 | 0 | 668.9479 - 668.9679 | 669.9530 - 669.9730 | IDA-13C2_PFTeDA | PFTeDA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.87 | 9.86 | Baseline | Baseline | 644185.2623 | 18048.2518 | 35.6924 | 159886.6158 | 4722.4855 | 33.8565 | 4.029 | 3.8218 | 159944.3772 | 4730.4573 | 0.99562 | 0.98927 | 9.8899 | 9.8919 | 0.20118 | 0.019857 | 9.7564 | 9.798 | 10.0848 | 9.987 | 0.32843 | 0.18905 | 0.065044 | 0.061077 | 852.1548 | 544.5926 | 0.0008835 | 0.0035132 | FALSE | 0.99979 | TRUE | [M-H]- | 100.6705 | 9.8317 | 9.9544 | 0.12271 | 9.8382 | 9.9472 | 0.10896 | 430.1028 | 1.14 | 1.304 | 1.0065 | 0 | 39 | 8 | C14HF27O2 | 712.9473 | 668.9579 | Gray | Pass | Pass | Pass | 0.02 | 29.8972 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 644185.2623 | No Control Sample | 0 | off | off | 644185.2623 | 0 | off | 0 | off | 0 | off | 0 | off | 712.9479 | 0.90792 | 0.6473 | 668.9574 | -0.78374 | -0.52429 | 0 | 0 | 0 | 0 | 0 | 3.0958 | FALSE | 25 | 35.6924
6091 | 6092 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFBS | 99.377868214754 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 29 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFBS | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 3.5 | 3.5 | Baseline | Manual | 990980.4016 | 37673.9751 | 26.3041 | 87322.4097 | 3319.7958 | 26.3035 | 11.3485 | 11.3483 | 87701.2803 | 3326.5591 | 0.98261 | 0.98816 | 3.8048 | 3.8046 | 8.7082 | 0.30479 | 3.5003 | 3.5503 | 4.1085 | 4.0828 | 0.60827 | 0.53254 | 0.17074 | 0.16797 | 825.0568 | 769.435 | 0.0069506 | 0.0020749 | FALSE | 1 | TRUE | [M-H]- | 99.3779 | 3.6288 | 4.0155 | 0.38677 | 3.6524 | 3.985 | 0.33262 | 997.8218 | 1.0962 | 1.1772 | 5.2275 | 0 | 95 | 28 | C4HF9SO3 | 298.943 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 990980.4016 | No Control Sample | 0 | off | off | 990980.4016 | 0 | off | 0 | off | 0 | off | 0 | off | 298.9434 | 1.2253 | 0.36628 | 79.9565 | -9.0733 | -0.72547 | 0 | 0 | 0 | 0 | 0 | 0.84165 | FALSE | 25 | 26.3041
6093 | 6094 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFPeS | 94.7068159879911 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 31 | 0 | 0 | 79.9472 - 79.9672 | 79.9474 - 79.9674 | IDA-13C3_PFBS | PFPeS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 6.38 | 3.5 | Baseline | Manual | 950904.3235 | 37673.9751 | 25.2404 | 167438.0891 | 3319.7958 | 50.4363 | 5.6791 | 11.3483 | 167624.1563 | 3326.5591 | 0.99091 | 0.98816 | 6.3755 | 3.8046 | 0.071042 | 0.0045325 | 6.208 | 3.5503 | 6.572 | 4.0828 | 0.36402 | 0.53254 | 0.09031 | 0.16797 | 285.71 | 769.435 | 0.00234 | 0.0020749 | FALSE | 1.6757 | TRUE | [M-H]- | 94.7068 | 6.2884 | 6.4676 | 0.17918 | 6.2997 | 6.4557 | 0.15593 | 1076.3426 | 0.99521 | 0.98011 | 4.5479 | 0 | 40 | 12 | C5HF11SO3 | 348.9398 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 950904.3235 | No Control Sample | 0 | off | off | 950904.3235 | 0 | off | 0 | off | 0 | off | 0 | off | 348.9398 | -0.088254 | -0.030795 | 79.9561 | -13.9808 | -1.1179 | 0 | 0 | 0 | 0 | 0 | 29.2972 | FALSE | 25 | 25.2404
6095 | 6096 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | L-PFHxS | 82.4432920259515 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 33 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 81.1 | 4 | 20.275 | 0 | 7.34 | 7.35 | Baseline | Baseline | 660456.3038 | 42886.951 | 15.3999 | 162289.2018 | 10186.5217 | 15.9318 | 4.0696 | 4.2102 | 164363.9031 | 10202.6847 | 0.71897 | 0.98756 | 7.3541 | 7.3527 | 0.19271 | 0.014145 | 7.2643 | 7.2496 | 7.4776 | 7.5381 | 0.2133 | 0.28842 | 0.064817 | 0.065844 | 101.1972 | 200.4257 | 0.013637 | 0.0014413 | FALSE | 1.0002 | TRUE | [M-H]- | 101.6563 | 7.2907 | 7.4225 | 0.13184 | 7.2981 | 7.4136 | 0.11545 | -10375.4945 | 1.032 | 1.045 | 5.2953 | 0 | 27 | 10 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 660456.3038 | No Control Sample | 0 | off | off | 660456.3038 | 0 | off | 0 | off | 0 | off | 0 | off | 398.937 | 0.86983 | 0.34701 | 79.9562 | -12.2779 | -0.98171 | 0 | 0 | 0 | 0 | 0 | 1.3446 | FALSE | 20.275 | 15.3999
6097 | 6098 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | Br-PFHxS | 19.9868562963248 | Qualifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 35 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHxS | 0 | Integration Quality, Total Width | 0 | 0 | 18.9 | 4 | 4.725 | 0 | 7.16 | 7.35 | Manual | Baseline | 217538.9627 | 42886.951 | 5.0724 | 39988.037 | 10186.5217 | 3.9256 | 5.4401 | 4.2102 | 164363.9031 | 10202.6847 | 0.23687 | 0.98756 | 7.1677 | 7.3527 | 0.1069 | 0.0076539 | 7.0584 | 7.2496 | 7.2647 | 7.5381 | 0.20627 | 0.28842 | 0.08696 | 0.065844 | 24.935 | 200.4257 | 0.025429 | 0.0014413 | TRUE | 0.97484 | TRUE | [M-H]- | 105.7506 | 7.0845 | 7.2725 | 0.18798 | 7.0927 | 7.2472 | 0.15453 | 4929.7842 | 1.1607 | 1.122 | 0.32938 | 0 | 25 | 11 | C6HF13SO3 | 398.9366 | 79.9572 | Gray | Pass | Pass | Fail | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 217538.9627 | No Control Sample | 0 | off | off | 660456.3038 | 0 | off | 0 | off | 0 | off | 0 | off | 398.937 | 0.86983 | 0.34701 | 79.9627 | 69.3791 | 5.5474 | 0 | 0 | 0 | 0 | 0 | 1.3446 | FALSE | 4.725 | 5.0724
6099 | 6100 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFHpS | 99.5967236439223 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 37 | 0 | 0 | 79.9472 - 79.9672 | 79.9472 - 79.9672 | IDA-13C3_PFHxS | PFHpS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7.98 | 7.35 | Baseline | Baseline | 650080.9018 | 42886.951 | 15.158 | 140457.8646 | 10186.5217 | 13.7886 | 4.6283 | 4.2102 | 140575.8783 | 10202.6847 | 0.99295 | 0.98756 | 7.9935 | 7.3527 | 0.16945 | 0.013522 | 7.8756 | 7.2496 | 8.2584 | 7.5381 | 0.38287 | 0.28842 | 0.072874 | 0.065844 | 123.3874 | 200.4257 | 0.00020438 | 0.0014413 | FALSE | 1.0872 | TRUE | [M-H]- | 99.5967 | 7.9259 | 8.076 | 0.15002 | 7.9345 | 8.0656 | 0.13113 | -74.9787 | 1.1228 | 1.2502 | 6.705 | 0 | 28 | 7 | C7HF15SO3 | 448.9334 | 79.9572 | Gray | Pass | Pass | Pass | 0.02 | 250.1338 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 650080.9018 | No Control Sample | 0 | off | off | 650080.9018 | 0 | off | 0 | off | 0 | off | 0 | off | 448.9342 | 1.8344 | 0.82351 | 79.9562 | -12.348 | -0.98731 | 0 | 0 | 0 | 0 | 0 | 1.3666 | FALSE | 25 | 15.158
6101 | 6102 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | L-PFOS | 85.0554613241239 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 39 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Integration Quality, Total Width | 0 | 0 | 78.8 | 4 | 19.7 | 0 | 8.46 | 8.45 | Manual | Baseline | 513091.0388 | 23207.9564 | 22.1084 | 99890.2236 | 4351.3575 | 22.9561 | 5.1365 | 5.3335 | 100999.486 | 4360.6476 | 0.59582 | 0.98086 | 8.4777 | 8.475 | 0.20904 | 0.017685 | 8.3969 | 8.3479 | 8.5911 | 8.6216 | 0.19419 | 0.27368 | 0.081299 | 0.084842 | 47.3916 | 146.4801 | 0.010162 | 0.0014837 | TRUE | 1.0003 | TRUE | [M-H]- | 107.9384 | 8.3948 | 8.5586 | 0.16383 | 8.3948 | 8.548 | 0.15323 | -5227.1818 | 1.0088 | 0.88708 | 6.8268 | 0 | 12 | 6 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Pass | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 513091.0388 | No Control Sample | 0 | off | off | 513091.0388 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9304 | 0.37851 | 0.18885 | 79.9562 | -13.8688 | -1.1089 | 0 | 0 | 0 | 0 | 0 | 1.2891 | FALSE | 19.7 | 22.1084
6103 | 6104 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | Br-PFOS | 24.2919993556582 | Qualifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 41 | 0 | 0 | 79.9473 - 79.9673 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFOS | 0 | Accuracy, Integration Quality, Total Width | 0 | 0 | 21.1 | 4 | 5.275 | 0 | 8.3 | 8.45 | Manual | Baseline | 313716.6167 | 23207.9564 | 13.5176 | 52416.158 | 4351.3575 | 12.0459 | 5.9851 | 5.3335 | 100999.486 | 4360.6476 | 0.36441 | 0.98086 | 8.3299 | 8.475 | 0.36071 | 0.029939 | 8.2197 | 8.3479 | 8.4054 | 8.6216 | 0.18571 | 0.27368 | 0.09855 | 0.084842 | 24.8681 | 146.4801 | 0.0056442 | 0.0014837 | TRUE | 0.98289 | TRUE | [M-H]- | 115.128 | 8.2403 | 8.4132 | 0.173 | 8.2488 | 8.4132 | 0.16449 | -1593.0664 | 0.86558 | 0.7992 | 0.61142 | 0 | 13 | 8 | C8HF17SO3 | 498.9302 | 79.9573 | Gray | Fail | Pass | Fail | 0.02 | 250.1335 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 313716.6167 | No Control Sample | 0 | off | off | 513091.0388 | 0 | off | 0 | off | 0 | off | 0 | off | 498.9304 | 0.37851 | 0.18885 | 79.9581 | 9.7603 | 0.78041 | 0 | 0 | 0 | 0 | 0 | 1.2891 | FALSE | 5.275 | 13.5176
6105 | 6106 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFNS | 112.491254373491 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 43 | 0 | 0 | 79.9471 - 79.9671 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFNS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 8.85 | 8.45 | Baseline | Baseline | 578649.7745 | 23207.9564 | 24.9332 | 112071.3139 | 4351.3575 | 25.7555 | 5.1632 | 5.3335 | 112112.9582 | 4360.6476 | 0.99389 | 0.98086 | 8.875 | 8.475 | 0.28245 | 0.024997 | 8.6318 | 8.3479 | 9.1158 | 8.6216 | 0.484 | 0.27368 | 0.080855 | 0.084842 | 135.7581 | 146.4801 | 0.0007273 | 0.0014837 | FALSE | 1.0472 | TRUE | [M-H]- | 112.4913 | 8.7946 | 8.9624 | 0.16779 | 8.8026 | 8.953 | 0.15043 | 168.4075 | 0.99313 | 0.96615 | 6.8141 | 0 | 26 | 6 | C9HF19SO3 | 548.927 | 79.9571 | Gray | Pass | Pass | Pass | 0.02 | 250.1341 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 578649.7745 | No Control Sample | 0 | off | off | 578649.7745 | 0 | off | 0 | off | 0 | off | 0 | off | 548.9271 | 0.10779 | 0.059166 | 79.9563 | -10.2552 | -0.81997 | 0 | 0 | 0 | 0 | 0 | 1.2221 | FALSE | 25 | 24.9332
6107 | 6108 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | PFDS | 116.398714749285 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 45 | 0 | 0 | 79.9469 - 79.9669 | 79.9472 - 79.9672 | IDA-13C8_PFOS | PFDS | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 0 | 9.19 | 8.45 | Baseline | Baseline | 615190.066 | 23207.9564 | 26.5077 | 124555.4073 | 4351.3575 | 28.6245 | 4.9391 | 5.3335 | 124582.3552 | 4360.6476 | 0.99461 | 0.98086 | 9.2032 | 8.475 | 0.14376 | 0.013211 | 9.0609 | 8.3479 | 9.412 | 8.6216 | 0.35108 | 0.27368 | 0.077143 | 0.084842 | 186.7007 | 146.4801 | 0.00052577 | 0.0014837 | FALSE | 1.0859 | TRUE | [M-H]- | 116.3987 | 9.1273 | 9.2882 | 0.1609 | 9.1381 | 9.2776 | 0.1395 | 186.5292 | 1.2185 | 1.5246 | 4.2759 | 0 | 20 | 6 | C10HF21SO3 | 598.9238 | 79.9569 | Gray | Fail | Pass | Pass | 0.02 | 250.1348 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 615190.066 | No Control Sample | 0 | off | off | 615190.066 | 0 | off | 0 | off | 0 | off | 0 | off | 598.9235 | -0.4761 | -0.28515 | 79.9563 | -6.9831 | -0.55835 | 0 | 0 | 0 | 0 | 0 | 2.7979 | FALSE | 25 | 26.5077
6111 | 6112 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FBSA | 106.595394440898 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 49 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FBSA | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 4.6 | 8.86 | Manual | Baseline | 1520001.608 | 54079.4915 | 28.1068 | 126229.4108 | 11066.6763 | 11.4063 | 12.0416 | 4.8867 | 126437.8029 | 11067.2276 | 0.99335 | 0.99528 | 4.942 | 9.0191 | 7.4351 | 0.34201 | 4.6228 | 8.8051 | 5.3271 | 9.2768 | 0.70439 | 0.47168 | 0.1815 | 0.075716 | 590.1756 | 100.876 | 0.00065763 | 0.00011312 | TRUE | 0.54795 | TRUE | [M-H]- | 106.5954 | 4.7365 | 5.1547 | 0.41817 | 4.766 | 5.1163 | 0.35029 | -117.8511 | 0.90459 | 0.73679 | 3.2222 | 0 | 129 | 35 | C4H2F9NO2S | 297.959 | 77.9653 | Gray | Pass | Pass | Fail | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1520001.608 | No Control Sample | 0 | off | off | 1520001.608 | 0 | off | 0 | off | 0 | off | 0 | off | 297.9595 | 1.8126 | 0.54009 | 77.9645 | -10.0202 | -0.78123 | 0 | 0 | 0 | 0 | 0 | 1.4852 | FALSE | 25 | 28.1068
6113 | 6114 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FPeSA | 93.0782074920682 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 51 | 0 | 0 | 77.9553 - 77.9753 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FPeSA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7 | 8.86 | Baseline | Baseline | 2061104.2153 | 54079.4915 | 38.1125 | 332561.6909 | 11066.6763 | 30.0507 | 6.1977 | 4.8867 | 332726.2314 | 11067.2276 | 0.99468 | 0.99528 | 7.1043 | 9.0191 | 1.4905 | 0.10433 | 6.9264 | 8.8051 | 7.3783 | 9.2768 | 0.45182 | 0.47168 | 0.099037 | 0.075716 | 195.215 | 100.876 | 0.0011993 | 0.00011312 | FALSE | 0.7877 | TRUE | [M-H]- | 93.0782 | 7.0161 | 7.1958 | 0.17964 | 7.0267 | 7.1849 | 0.15814 | 882.7545 | 1.0605 | 1.1345 | 1.7239 | 0 | 51 | 12 | C5H2O2NSF11 | 347.9558 | 77.9653 | Gray | Pass | Pass | Pass | 0.02 | 256.5244 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 2061104.2153 | No Control Sample | 0 | off | off | 2061104.2153 | 0 | off | 0 | off | 0 | off | 0 | off | 347.9561 | 0.81386 | 0.28319 | 77.9639 | -17.7902 | -1.387 | 0 | 0 | 0 | 0 | 0 | 1.2716 | FALSE | 25 | 38.1125
6115 | 6116 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FHxSA | 99.7610937327128 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 53 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHxSA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 7.86 | 8.86 | Baseline | Baseline | 1102383.056 | 54079.4915 | 20.3845 | 236119.3108 | 11066.6763 | 21.3361 | 4.6688 | 4.8867 | 236244.3706 | 11067.2276 | 0.9945 | 0.99528 | 7.9979 | 9.0191 | 1.7551 | 0.13795 | 7.8338 | 8.8051 | 8.2567 | 9.2768 | 0.42292 | 0.47168 | 0.072534 | 0.075716 | 135.5452 | 100.876 | 0.00083315 | 0.00011312 | FALSE | 0.88678 | TRUE | [M-H]- | 99.7611 | 7.9274 | 8.0821 | 0.15475 | 7.9371 | 8.0717 | 0.13454 | 465.1599 | 1.0169 | 1.0279 | 3.3851 | 0 | 31 | 8 | C6H2F13NO2S | 397.9526 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1102383.056 | No Control Sample | 0 | off | off | 1102383.056 | 0 | off | 0 | off | 0 | off | 0 | off | 397.9535 | 2.259 | 0.89897 | 77.9641 | -16.4254 | -1.2806 | 0 | 0 | 0 | 0 | 0 | 1.4869 | FALSE | 25 | 20.3845
6117 | 6118 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FHpSA | 74.1519979008633 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 55 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FHpSA | 0 | Accuracy | 0 | 0 | 100 | 4 | 25 | 0 | 8.58 | 8.86 | Baseline | Baseline | 1697195.0339 | 54079.4915 | 31.3833 | 308807.929 | 11066.6763 | 27.9043 | 5.496 | 4.8867 | 309715.378 | 11067.2276 | 0.98456 | 0.99528 | 8.5797 | 9.0191 | 0.0034598 | 0.00029685 | 8.4534 | 8.8051 | 8.7534 | 9.2768 | 0.3 | 0.47168 | 0.088562 | 0.075716 | 106.3477 | 100.876 | 0.0018253 | 0.00011312 | FALSE | 0.95129 | FALSE | [M-H]- | 74.152 | 8.4983 | 8.6625 | 0.16421 | 8.5071 | 8.6543 | 0.14719 | -1878.8654 | 0.99435 | 0.99419 | 1.8668 | 0 | 18 | 7 | C7H2F15NO2S | 447.9494 | 77.9654 | Gray | Fail | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1697195.0339 | No Control Sample | 0 | off | off | 1697195.0339 | 0 | off | 0 | off | 0 | off | 0 | off | 447.9493 | -0.27564 | -0.12347 | 77.9638 | -20.9688 | -1.6348 | 0 | 0 | 0 | 0 | 0 | 0.84492 | FALSE | 25 | 31.3833
6119 | 6120 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | FOSA | 99.7708082664967 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 57 | 0 | 0 | 77.9554 - 77.9754 | 77.9554 - 77.9754 | IDA-13C8_FOSA | FOSA | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 8.88 | 8.86 | Baseline | Baseline | 1227669.7693 | 54079.4915 | 22.7012 | 254315.4654 | 11066.6763 | 22.9803 | 4.8274 | 4.8867 | 254397.8553 | 11067.2276 | 0.99336 | 0.99528 | 9.0219 | 9.0191 | 1.5975 | 0.14186 | 8.8365 | 8.8051 | 9.25 | 9.2768 | 0.41352 | 0.47168 | 0.074821 | 0.075716 | 154.6774 | 100.876 | 0.00068777 | 0.00011312 | FALSE | 1.0003 | TRUE | [M-H]- | 99.7708 | 8.9402 | 9.1021 | 0.16192 | 8.9482 | 9.0915 | 0.1433 | 422.9813 | 0.99893 | 0.96146 | 2.9059 | 0 | 22 | 6 | C8H2F17NO2S | 497.9462 | 77.9654 | Gray | Pass | Pass | Pass | 0.02 | 256.524 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 1227669.7693 | No Control Sample | 0 | off | off | 1227669.7693 | 0 | off | 0 | off | 0 | off | 0 | off | 497.9463 | 0.27516 | 0.13701 | 77.9641 | -16.1448 | -1.2587 | 0 | 0 | 0 | 0 | 0 | 0.72321 | FALSE | 25 | 22.7012
6138 | 6139 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 4:2 FTS | 98.2383955826609 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 76 | 0 | 0 | 306.9576 - 306.9776 | 80.9549 - 80.9749 | IDA-13C2 4:2 FTS | 4:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 5.94 | 5.92 | Baseline | Baseline | 647982.3121 | 12990.7022 | 49.8805 | 106358.6553 | 2248.4015 | 47.3041 | 6.0924 | 5.7778 | 106413.466 | 2250.1108 | 0.99571 | 0.99423 | 5.9279 | 5.9259 | 0.20305 | 0.012061 | 5.6964 | 5.7674 | 6.1084 | 6.0927 | 0.41202 | 0.32538 | 0.091444 | 0.085315 | 224.4744 | 174.1344 | 0.0009155 | 0.0013237 | FALSE | 1.0003 | TRUE | [M-H]- | 98.2384 | 5.8303 | 6.0294 | 0.1991 | 5.8442 | 6.0169 | 0.17271 | 236.3284 | 1.083 | 1.2135 | 3.0223 | 0 | 52 | 13 | C6H5F9SO3 | 326.9743 | 306.9676 | Gray | Pass | Pass | Pass | 0.02 | 65.1535 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 647982.3121 | No Control Sample | 0 | off | off | 647982.3121 | 0 | off | 0 | off | 0 | off | 0 | off | 326.9749 | 1.8089 | 0.59146 | 306.967 | -1.937 | -0.59461 | 0 | 0 | 0 | 0 | 0 | 1.9704 | FALSE | 25 | 49.8805
6140 | 6141 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 6:2 FTS | 97.5174225021857 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 78 | 0 | 0 | 406.9515 - 406.9715 | 80.9551 - 80.9751 | IDA-13C2 6:2 FTS | 6:2 FTS | 0 | Total Width | 0 | 0 | 100 | 4 | 25 | 0 | 7.92 | 7.92 | Baseline | Baseline | 472226.8919 | 7981.3432 | 59.1663 | 114967.4205 | 1964.7726 | 58.5144 | 4.1075 | 4.0622 | 114985.4487 | 1964.7765 | 0.99737 | 0.98553 | 7.928 | 7.9295 | 0.10159 | 0.0080456 | 7.7555 | 7.7757 | 8.2602 | 8.0936 | 0.50467 | 0.31795 | 0.062285 | 0.061136 | 104.8617 | 92.1532 | 0.00045636 | 4.0872e-06 | FALSE | 0.99982 | TRUE | [M-H]- | 97.5174 | 7.8567 | 7.9988 | 0.1421 | 7.8674 | 7.9885 | 0.12119 | 103.9617 | 1.0211 | 1.0554 | 2.1294 | 0 | 37 | 7 | C8H5F13SO3 | 426.9679 | 406.9615 | Gray | Pass | Pass | Fail | 0.02 | 49.1447 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 472226.8919 | No Control Sample | 0 | off | off | 472226.8919 | 0 | off | 0 | off | 0 | off | 0 | off | 426.9697 | 4.2317 | 1.8068 | 406.9609 | -1.3548 | -0.55133 | 0 | 0 | 0 | 0 | 0 | 1.6013 | FALSE | 25 | 59.1663
6142 | 6143 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 8:2 FTS | 99.4482338506178 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 80 | 0 | 0 | 506.9457 - 506.9657 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 8:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 8.88 | 8.89 | Baseline | Baseline | 386036.4428 | 7911.4781 | 48.7945 | 70255.6878 | 1388.3519 | 50.6037 | 5.4947 | 5.6985 | 70260.6039 | 1390.7715 | 0.99852 | 0.98966 | 8.8892 | 8.8898 | 0.10366 | 0.0092046 | 8.7197 | 8.7465 | 9.2136 | 9.0144 | 0.49388 | 0.26785 | 0.085269 | 0.091056 | 112.2316 | 116.7455 | 0.00020159 | 0.0030845 | FALSE | 0.99994 | TRUE | [M-H]- | 99.4482 | 8.8008 | 8.9824 | 0.18166 | 8.8109 | 8.9703 | 0.15944 | 28.6771 | 1.135 | 1.2791 | 1.5566 | 0 | 26 | 6 | C10H5F17SO3 | 526.9615 | 506.9557 | Gray | Pass | Pass | Pass | 0.02 | 39.4512 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 386036.4428 | No Control Sample | 0 | off | off | 386036.4428 | 0 | off | 0 | off | 0 | off | 0 | off | 526.9637 | 4.088 | 2.1542 | 506.9554 | -0.68764 | -0.3486 | 0 | 0 | 0 | 0 | 0 | 3.8543 | FALSE | 25 | 48.7945
6144 | 6145 | PFAS CS10 400ng/ml | 44 | FALSE | (Path: D:\SCIEX OS Data\Justin Sankey\Data\20231026 PFAS Calibrations.wiff2), (sample Index: 45) | CS10 | 0 | 0 | Standard | 0 | 0 | 0 | 11 | 1 | 10:2 FTS | 101.002754233556 | Quantifiers | 10/27/2023 07:01:47 | 20 | ABSCIEX | 0 | FALSE | 82 | 0 | 0 | 606.9379 - 606.9579 | 80.9548 - 80.9748 | IDA-13C2 8:2 FTS | 10:2 FTS | 0 | 0 | 0 | 0 | 100 | 4 | 25 | 0 | 9.5 | 8.89 | Baseline | Baseline | 408153.8398 | 7911.4781 | 51.5901 | 89544.8789 | 1388.3519 | 64.4972 | 4.5581 | 5.6985 | 89569.1865 | 1390.7715 | 0.99616 | 0.98966 | 9.514 | 8.8898 | 0.1472 | 0.013984 | 9.3632 | 8.7465 | 9.6894 | 9.0144 | 0.3262 | 0.26785 | 0.071624 | 0.091056 | 300.9399 | 116.7455 | 0.00060572 | 0.0030845 | FALSE | 1.0702 | TRUE | [M-H]- | 101.0028 | 9.4403 | 9.5855 | 0.14522 | 9.4486 | 9.5761 | 0.12752 | 166.2753 | 1.0767 | 1.1568 | 2.242 | 0 | 23 | 6 | C12F21H5SO3 | 626.9551 | 606.9479 | Gray | Pass | Pass | Pass | 0.02 | 32.9518 | PFAS_MRM HR_TOF qual_suspects_core list_4060 | PFAS_LC_40_60 | 0 | 0 | 0 | TRUE | PFAS MRM HR Core_confirm TOF_IPS_Suspects_20231026.qmethod | 408153.8398 | No Control Sample | 0 | off | off | 408153.8398 | 0 | off | 0 | off | 0 | off | 0 | off | 626.957 | 2.931 | 1.8376 | 606.949 | 1.874 | 1.1374 | 0 | 0 | 0 | 0 | 0 | 0.34342 | FALSE | 25 | 51.5901

## Calculated concentration
Sample Name | PFBA | 10:2 FTS | 8:2 FTS | 6:2 FTS | 4:2 FTS | FOSA | PFTeDA | PFTrDA | PFDoA | PFUdA | PFDA | PFNA | PFOA | PFHpA | PFHxA | PFPeA | PFBS | FHpSA | PFNS | FHxSA | FPeSA | FBSA | PFDS | Br-PFOS | L-PFOS | PFHpS | Br-PFHxS | L-PFHxS | PFPeS
PFAS QCS0 0ng/ml | [12 empty] | 0.0162671931307739 | 0.0353564961081969
PFBA_72hr_Dup | 1.78099064689806 | [12 empty] | 0.00329909834357048 | 0.012862582429304
PFBS_72hr_Dup | [13 empty] | 0.00350146886802616 |  |  | 0.995083222314166
PFOA_72hr | [12 empty] | 4.2034233950993 | 0.00974767629483499 | 0.0162205061249421
PFOS_72hr_Dup | [13 empty] | 0.00545501410327587 | 0.0208537360611477 | [8 empty] | 0.148578268365655 | 0.983040087137568
Standard_check_327ai_343g
Standard_check_327aj_343g | [5 empty] | 0.0139661654663983 | [8 empty] | 0.00966781769528913
Standard_check_327aj_343o

## TOF Calculated concentration
Sample Name | PFBA_TOF MS | 10:2 FTS_TOF MS | 8:2 FTS_TOF MS | 6:2 FTS_TOF MS | 4:2 FTS_TOF MS | FOSA_TOF MS | PFTeDA_TOF MS | PFTrDA_TOF MS | PFDoA_TOF MS | PFUdA_TOF MS | PFDA_TOF MS | PFNA_TOF MS | PFOA_TOF MS | PFHpA_TOF MS | PFHxA_TOF MS | PFPeA_TOF MS | PFBS_TOF MS | FHpSA_TOF MS | PFNS_TOF MS | FHxSA_TOF MS | FPeSA_TOF MS | FBSA_TOF MS | PFDS_TOF MS | Br-PFOS_TOF MS | L-PFOS_TOF MS | PFHpS_TOF MS | Br-PFHxS_TOF MS | L-PFHxS_TOF MS | PFPeS_TOF MS
PFAS QCS0 0ng/ml |  |  |  |  | 0.0489599086416034 | [7 empty] | 0.00783657244185009 | 0.0220117239549493 | 0.0151374996161742 |  |  | 0.0136137807512454 | [5 empty] | 0.00220304523338136 | 0.0060187568061587 |  | 0.0190736060125838 | 0.0202188283844378
PFBA_72hr_Dup | 1.36788815947709 |  |  |  | 0.0563922085361553 | [9 empty] | 0.0103553126070929 |  |  | 0.0131971597459451 | [8 empty] | 0.00175802388493325 | 0.00214384479019904
PFBS_72hr_Dup |  |  |  |  | 0.053646516584077 | 0.00208720345773131 |  |  | 0.00474273815948231 | [7 empty] | 0.908592111628908 | 0.00716342761763115 | [8 empty] | 0.00188911068613336 | 0.00200253715720763
PFOA_72hr |  |  |  |  | 0.0620721312888851 | [7 empty] | 4.08329257635442 | 0.00833436020033182 | 0.0110691684000553 |  |  | 0.0148289313319647 | [8 empty] | 0.00344476232055755 | 0.00365159362831443
PFOS_72hr_Dup |  |  |  |  | 0.0520494328883192 |  |  |  | 0.00331126654452556 |  |  |  | 0.00309929386493177 | 0.00484133986096783 | 0.0127002798609048 |  |  | 0.00690965068312112 | [5 empty] | 0.073582676293974 | 0.673828768873611 | 0.0178095398189149 | 0.002030942553406 | 0.00215288493003768
Standard_check_327ai_343g |  |  |  |  | 0.0279218909711343 | 0.00278076392024542 | [7 empty] | 0.00281690457070918 |  |  |  | 0.00932967617931215 | [6 empty] | 0.000442888899288632 |  | 0.0023068104781044 | 0.00244531658782531
Standard_check_327aj_343g |  |  |  |  | 0.0488814210221799 | 0.00934295988508168 | [8 empty] | 0.00668749068038553 |  |  | 0.01457117491849 | [5 empty] | 0.0021418480474113 | 0.0010101403373563 |  | 0.00206798027980715 | 0.00219214648516054
Standard_check_327aj_343o |  |  |  |  | 0.0560468957535074 | 0.00723652513692072 | [7 empty] | 0.00347912761268485 |  |  |  | 0.00937177636320419 | [5 empty] | 0.0022169238116042 | 0.00263855889588486 |  | 0.00231331620382041 | 0.00245221293200278

## Concentration Difference (%)
Sample Name | PFBA | 10:2 FTS | 8:2 FTS | 6:2 FTS | 4:2 FTS | FOSA | PFTeDA | PFTrDA | PFDoA | PFUdA | PFDA | PFNA | PFOA | PFHpA | PFHxA | PFPeA | PFBS | FHpSA | PFNS | FHxSA | FPeSA | FBSA | PFDS | Br-PFOS | L-PFOS | PFHpS | Br-PFHxS | L-PFHxS | PFPeS
PFAS QCS0 0ng/ml | [12 empty] | 70 | 46.5
PFBA_72hr_Dup | 26.2 | [13 empty] | 21.6
PFBS_72hr_Dup | [16 empty] | 9.1
PFOA_72hr | [12 empty] | 2.9 | 15.6 | 37.8
PFOS_72hr_Dup | [13 empty] | 11.9 | 48.6 | [8 empty] | 67.5 | 37.3
Standard_check_327ai_343g
Standard_check_327aj_343g | [5 empty] | 39.7 | [8 empty] | 36.4
Standard_check_327aj_343o